## Instructions
Picarro Isotope Analyzer
Data Processor 3.0 (KAP)
How to Use this Spreadsheet:
Step 1: Retrieve raw result file from Piccaro CPU (C:Piccaro\IsotopeData). Add initials and project name to end
of file name. Copy file into your folder on Water Server.
Step 2: On syncing computer, move raw file into Raw Data folder on Google Drive:
Instruments and Analyses --> Picarro Water Isotope Analyzer --> Raw Data
Step 3: Save this spreadsheet in the Processed Data folder on Google Drive:
Instruments and Analyses --> Picarro Water Isotope Analyzer --> Processed Data
Use file name: YYYYMMDD_Initials_(anything else to identify ex. project, run number, etc.)
Step 4: Enter run information and sample positions in the "Tray Configuration" tab. (orange cells)
(It is important to use the same names for the standards throughout this spreadsheet)
Step 5: Copy text from raw result file and paste into "Picarro Output" tab.
Step 6: In Run Summary tab, check that H2O Mean is within "good" bounds. (between 17,000 - 23,000)
Step 7: In 18O tab, verify that the reference defined values match the references used in run (cells U39 - 42). Run Solver and
verify stdev memory corrected are reduced (column K values less than corresponding column I values). Verify stdev of
drift corrected reference is lower than memory corrected. Verify normalization R2 is 0.99+. Verify QC checks are within
0.15 tolerance.
Step 8: In dD tab, verify that the reference defined values match the references used in run (cells U39 - 42. Run Solver and
verify stdev memory corrected are reduced (column K values less than corresponding column I values). Verify stdev of
drift corrected reference is lower than memory corrected. Verify normalization R2 is 0.99+. Verify QC checks are within
1.5 tolerance.
Step 9: Fill in Quality Assurance Verification chart on RESULTS tab. Dr. Scott or Kelly Peeler will fill in as reviewer 2.
Step 10: Add reference results to Water Isotope Analyzer run log.

## Tray Configuration
| Analyst |  | DWH
| PI |  | Durelle Scott
| Project |  | CCR
| Sampling Date |  | 2024 - 2025
| Sampling Location |  | Carvins Cove Reservoir
| Analysis Date |  | 2025-03-14
Notes
Rear Rack
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
1 | Blacksburg | Myrtle | Homer | Blacksburg | Hawaii | CC4 17feb25 BOT | C50 17feb25 20m | C50 17feb25 6m | Blacksburg | 9
10 | CC4 17feb25 1.5m | CC3 15aug24 0.1m | CP1 17feb25 BOT | CC5 22may24 13m | CC2 22may24 BOT | CP2 17feb25 BOT | CS1 15aug24 0.1m | CC2 17feb25 BOT | Blacksburg | 18
19 | CS1 17feb25 0.1m | CC2 17feb25 0.1m | CS1 22may24 0.1m | CP1 17feb25 0.1m | CP2 17feb25 0.1m | CC3 17feb25 BOT | CS1 15aug24 0.1m - DUP | Hawaii | Blacksburg | 27
28 | [9 empty] | 36
37 | [9 empty] | 45
46 | [9 empty] | 54
Rear Rack
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
1 | Drift | High reference | Low Reference | Drift | QC Check | Sample | Sample | Sample | Drift | 9
10 | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Drift | 18
19 | Sample | Sample | Sample | Sample | Sample | Sample | Sample | QC Check | Drift | 27
28 | [9 empty] | 36
37 | [9 empty] | 45
46 | [9 empty] | 54
Standards
|  | Standard ID | Actual d(18_16) | Actual d(D_H)
| #1 | Denali | -19.9287 | -153.7172 | [6 empty] | -3.37
| #2 | Alberta | -18.8392 | -146.8245 | [6 empty] | -20.92
| #3 | MT BHB | -18.431 | -140.5778 | [6 empty] | -17.35
| #4 | Anchorage | -16.4204 | -129.9028 | [6 empty] | -7.39
| #5 | Homer | -14.5399 | -110.5169
| #6 | MT Creek | -13.3562 | -97.7428
| #7 | Nickel Mines | -7.876 | -48.7951
| #8 | Blacksburg | -7.5615 | -47.446
| #9 | Denton's Well | -6.9583 | -42.6011
| #10 | Hawaii | -3.6108 | -13.2294
| #11 | Myrtle Beach | -2.2908 | -10.453
| #12 | Heavy Oxygen | 3.1753 | -15.7881
| #13 | Heavy Hydrogen | 11.1372 | 4.0332
| #14 | Andres Heavy Water | 17.5664 | 20.5382
| #15 | Picarro ZERO | 0.3 | 1.8
| #16 | Picarro MID | -20.6 | -159
| #17 | Picarro DEPL | -29.6 | -235
#18
#19
#20
|  | DI | -6 | -36
Note on C34: Nate: The Standard ID must match the Standard ID used above in the trey protocol. For example, if you are using the Boulder Colorado Standard, and you labeled it "B" above, it must be labled "B" in this collum.

## Timing
| Enter start time and date --> | 2019-01-10 09:30
Tray Position | Sample ID | Start Time | # Injections
1 | Blacksburg | 2019-01-10 10:58 | 10
2 | Myrtle | 2019-01-10 12:27 | 10
3 | Homer | 2019-01-10 13:55 | 10
4 | Blacksburg | 2019-01-10 15:23 | 10
5 | Hawaii | 2019-01-10 15:59 | 4
6 | CC4 17feb25 BOT | 2019-01-10 16:34 | 4
7 | C50 17feb25 20m | 2019-01-10 17:09 | 4
8 | C50 17feb25 6m | 2019-01-10 17:45 | 4
9 | Blacksburg | 2019-01-10 18:20 | 4
10 | CC4 17feb25 1.5m | 2019-01-10 18:55 | 4
11 | CC3 15aug24 0.1m | 2019-01-10 19:31 | 4
12 | CP1 17feb25 BOT | 2019-01-10 20:06 | 4
13 | CC5 22may24 13m | 2019-01-10 20:41 | 4
14 | CC2 22may24 BOT | 2019-01-10 21:17 | 4
15 | CP2 17feb25 BOT | 2019-01-10 21:52 | 4
16 | CS1 15aug24 0.1m | 2019-01-10 22:27 | 4
17 | CC2 17feb25 BOT | 2019-01-10 23:03 | 4
18 | Blacksburg | 2019-01-10 23:38 | 4
19 | CS1 17feb25 0.1m | 2019-01-11 00:13 | 4
20 | CC2 17feb25 0.1m | 2019-01-11 00:49 | 4
21 | CS1 22may24 0.1m | 2019-01-11 01:24 | 4
22 | CP1 17feb25 0.1m | 2019-01-11 01:59 | 4
23 | CP2 17feb25 0.1m | 2019-01-11 02:35 | 4
24 | CC3 17feb25 BOT | 2019-01-11 03:10 | 4
25 | CS1 15aug24 0.1m - DUP | 2019-01-11 03:45 | 4
26 | Hawaii | 2019-01-11 04:21 | 4
27 | Blacksburg | 2019-01-11 04:56 | 4
Formulas:
Start Time = [Start Time, row above]+"0:08:50"*10

## Picarro Output
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | days_Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | BASELINE_Mean | BASELINE_SD | DAS Temp | Tray | Sample | Job | Method
1 | P-11566 | 2025/03/14 09:56:57 | MT1-Rear-01 | 1 | -10.405 | -46.647 | 21368 | -1 | 1 |  |  | H2O | 63.4329 | 0.133 | 0.443 | 89 | -0.002 | -0.109 | 29.41 | 845 | 0 | 30.625 | MT1-Rear | 1 | 66 | BSE
2 | P-11566 | 2025/03/14 10:05:43 | MT1-Rear-01 | 2 | -10.211 | -43.011 | 20722 | -1 | 1 |  |  | H2O | 63.4391 | 0.101 | 0.284 | 90 | -0.024 | 0.042 | 29.59 | 845 | 0 | 29.375 | MT1-Rear | 1 | 66 | BSE
3 | P-11566 | 2025/03/14 10:14:30 | MT1-Rear-01 | 3 | -10.116 | -42.545 | 19833 | -1 | 1 |  |  | H2O | 63.4451 | 0.182 | 0.494 | 86 | -0.032 | -0.064 | 28.47 | 844 | 0 | 30.25 | MT1-Rear | 1 | 66 | BSE
4 | P-11566 | 2025/03/14 10:23:16 | MT1-Rear-01 | 4 | -10.119 | -42.237 | 19730 | 0 | 1 |  |  | H2O | 63.4513 | 0.105 | 0.438 | 87 | -0.008 | -0.091 | 28.68 | 844 | 0 | 30.25 | MT1-Rear | 1 | 66 | BSE
5 | P-11566 | 2025/03/14 10:32:03 | MT1-Rear-01 | 5 | -10.105 | -41.934 | 19754 | 0 | 1 |  |  | H2O | 63.4574 | 0.187 | 0.42 | 90 | -0.054 | 0.003 | 29.58 | 844 | 0 | 29.938 | MT1-Rear | 1 | 66 | BSE
6 | P-11566 | 2025/03/14 10:40:49 | MT1-Rear-01 | 6 | -10.045 | -42.269 | 19729 | 0 | 1 |  |  | H2O | 63.4634 | 0.145 | 0.341 | 86 | -0.025 | -0.046 | 28.41 | 845 | 0 | 29.938 | MT1-Rear | 1 | 66 | BSE
7 | P-11566 | 2025/03/14 10:49:35 | MT1-Rear-01 | 7 | -9.944 | -42.343 | 19739 | 0 | 1 |  |  | H2O | 63.4695 | 0.113 | 0.281 | 90 | -0.022 | -0.034 | 29.56 | 844 | 0 | 30.313 | MT1-Rear | 1 | 66 | BSE
8 | P-11566 | 2025/03/14 10:58:22 | MT1-Rear-01 | 8 | -10.057 | -42.167 | 19680 | 0 | 1 |  |  | H2O | 63.4757 | 0.103 | 0.335 | 87 | -0.004 | -0.048 | 28.82 | 844 | 0 | 30.188 | MT1-Rear | 1 | 66 | BSE
9 | P-11566 | 2025/03/14 11:07:08 | MT1-Rear-01 | 9 | -10.057 | -42.147 | 19576 | 0 | 1 |  |  | H2O | 63.4817 | 0.191 | 0.341 | 87 | -0.048 | -0.025 | 28.68 | 844 | 0 | 30.75 | MT1-Rear | 1 | 66 | BSE
10 | P-11566 | 2025/03/14 11:15:55 | MT1-Rear-01 | 10 | -9.999 | -42.018 | 19642 | 0 | 1 |  |  | H2O | 63.4878 | 0.189 | 0.289 | 90 | -0.041 | 0.014 | 29.78 | 844 | 0 | 29.875 | MT1-Rear | 1 | 66 | BSE
11 | P-11567 | 2025/03/14 11:24:42 | MT1-Rear-02 | 1 | -5.096 | -8.127 | 19914 | -1 | 1 |  |  | H2O | 63.494 | 0.154 | 0.206 | 89 | -0.028 | -0.013 | 29.46 | 844 | 0 | 30.125 | MT1-Rear | 2 | 66 | BSE
12 | P-11567 | 2025/03/14 11:33:28 | MT1-Rear-02 | 2 | -4.803 | -4.225 | 19586 | -1 | 1 |  |  | H2O | 63.5 | 0.148 | 0.298 | 82 | -0.037 | -0.014 | 29.76 | 844 | 0 | 30.25 | MT1-Rear | 2 | 66 | BSE
13 | P-11567 | 2025/03/14 11:42:15 | MT1-Rear-02 | 3 | -4.673 | -3.581 | 19629 | -1 | 1 |  |  | H2O | 63.5061 | 0.084 | 0.26 | 89 | -0.003 | -0.015 | 29.39 | 844 | 0 | 30.25 | MT1-Rear | 2 | 66 | BSE
14 | P-11567 | 2025/03/14 11:51:02 | MT1-Rear-02 | 4 | -4.676 | -3.305 | 19726 | 0 | 1 |  |  | H2O | 63.5122 | 0.151 | 0.339 | 87 | -0.035 | -0.066 | 28.62 | 845 | 0 | 30.125 | MT1-Rear | 2 | 66 | BSE
15 | P-11567 | 2025/03/14 11:59:50 | MT1-Rear-02 | 5 | -4.796 | -2.96 | 19941 | 0 | 1 |  |  | H2O | 63.5184 | 0.144 | 0.454 | 88 | -0.018 | -0.064 | 28.91 | 844 | 0 | 30 | MT1-Rear | 2 | 66 | BSE
16 | P-11567 | 2025/03/14 12:08:36 | MT1-Rear-02 | 6 | -4.642 | -3.021 | 19713 | 0 | 1 |  |  | H2O | 63.5243 | 0.076 | 0.404 | 86 | -0.021 | -0.076 | 28.39 | 844 | 0 | 30.313 | MT1-Rear | 2 | 66 | BSE
17 | P-11567 | 2025/03/14 12:17:23 | MT1-Rear-02 | 7 | -4.651 | -2.744 | 19690 | 0 | 1 |  |  | H2O | 63.5305 | 0.156 | 0.284 | 86 | -0.021 | -0.072 | 28.44 | 844 | 0 | 30.25 | MT1-Rear | 2 | 66 | BSE
18 | P-11567 | 2025/03/14 12:26:10 | MT1-Rear-02 | 8 | -4.644 | -2.876 | 19541 | 0 | 1 |  |  | H2O | 63.5366 | 0.095 | 0.374 | 87 | -0.017 | -0.056 | 28.65 | 844 | 0 | 30.5 | MT1-Rear | 2 | 66 | BSE
19 | P-11567 | 2025/03/14 12:34:57 | MT1-Rear-02 | 9 | -4.611 | -2.793 | 19664 | 0 | 1 |  |  | H2O | 63.5427 | 0.171 | 0.364 | 87 | -0.036 | -0.084 | 28.66 | 844 | 0 | 30.375 | MT1-Rear | 2 | 66 | BSE
20 | P-11567 | 2025/03/14 12:43:43 | MT1-Rear-02 | 10 | -4.572 | -2.983 | 19627 | 0 | 1 |  |  | H2O | 63.5487 | 0.147 | 0.407 | 87 | -0.022 | -0.094 | 28.54 | 845 | 0 | 30.25 | MT1-Rear | 2 | 66 | BSE
21 | P-11568 | 2025/03/14 12:52:30 | MT1-Rear-03 | 1 | -16.374 | -93.746 | 19941 | -1 | 1 |  |  | H2O | 63.5549 | 0.122 | 0.548 | 88 | -0.025 | 0.144 | 29.09 | 844 | 0 | 30.125 | MT1-Rear | 3 | 66 | BSE
22 | P-11568 | 2025/03/14 13:01:16 | MT1-Rear-03 | 2 | -17.045 | -103.773 | 19770 | -1 | 1 |  |  | H2O | 63.561 | 0.149 | 0.475 | 87 | -0.006 | 0.115 | 28.64 | 844 | 0 | 30.25 | MT1-Rear | 3 | 66 | BSE
23 | P-11568 | 2025/03/14 13:10:03 | MT1-Rear-03 | 3 | -17.283 | -105.376 | 19788 | -1 | 1 |  |  | H2O | 63.5671 | 0.145 | 0.497 | 90 | -0.009 | 0.114 | 29.71 | 844 | 0 | 30.688 | MT1-Rear | 3 | 66 | BSE
24 | P-11568 | 2025/03/14 13:18:50 | MT1-Rear-03 | 4 | -17.398 | -106.349 | 19952 | 0 | 1 |  |  | H2O | 63.5733 | 0.164 | 0.337 | 89 | 0.022 | 0.082 | 29.38 | 844 | 0 | 30.625 | MT1-Rear | 3 | 66 | BSE
25 | P-11568 | 2025/03/14 13:27:35 | MT1-Rear-03 | 5 | -17.308 | -106.773 | 19931 | 0 | 1 |  |  | H2O | 63.5792 | 0.081 | 0.395 | 101 | -0.006 | 0.074 | 30.37 | 844 | 0 | 30.688 | MT1-Rear | 3 | 66 | BSE
26 | P-11568 | 2025/03/14 13:36:22 | MT1-Rear-03 | 6 | -17.227 | -107.201 | 19835 | 0 | 1 |  |  | H2O | 63.5854 | 0.153 | 0.42 | 88 | -0.033 | 0.08 | 28.9 | 844 | 0 | 30.5 | MT1-Rear | 3 | 66 | BSE
27 | P-11568 | 2025/03/14 13:45:09 | MT1-Rear-03 | 7 | -17.284 | -107.375 | 19815 | 0 | 1 |  |  | H2O | 63.5915 | 0.109 | 0.649 | 89 | 0.012 | 0.164 | 29.5 | 844 | 0 | 30.188 | MT1-Rear | 3 | 66 | BSE
28 | P-11568 | 2025/03/14 13:53:57 | MT1-Rear-03 | 8 | -17.33 | -107.753 | 20367 | 0 | 1 |  |  | H2O | 63.5976 | 0.109 | 0.356 | 91 | 0.005 | 0.061 | 29.98 | 844 | 0 | 30.188 | MT1-Rear | 3 | 66 | BSE
29 | P-11568 | 2025/03/14 14:02:45 | MT1-Rear-03 | 9 | -17.325 | -107.663 | 19880 | 0 | 1 |  |  | H2O | 63.6038 | 0.13 | 0.353 | 89 | -0.016 | 0.098 | 29.27 | 844 | 0 | 30.438 | MT1-Rear | 3 | 66 | BSE
30 | P-11568 | 2025/03/14 14:11:32 | MT1-Rear-03 | 10 | -17.342 | -107.951 | 19705 | 0 | 1 |  |  | H2O | 63.6097 | 0.105 | 0.464 | 88 | -0.023 | 0.088 | 29.18 | 844 | 0 | 30.375 | MT1-Rear | 3 | 66 | BSE
31 | P-11569 | 2025/03/14 14:20:19 | MT1-Rear-04 | 1 | -10.498 | -50.908 | 19954 | -1 | 1 |  |  | H2O | 63.6159 | 0.104 | 0.267 | 88 | -0.006 | 0.015 | 29.06 | 844 | 0 | 30.063 | MT1-Rear | 4 | 66 | BSE
32 | P-11569 | 2025/03/14 14:29:07 | MT1-Rear-04 | 2 | -10.178 | -44.518 | 19963 | -1 | 1 |  |  | H2O | 63.622 | 0.1 | 0.292 | 88 | -0.008 | 0 | 29.2 | 844 | 0 | 30.188 | MT1-Rear | 4 | 66 | BSE
33 | P-11569 | 2025/03/14 14:37:54 | MT1-Rear-04 | 3 | -10.06 | -43.475 | 20272 | -1 | 1 |  |  | H2O | 63.6281 | 0.113 | 0.561 | 105 | 0.001 | 0.119 | 31.62 | 844 | 0 | 30.5 | MT1-Rear | 4 | 66 | BSE
34 | P-11569 | 2025/03/14 14:46:42 | MT1-Rear-04 | 4 | -10.117 | -42.532 | 20632 | 0 | 1 |  |  | H2O | 63.6343 | 0.173 | 0.247 | 90 | 0.005 | -0.01 | 29.87 | 844 | 0 | 30.438 | MT1-Rear | 4 | 66 | BSE
35 | P-11569 | 2025/03/14 14:55:29 | MT1-Rear-04 | 5 | -10.071 | -42.567 | 19981 | 0 | 1 |  |  | H2O | 63.6403 | 0.113 | 0.457 | 87 | -0.015 | 0.012 | 28.67 | 844 | 0 | 30.813 | MT1-Rear | 4 | 66 | BSE
36 | P-11569 | 2025/03/14 15:04:17 | MT1-Rear-04 | 6 | -10.048 | -42.225 | 19553 | 0 | 1 |  |  | H2O | 63.6464 | 0.122 | 0.407 | 87 | -0.011 | -0.069 | 28.57 | 844 | 0 | 30.188 | MT1-Rear | 4 | 66 | BSE
37 | P-11569 | 2025/03/14 15:13:04 | MT1-Rear-04 | 7 | -10.009 | -42.175 | 19565 | 0 | 1 |  |  | H2O | 63.6525 | 0.083 | 0.399 | 88 | -0.011 | 0.029 | 28.94 | 844 | 0 | 30.813 | MT1-Rear | 4 | 66 | BSE
38 | P-11569 | 2025/03/14 15:21:52 | MT1-Rear-04 | 8 | -10.004 | -42.105 | 19924 | 0 | 1 |  |  | H2O | 63.6586 | 0.124 | 0.212 | 86 | -0.005 | -0.011 | 28.45 | 844 | 0 | 30.688 | MT1-Rear | 4 | 66 | BSE
39 | P-11569 | 2025/03/14 15:30:40 | MT1-Rear-04 | 9 | -10.096 | -42.021 | 19566 | 0 | 1 |  |  | H2O | 63.6647 | 0.151 | 0.328 | 85 | -0.027 | -0.017 | 28.06 | 844 | 0 | 30.563 | MT1-Rear | 4 | 66 | BSE
40 | P-11569 | 2025/03/14 15:39:28 | MT1-Rear-04 | 10 | -9.944 | -42.073 | 19989 | 0 | 1 |  |  | H2O | 63.6708 | 0.106 | 0.232 | 88 | -0.02 | 0.058 | 28.97 | 844 | 0 | 30.25 | MT1-Rear | 4 | 66 | BSE
41 | P-11570 | 2025/03/14 15:48:15 | MT1-Rear-05 | 1 | -6.27 | -10.476 | 19972 | -1 | 1 |  |  | H2O | 63.6769 | 0.135 | 0.409 | 90 | -0.035 | 0.013 | 29.76 | 844 | 0 | 30 | MT1-Rear | 5 | 67 | BSE
42 | P-11570 | 2025/03/14 15:57:02 | MT1-Rear-05 | 2 | -6.124 | -6.728 | 19944 | -1 | 1 |  |  | H2O | 63.6831 | 0.083 | 0.392 | 95 | 0.003 | 0.085 | 31.22 | 844 | 0 | 30.25 | MT1-Rear | 5 | 67 | BSE
43 | P-11570 | 2025/03/14 16:05:50 | MT1-Rear-05 | 3 | -6.131 | -5.763 | 19959 | -1 | 1 |  |  | H2O | 63.6892 | 0.123 | 0.241 | 86 | 0.001 | 0.004 | 28.42 | 844 | 0 | 30.938 | MT1-Rear | 5 | 67 | BSE
44 | P-11570 | 2025/03/14 16:14:39 | MT1-Rear-05 | 4 | -6.032 | -5.573 | 19939 | 0 | 1 |  |  | H2O | 63.6953 | 0.142 | 0.32 | 89 | -0.027 | -0.034 | 29.29 | 844 | 0 | 30.938 | MT1-Rear | 5 | 67 | BSE
45 | P-11571 | 2025/03/14 16:23:27 | MT1-Rear-06 | 1 | -7.551 | -22.036 | 19966 | -1 | 1 |  |  | H2O | 63.7014 | 0.123 | 0.223 | 90 | -0.026 | 0.004 | 29.57 | 844 | 0 | 30.688 | MT1-Rear | 6 | 67 | BSE
46 | P-11571 | 2025/03/14 16:32:16 | MT1-Rear-06 | 2 | -7.643 | -23.21 | 19971 | -1 | 1 |  |  | H2O | 63.7075 | 0.192 | 0.339 | 88 | -0.038 | 0.006 | 29.1 | 844 | 0 | 30.563 | MT1-Rear | 6 | 67 | BSE
47 | P-11571 | 2025/03/14 16:41:04 | MT1-Rear-06 | 3 | -7.636 | -23.819 | 19965 | -1 | 1 |  |  | H2O | 63.7137 | 0.114 | 0.368 | 90 | -0.026 | 0.024 | 29.57 | 844 | 0 | 29.813 | MT1-Rear | 6 | 67 | BSE
48 | P-11571 | 2025/03/14 16:49:52 | MT1-Rear-06 | 4 | -7.689 | -23.668 | 19838 | 0 | 1 |  |  | H2O | 63.7198 | 0.094 | 0.306 | 89 | -0.008 | -0.038 | 29.22 | 844 | 0 | 30.438 | MT1-Rear | 6 | 67 | BSE
49 | P-11572 | 2025/03/14 16:58:40 | MT1-Rear-07 | 1 | -7.558 | -23.929 | 19899 | -1 | 1 |  |  | H2O | 63.7259 | 0.1 | 0.436 | 89 | -0.008 | 0.082 | 29.43 | 844 | 0 | 30.5 | MT1-Rear | 7 | 67 | BSE
50 | P-11572 | 2025/03/14 17:07:29 | MT1-Rear-07 | 2 | -7.565 | -23.985 | 19915 | -1 | 1 |  |  | H2O | 63.7321 | 0.128 | 0.491 | 91 | -0.017 | 0.071 | 30.16 | 844 | 0 | 30.5 | MT1-Rear | 7 | 67 | BSE
51 | P-11572 | 2025/03/14 17:16:17 | MT1-Rear-07 | 3 | -7.658 | -23.639 | 20048 | -1 | 1 |  |  | H2O | 63.7381 | 0.098 | 0.252 | 97 | -0.016 | 0.014 | 29.22 | 844 | 0 | 30.625 | MT1-Rear | 7 | 67 | BSE
52 | P-11572 | 2025/03/14 17:25:05 | MT1-Rear-07 | 4 | -7.66 | -23.926 | 19926 | 0 | 1 |  |  | H2O | 63.7442 | 0.118 | 0.274 | 87 | 0.02 | 0.04 | 28.57 | 844 | 0 | 30.938 | MT1-Rear | 7 | 67 | BSE
53 | P-11573 | 2025/03/14 17:33:54 | MT1-Rear-08 | 1 | -7.706 | -24.015 | 19953 | -1 | 1 |  |  | H2O | 63.7503 | 0.174 | 0.721 | 96 | 0.032 | 0.218 | 31.74 | 844 | 0 | 30.625 | MT1-Rear | 8 | 67 | BSE
54 | P-11573 | 2025/03/14 17:42:43 | MT1-Rear-08 | 2 | -7.631 | -23.581 | 19919 | -1 | 1 |  |  | H2O | 63.7564 | 0.106 | 0.362 | 89 | -0.001 | 0.041 | 29.23 | 844 | 0 | 30.375 | MT1-Rear | 8 | 67 | BSE
55 | P-11573 | 2025/03/14 17:51:31 | MT1-Rear-08 | 3 | -7.628 | -23.795 | 19929 | -1 | 1 |  |  | H2O | 63.7625 | 0.173 | 0.472 | 88 | 0.012 | 0.044 | 29.12 | 844 | 0 | 31 | MT1-Rear | 8 | 67 | BSE
56 | P-11573 | 2025/03/14 18:00:20 | MT1-Rear-08 | 4 | -7.718 | -24.117 | 19937 | 0 | 1 |  |  | H2O | 63.7687 | 0.065 | 0.352 | 87 | -0.008 | 0.035 | 28.7 | 844 | 0 | 30.438 | MT1-Rear | 8 | 67 | BSE
57 | P-11574 | 2025/03/14 18:09:08 | MT1-Rear-09 | 1 | -9.987 | -38.838 | 19871 | -1 | 1 |  |  | H2O | 63.7748 | 0.151 | 0.348 | 89 | -0.005 | 0.024 | 29.26 | 844 | 0 | 30.563 | MT1-Rear | 9 | 67 | BSE
58 | P-11574 | 2025/03/14 18:17:57 | MT1-Rear-09 | 2 | -10.087 | -40.646 | 19960 | -1 | 1 |  |  | H2O | 63.7809 | 0.078 | 0.41 | 87 | 0.006 | 0.047 | 28.78 | 844 | 0 | 30.813 | MT1-Rear | 9 | 67 | BSE
59 | P-11574 | 2025/03/14 18:26:45 | MT1-Rear-09 | 3 | -10.137 | -41.022 | 19953 | -1 | 1 |  |  | H2O | 63.7871 | 0.077 | 0.423 | 89 | -0.003 | 0.074 | 29.5 | 844 | 0 | 30.75 | MT1-Rear | 9 | 67 | BSE
60 | P-11574 | 2025/03/14 18:35:34 | MT1-Rear-09 | 4 | -10.123 | -41.1 | 19935 | 0 | 1 |  |  | H2O | 63.7932 | 0.103 | 0.301 | 88 | -0.006 | 0.039 | 29.02 | 844 | 0 | 30.688 | MT1-Rear | 9 | 67 | BSE
61 | P-11575 | 2025/03/14 18:44:22 | MT1-Rear-10 | 1 | -7.932 | -26.007 | 19938 | -1 | 1 |  |  | H2O | 63.7993 | 0.086 | 0.332 | 89 | -0.014 | 0.029 | 29.39 | 844 | 0 | 30.563 | MT1-Rear | 10 | 67 | BSE
62 | P-11575 | 2025/03/14 18:53:10 | MT1-Rear-10 | 2 | -7.695 | -24.471 | 19915 | -1 | 1 |  |  | H2O | 63.8055 | 0.203 | 0.335 | 89 | 0.035 | 0.061 | 29.32 | 844 | 0 | 30.375 | MT1-Rear | 10 | 67 | BSE
63 | P-11575 | 2025/03/14 19:02:00 | MT1-Rear-10 | 3 | -7.667 | -24.288 | 19922 | -1 | 1 |  |  | H2O | 63.8116 | 0.145 | 0.47 | 91 | 0.013 | 0.082 | 30.09 | 844 | 0 | 30.625 | MT1-Rear | 10 | 67 | BSE
64 | P-11575 | 2025/03/14 19:10:48 | MT1-Rear-10 | 4 | -7.685 | -23.988 | 19890 | 0 | 1 |  |  | H2O | 63.8177 | 0.081 | 0.349 | 92 | -0.015 | -0.011 | 30.23 | 844 | 0 | 30.5 | MT1-Rear | 10 | 67 | BSE
65 | P-11576 | 2025/03/14 19:19:37 | MT1-Rear-11 | 1 | -6.617 | -19.647 | 19961 | -1 | 1 |  |  | H2O | 63.8237 | 0.089 | 0.279 | 81 | -0.012 | 0.034 | 29.45 | 844 | 0 | 30.438 | MT1-Rear | 11 | 67 | BSE
66 | P-11576 | 2025/03/14 19:28:26 | MT1-Rear-11 | 2 | -6.431 | -19.078 | 19925 | -1 | 1 |  |  | H2O | 63.8299 | 0.127 | 0.515 | 96 | 0.034 | 0.155 | 31.76 | 844 | 0 | 30.063 | MT1-Rear | 11 | 67 | BSE
67 | P-11576 | 2025/03/14 19:37:15 | MT1-Rear-11 | 3 | -6.438 | -19.037 | 19805 | -1 | 1 |  |  | H2O | 63.8361 | 0.162 | 0.382 | 90 | -0.009 | 0.058 | 29.66 | 844 | 0 | 30.313 | MT1-Rear | 11 | 67 | BSE
68 | P-11576 | 2025/03/14 19:46:05 | MT1-Rear-11 | 4 | -6.538 | -19.076 | 19941 | 0 | 1 |  |  | H2O | 63.8421 | 0.175 | 0.309 | 82 | -0.04 | 0.041 | 29.77 | 844 | 0 | 30.375 | MT1-Rear | 11 | 67 | BSE
69 | P-11577 | 2025/03/14 19:54:54 | MT1-Rear-12 | 1 | -8.695 | -27.748 | 19819 | -1 | 1 |  |  | H2O | 63.8483 | 0.131 | 0.46 | 90 | -0.009 | 0.121 | 29.57 | 844 | 0 | 30.5 | MT1-Rear | 12 | 67 | BSE
70 | P-11577 | 2025/03/14 20:03:43 | MT1-Rear-12 | 2 | -8.67 | -28.683 | 19803 | -1 | 1 |  |  | H2O | 63.8544 | 0.143 | 0.397 | 89 | -0.01 | -0.04 | 29.29 | 844 | 0 | 31.313 | MT1-Rear | 12 | 67 | BSE
71 | P-11577 | 2025/03/14 20:12:33 | MT1-Rear-12 | 3 | -8.803 | -29.025 | 19902 | -1 | 1 |  |  | H2O | 63.8606 | 0.161 | 0.345 | 86 | -0.013 | -0.021 | 28.37 | 844 | 0 | 30.688 | MT1-Rear | 12 | 67 | BSE
72 | P-11577 | 2025/03/14 20:21:23 | MT1-Rear-12 | 4 | -8.77 | -29.073 | 19880 | 0 | 1 |  |  | H2O | 63.8667 | 0.123 | 0.288 | 89 | -0.019 | 0.043 | 29.45 | 844 | 0 | 30.5 | MT1-Rear | 12 | 67 | BSE
73 | P-11578 | 2025/03/14 20:30:13 | MT1-Rear-13 | 1 | -7.872 | -25.87 | 19925 | -1 | 1 |  |  | H2O | 63.8728 | 0.092 | 0.4 | 89 | 0.003 | 0.094 | 29.42 | 844 | 0 | 30.5 | MT1-Rear | 13 | 67 | BSE
74 | P-11578 | 2025/03/14 20:39:02 | MT1-Rear-13 | 2 | -7.818 | -25.633 | 19923 | -1 | 1 |  |  | H2O | 63.879 | 0.129 | 0.525 | 89 | 0.001 | 0.109 | 29.49 | 844 | 0 | 30.313 | MT1-Rear | 13 | 67 | BSE
75 | P-11578 | 2025/03/14 20:47:51 | MT1-Rear-13 | 3 | -7.918 | -25.576 | 19906 | -1 | 1 |  |  | H2O | 63.8851 | 0.062 | 0.385 | 88 | -0.003 | -0.022 | 29.16 | 844 | 0 | 30.625 | MT1-Rear | 13 | 67 | BSE
76 | P-11578 | 2025/03/14 20:56:41 | MT1-Rear-13 | 4 | -7.815 | -25.774 | 19815 | 0 | 1 |  |  | H2O | 63.8912 | 0.124 | 0.339 | 88 | -0.006 | 0.059 | 28.98 | 844 | 0 | 30.625 | MT1-Rear | 13 | 67 | BSE
77 | P-11579 | 2025/03/14 21:05:30 | MT1-Rear-14 | 1 | -7.578 | -24.492 | 19975 | -1 | 1 |  |  | H2O | 63.8973 | 0.087 | 0.547 | 96 | 0.019 | 0.13 | 31.62 | 844 | 0 | 30.5 | MT1-Rear | 14 | 67 | BSE
78 | P-11579 | 2025/03/14 21:14:19 | MT1-Rear-14 | 2 | -7.463 | -24.32 | 19873 | -1 | 1 |  |  | H2O | 63.9034 | 0.172 | 0.279 | 89 | -0.01 | 0.026 | 29.51 | 844 | 0 | 30.375 | MT1-Rear | 14 | 67 | BSE
79 | P-11579 | 2025/03/14 21:23:08 | MT1-Rear-14 | 3 | -7.541 | -24.359 | 19910 | -1 | 1 |  |  | H2O | 63.9096 | 0.103 | 0.258 | 89 | -0.004 | 0.008 | 29.25 | 844 | 0 | 30.438 | MT1-Rear | 14 | 67 | BSE
80 | P-11579 | 2025/03/14 21:31:58 | MT1-Rear-14 | 4 | -7.552 | -24.509 | 19948 | 0 | 1 |  |  | H2O | 63.9157 | 0.123 | 0.333 | 88 | 0.001 | 0.021 | 29 | 844 | 0 | 30.688 | MT1-Rear | 14 | 67 | BSE
81 | P-11580 | 2025/03/14 21:40:48 | MT1-Rear-15 | 1 | -7.648 | -24.283 | 19973 | -1 | 1 |  |  | H2O | 63.9218 | 0.181 | 0.226 | 89 | -0.002 | 0.009 | 29.31 | 844 | 0 | 30.25 | MT1-Rear | 15 | 67 | BSE
82 | P-11580 | 2025/03/14 21:49:39 | MT1-Rear-15 | 2 | -7.544 | -24.64 | 19888 | -1 | 1 |  |  | H2O | 63.9279 | 0.121 | 0.375 | 89 | 0.02 | 0.084 | 29.23 | 844 | 0 | 30.313 | MT1-Rear | 15 | 67 | BSE
83 | P-11580 | 2025/03/14 21:58:29 | MT1-Rear-15 | 3 | -7.47 | -24.338 | 19858 | -1 | 1 |  |  | H2O | 63.934 | 0.124 | 0.383 | 89 | -0.001 | 0.021 | 29.4 | 844 | 0 | 30.375 | MT1-Rear | 15 | 67 | BSE
84 | P-11580 | 2025/03/14 22:07:18 | MT1-Rear-15 | 4 | -7.423 | -24.918 | 19852 | 0 | 1 |  |  | H2O | 63.9402 | 0.094 | 0.226 | 86 | -0.021 | -0.005 | 28.31 | 844 | 0 | 30.438 | MT1-Rear | 15 | 67 | BSE
85 | P-11581 | 2025/03/14 22:16:08 | MT1-Rear-16 | 1 | -8.643 | -30.102 | 19921 | -1 | 1 |  |  | H2O | 63.9463 | 0.111 | 0.437 | 82 | 0.008 | 0 | 27.08 | 844 | 0 | 30.813 | MT1-Rear | 16 | 67 | BSE
86 | P-11581 | 2025/03/14 22:24:57 | MT1-Rear-16 | 2 | -8.702 | -31.291 | 19916 | -1 | 1 |  |  | H2O | 63.9524 | 0.159 | 0.382 | 85 | -0.036 | 0.026 | 28.1 | 844 | 0 | 30.188 | MT1-Rear | 16 | 67 | BSE
87 | P-11581 | 2025/03/14 22:33:48 | MT1-Rear-16 | 3 | -8.718 | -31.056 | 19993 | -1 | 1 |  |  | H2O | 63.9586 | 0.132 | 0.282 | 98 | -0.012 | 0.024 | 29.6 | 844 | 0 | 30.563 | MT1-Rear | 16 | 67 | BSE
88 | P-11581 | 2025/03/14 22:42:37 | MT1-Rear-16 | 4 | -8.693 | -31.259 | 19967 | 0 | 1 |  |  | H2O | 63.9648 | 0.075 | 0.345 | 92 | -0.006 | 0.078 | 30.22 | 844 | 0 | 30.875 | MT1-Rear | 16 | 67 | BSE
89 | P-11582 | 2025/03/14 22:51:27 | MT1-Rear-17 | 1 | -7.483 | -24.989 | 19928 | -1 | 1 |  |  | H2O | 63.9709 | 0.157 | 0.309 | 86 | -0.026 | -0.061 | 28.32 | 844 | 0 | 30.188 | MT1-Rear | 17 | 67 | BSE
90 | P-11582 | 2025/03/14 23:00:18 | MT1-Rear-17 | 2 | -7.454 | -24.606 | 19913 | -1 | 1 |  |  | H2O | 63.977 | 0.101 | 0.471 | 87 | 0.003 | -0.041 | 28.64 | 844 | 0 | 30.313 | MT1-Rear | 17 | 67 | BSE
91 | P-11582 | 2025/03/14 23:09:07 | MT1-Rear-17 | 3 | -7.439 | -24.546 | 19884 | -1 | 1 |  |  | H2O | 63.9832 | 0.125 | 0.219 | 88 | -0.012 | -0.026 | 28.97 | 844 | 0 | 30.375 | MT1-Rear | 17 | 67 | BSE
92 | P-11582 | 2025/03/14 23:17:57 | MT1-Rear-17 | 4 | -7.54 | -24.428 | 19837 | 0 | 1 |  |  | H2O | 63.9893 | 0.178 | 0.383 | 86 | 0 | -0.079 | 28.31 | 844 | 0 | 30.438 | MT1-Rear | 17 | 67 | BSE
93 | P-11583 | 2025/03/14 23:26:47 | MT1-Rear-18 | 1 | -9.758 | -39.313 | 19945 | -1 | 1 |  |  | H2O | 63.9954 | 0.093 | 0.316 | 89 | -0.016 | 0.006 | 29.33 | 844 | 0 | 29.938 | MT1-Rear | 18 | 67 | BSE
94 | P-11583 | 2025/03/14 23:35:36 | MT1-Rear-18 | 2 | -9.947 | -40.964 | 19893 | -1 | 1 |  |  | H2O | 64.0015 | 0.134 | 0.379 | 86 | -0.003 | 0.098 | 28.29 | 844 | 0 | 29.438 | MT1-Rear | 18 | 67 | BSE
95 | P-11583 | 2025/03/14 23:44:27 | MT1-Rear-18 | 3 | -9.907 | -41.287 | 19945 | -1 | 1 |  |  | H2O | 64.0078 | 0.12 | 0.222 | 90 | 0.004 | 0.032 | 29.83 | 844 | 0 | 30.25 | MT1-Rear | 18 | 67 | BSE
96 | P-11583 | 2025/03/14 23:53:18 | MT1-Rear-18 | 4 | -9.908 | -41.448 | 19943 | 0 | 1 |  |  | H2O | 64.0139 | 0.152 | 0.402 | 94 | 0.019 | 0.101 | 30.95 | 844 | 0 | 30.25 | MT1-Rear | 18 | 67 | BSE
97 | P-11584 | 2025/03/15 00:02:08 | MT1-Rear-19 | 1 | -9.046 | -32.733 | 19918 | -1 | 1 |  |  | H2O | 64.02 | 0.087 | 0.381 | 86 | -0.006 | -0.053 | 28.32 | 844 | 0 | 29.875 | MT1-Rear | 19 | 67 | BSE
98 | P-11584 | 2025/03/15 00:10:58 | MT1-Rear-19 | 2 | -9.066 | -31.352 | 19907 | -1 | 1 |  |  | H2O | 64.0261 | 0.085 | 0.233 | 87 | -0.004 | 0.01 | 28.82 | 844 | 0 | 30.188 | MT1-Rear | 19 | 67 | BSE
99 | P-11584 | 2025/03/15 00:19:50 | MT1-Rear-19 | 3 | -8.973 | -31.484 | 19904 | -1 | 1 |  |  | H2O | 64.0322 | 0.044 | 0.351 | 89 | 0.004 | 0.001 | 29.28 | 844 | 0 | 30.375 | MT1-Rear | 19 | 67 | BSE
100 | P-11584 | 2025/03/15 00:28:41 | MT1-Rear-19 | 4 | -9.048 | -31.193 | 19957 | 0 | 1 |  |  | H2O | 64.0384 | 0.093 | 0.241 | 92 | -0.017 | -0.05 | 30.28 | 844 | 0 | 30.25 | MT1-Rear | 19 | 67 | BSE
101 | P-11585 | 2025/03/15 00:37:31 | MT1-Rear-20 | 1 | -7.547 | -25.034 | 19953 | -1 | 1 |  |  | H2O | 64.0446 | 0.135 | 0.295 | 88 | -0.034 | 0.04 | 29.13 | 844 | 0 | 30.75 | MT1-Rear | 20 | 67 | BSE
102 | P-11585 | 2025/03/15 00:46:22 | MT1-Rear-20 | 2 | -7.525 | -24.712 | 19922 | -1 | 1 |  |  | H2O | 64.0507 | 0.109 | 0.361 | 90 | 0 | 0.054 | 29.8 | 844 | 0 | 30.75 | MT1-Rear | 20 | 67 | BSE
103 | P-11585 | 2025/03/15 00:55:14 | MT1-Rear-20 | 3 | -7.456 | -24.384 | 19930 | -1 | 1 |  |  | H2O | 64.0568 | 0.141 | 0.3 | 89 | -0.022 | -0.02 | 29.3 | 844 | 0 | 30.063 | MT1-Rear | 20 | 67 | BSE
104 | P-11585 | 2025/03/15 01:04:05 | MT1-Rear-20 | 4 | -7.496 | -24.735 | 19916 | 0 | 1 |  |  | H2O | 64.063 | 0.151 | 0.654 | 98 | 0.028 | 0.16 | 32.27 | 844 | 0 | 30.313 | MT1-Rear | 20 | 67 | BSE
105 | P-11586 | 2025/03/15 01:12:56 | MT1-Rear-21 | 1 | -9.005 | -31.431 | 19913 | -1 | 1 |  |  | H2O | 64.0691 | 0.111 | 0.353 | 88 | -0.02 | 0.009 | 29.02 | 844 | 0 | 30.438 | MT1-Rear | 21 | 67 | BSE
106 | P-11586 | 2025/03/15 01:21:48 | MT1-Rear-21 | 2 | -9.212 | -32.092 | 19936 | -1 | 1 |  |  | H2O | 64.0752 | 0.157 | 0.173 | 87 | -0.007 | 0.005 | 28.55 | 844 | 0 | 30.063 | MT1-Rear | 21 | 67 | BSE
107 | P-11586 | 2025/03/15 01:30:38 | MT1-Rear-21 | 3 | -9.208 | -32.556 | 19958 | -1 | 1 |  |  | H2O | 64.0815 | 0.167 | 0.395 | 86 | -0.006 | 0.026 | 28.52 | 844 | 0 | 29.813 | MT1-Rear | 21 | 67 | BSE
108 | P-11586 | 2025/03/15 01:39:29 | MT1-Rear-21 | 4 | -9.204 | -32.682 | 19921 | 0 | 1 |  |  | H2O | 64.0876 | 0.125 | 0.422 | 85 | -0.02 | 0.052 | 27.97 | 844 | 0 | 30.25 | MT1-Rear | 21 | 67 | BSE
109 | P-11587 | 2025/03/15 01:48:22 | MT1-Rear-22 | 1 | -8.966 | -31.485 | 19956 | -1 | 1 |  |  | H2O | 64.0937 | 0.11 | 0.316 | 87 | 0.003 | 0.007 | 28.62 | 844 | 0 | 30.625 | MT1-Rear | 22 | 67 | BSE
110 | P-11587 | 2025/03/15 01:57:13 | MT1-Rear-22 | 2 | -9.062 | -31.048 | 19960 | -1 | 1 |  |  | H2O | 64.0999 | 0.122 | 0.354 | 98 | -0.018 | 0.047 | 29.5 | 844 | 0 | 31.125 | MT1-Rear | 22 | 67 | BSE
111 | P-11587 | 2025/03/15 02:06:05 | MT1-Rear-22 | 3 | -9.012 | -31.091 | 19944 | -1 | 1 |  |  | H2O | 64.1061 | 0.106 | 0.484 | 92 | 0.004 | 0.089 | 30.35 | 844 | 0 | 30.313 | MT1-Rear | 22 | 67 | BSE
112 | P-11587 | 2025/03/15 02:14:56 | MT1-Rear-22 | 4 | -8.965 | -31.591 | 19938 | 0 | 1 |  |  | H2O | 64.1122 | 0.142 | 0.6 | 100 | 0.037 | 0.187 | 32.89 | 844 | 0 | 29.938 | MT1-Rear | 22 | 67 | BSE
113 | P-11588 | 2025/03/15 02:23:48 | MT1-Rear-23 | 1 | -8.694 | -30.28 | 19932 | -1 | 1 |  |  | H2O | 64.1184 | 0.14 | 0.332 | 88 | -0.024 | 0.045 | 29.05 | 844 | 0 | 30.438 | MT1-Rear | 23 | 67 | BSE
114 | P-11588 | 2025/03/15 02:32:40 | MT1-Rear-23 | 2 | -8.757 | -30.137 | 19944 | -1 | 1 |  |  | H2O | 64.1245 | 0.103 | 0.239 | 89 | 0.001 | -0.04 | 29.2 | 844 | 0 | 30.438 | MT1-Rear | 23 | 67 | BSE
115 | P-11588 | 2025/03/15 02:41:31 | MT1-Rear-23 | 3 | -8.721 | -30.335 | 19929 | -1 | 1 |  |  | H2O | 64.1306 | 0.103 | 0.421 | 89 | -0.02 | 0.043 | 29.27 | 844 | 0 | 29.688 | MT1-Rear | 23 | 67 | BSE
116 | P-11588 | 2025/03/15 02:50:21 | MT1-Rear-23 | 4 | -8.794 | -30.191 | 19965 | 0 | 1 |  |  | H2O | 64.1368 | 0.121 | 0.179 | 91 | 0.022 | 0.033 | 29.93 | 844 | 0 | 30.375 | MT1-Rear | 23 | 67 | BSE
117 | P-11589 | 2025/03/15 02:59:13 | MT1-Rear-24 | 1 | -7.458 | -25.066 | 19963 | -1 | 1 |  |  | H2O | 64.1429 | 0.096 | 0.188 | 91 | -0.021 | 0.03 | 29.98 | 844 | 0 | 30.313 | MT1-Rear | 24 | 67 | BSE
118 | P-11589 | 2025/03/15 03:08:05 | MT1-Rear-24 | 2 | -7.535 | -24.528 | 19943 | -1 | 1 |  |  | H2O | 64.149 | 0.107 | 0.464 | 93 | 0.004 | 0.094 | 30.62 | 844 | 0 | 30 | MT1-Rear | 24 | 67 | BSE
119 | P-11589 | 2025/03/15 03:16:56 | MT1-Rear-24 | 3 | -7.37 | -23.994 | 19927 | -1 | 1 |  |  | H2O | 64.1553 | 0.116 | 0.225 | 90 | 0.026 | 0.018 | 29.85 | 844 | 0 | 30.188 | MT1-Rear | 24 | 67 | BSE
120 | P-11589 | 2025/03/15 03:25:47 | MT1-Rear-24 | 4 | -7.402 | -24.411 | 19924 | 0 | 1 |  |  | H2O | 64.1614 | 0.098 | 0.232 | 87 | -0.002 | -0.006 | 28.79 | 844 | 0 | 30.938 | MT1-Rear | 24 | 67 | BSE
121 | P-11590 | 2025/03/15 03:34:40 | MT1-Rear-25 | 1 | -8.564 | -30.524 | 19946 | -1 | 1 |  |  | H2O | 64.1675 | 0.134 | 0.194 | 88 | -0.015 | -0.03 | 29.03 | 844 | 0 | 30.375 | MT1-Rear | 25 | 67 | BSE
122 | P-11590 | 2025/03/15 03:43:32 | MT1-Rear-25 | 2 | -8.53 | -30.945 | 19922 | -1 | 1 |  |  | H2O | 64.1737 | 0.125 | 0.392 | 90 | -0.023 | 0.047 | 29.75 | 844 | 0 | 30.438 | MT1-Rear | 25 | 67 | BSE
123 | P-11590 | 2025/03/15 03:52:24 | MT1-Rear-25 | 3 | -8.612 | -31.158 | 19763 | -1 | 1 |  |  | H2O | 64.1798 | 0.112 | 0.28 | 83 | -0.02 | 0.012 | 27.47 | 844 | 0 | 31.063 | MT1-Rear | 25 | 67 | BSE
124 | P-11590 | 2025/03/15 04:01:15 | MT1-Rear-25 | 4 | -8.645 | -31.01 | 19969 | 0 | 1 |  |  | H2O | 64.1861 | 0.088 | 0.408 | 90 | 0.007 | 0.007 | 29.67 | 844 | 0 | 30.125 | MT1-Rear | 25 | 67 | BSE
125 | P-11591 | 2025/03/15 04:10:06 | MT1-Rear-26 | 1 | -6.035 | -8.828 | 19944 | -1 | 1 |  |  | H2O | 64.1922 | 0.17 | 0.192 | 85 | -0.006 | 0.037 | 28.06 | 844 | 0 | 30.25 | MT1-Rear | 26 | 67 | BSE
126 | P-11591 | 2025/03/15 04:18:58 | MT1-Rear-26 | 2 | -5.797 | -6.491 | 19951 | -1 | 1 |  |  | H2O | 64.1983 | 0.13 | 0.265 | 95 | -0.022 | -0.028 | 28.51 | 844 | 0 | 30.625 | MT1-Rear | 26 | 67 | BSE
127 | P-11591 | 2025/03/15 04:27:50 | MT1-Rear-26 | 3 | -5.921 | -5.816 | 19921 | -1 | 1 |  |  | H2O | 64.2045 | 0.088 | 0.286 | 89 | -0.004 | 0.06 | 29.28 | 844 | 0 | 30.25 | MT1-Rear | 26 | 67 | BSE
128 | P-11591 | 2025/03/15 04:36:42 | MT1-Rear-26 | 4 | -5.708 | -5.981 | 19894 | 0 | 1 |  |  | H2O | 64.2107 | 0.115 | 0.412 | 96 | -0.009 | -0.019 | 28.98 | 844 | 0 | 30.438 | MT1-Rear | 26 | 67 | BSE
129 | P-11592 | 2025/03/15 04:45:33 | MT1-Rear-27 | 1 | -9.51 | -37.189 | 19953 | -1 | 1 |  |  | H2O | 64.2168 | 0.141 | 0.217 | 86 | -0.012 | -0.04 | 28.47 | 844 | 0 | 30.125 | MT1-Rear | 27 | 67 | BSE
130 | P-11592 | 2025/03/15 04:54:25 | MT1-Rear-27 | 2 | -9.783 | -40.283 | 19941 | -1 | 1 |  |  | H2O | 64.2229 | 0.111 | 0.372 | 88 | -0.02 | 0.072 | 29.1 | 844 | 0 | 30.25 | MT1-Rear | 27 | 67 | BSE
131 | P-11592 | 2025/03/15 05:03:17 | MT1-Rear-27 | 3 | -9.836 | -41.103 | 19831 | -1 | 1 |  |  | H2O | 64.2292 | 0.159 | 0.379 | 88 | -0.025 | 0.101 | 28.88 | 844 | 0 | 30.813 | MT1-Rear | 27 | 67 | BSE
132 | P-11592 | 2025/03/15 05:12:10 | MT1-Rear-27 | 4 | -9.82 | -41.279 | 19923 | 0 | 1 |  |  | H2O | 64.2352 | 0.099 | 0.473 | 90 | 0.002 | 0.131 | 29.76 | 844 | 0 | 29.5 | MT1-Rear | 27 | 67 | BSE

## Run Summary
Line | Sample | Inj Nr | Sample Name | H2O_Mean | Good | Ignore | Time Code | d(18_16)Mean | d(D_H)Mean
1 | 1 | 1 | Blacksburg | 21368 | 1 | 0 | 2025/03/14 09:56:57 | -10.405 | -46.647 |  | Min | Max
2 | 1 | 2 | Blacksburg | 20722 | 1 | 0 | 2025/03/14 10:05:43 | -10.211 | -43.011 | 0 | 17000 | 23000
3 | 1 | 3 | Blacksburg | 19833 | 1 | 0 | 2025/03/14 10:14:30 | -10.116 | -42.545 | 500 | 17000 | 23000
4 | 1 | 4 | Blacksburg | 19730 | 1 | 0 | 2025/03/14 10:23:16 | -10.119 | -42.237
5 | 1 | 5 | Blacksburg | 19754 | 1 | 0 | 2025/03/14 10:32:03 | -10.105 | -41.934
6 | 1 | 6 | Blacksburg | 19729 | 1 | 0 | 2025/03/14 10:40:49 | -10.045 | -42.269
7 | 1 | 7 | Blacksburg | 19739 | 1 | 0 | 2025/03/14 10:49:35 | -9.944 | -42.343
8 | 1 | 8 | Blacksburg | 19680 | 1 | 0 | 2025/03/14 10:58:22 | -10.057 | -42.167
9 | 1 | 9 | Blacksburg | 19576 | 1 | 0 | 2025/03/14 11:07:08 | -10.057 | -42.147
10 | 1 | 10 | Blacksburg | 19642 | 1 | 0 | 2025/03/14 11:15:55 | -9.999 | -42.018
11 | 2 | 1 | Myrtle | 19914 | 1 | 0 | 2025/03/14 11:24:42 | -5.096 | -8.127
12 | 2 | 2 | Myrtle | 19586 | 1 | 0 | 2025/03/14 11:33:28 | -4.803 | -4.225
13 | 2 | 3 | Myrtle | 19629 | 1 | 0 | 2025/03/14 11:42:15 | -4.673 | -3.581
14 | 2 | 4 | Myrtle | 19726 | 1 | 0 | 2025/03/14 11:51:02 | -4.676 | -3.305
15 | 2 | 5 | Myrtle | 19941 | 1 | 0 | 2025/03/14 11:59:50 | -4.796 | -2.96
16 | 2 | 6 | Myrtle | 19713 | 1 | 0 | 2025/03/14 12:08:36 | -4.642 | -3.021
17 | 2 | 7 | Myrtle | 19690 | 1 | 0 | 2025/03/14 12:17:23 | -4.651 | -2.744 | Concentration of water vapor should be about 20,000 ppmV (roughly +/- 1000; but not outside 17,000 and 23,000 ppmV).
18 | 2 | 8 | Myrtle | 19541 | 1 | 0 | 2025/03/14 12:26:10 | -4.644 | -2.876 | Large scatter may indicate a problem with the syringe or the septa.
19 | 2 | 9 | Myrtle | 19664 | 1 | 0 | 2025/03/14 12:34:57 | -4.611 | -2.793
20 | 2 | 10 | Myrtle | 19627 | 1 | 0 | 2025/03/14 12:43:43 | -4.572 | -2.983
21 | 3 | 1 | Homer | 19941 | 1 | 0 | 2025/03/14 12:52:30 | -16.374 | -93.746
22 | 3 | 2 | Homer | 19770 | 1 | 0 | 2025/03/14 13:01:16 | -17.045 | -103.773
23 | 3 | 3 | Homer | 19788 | 1 | 0 | 2025/03/14 13:10:03 | -17.283 | -105.376
24 | 3 | 4 | Homer | 19952 | 1 | 0 | 2025/03/14 13:18:50 | -17.398 | -106.349
25 | 3 | 5 | Homer | 19931 | 1 | 0 | 2025/03/14 13:27:35 | -17.308 | -106.773
26 | 3 | 6 | Homer | 19835 | 1 | 0 | 2025/03/14 13:36:22 | -17.227 | -107.201
27 | 3 | 7 | Homer | 19815 | 1 | 0 | 2025/03/14 13:45:09 | -17.284 | -107.375
28 | 3 | 8 | Homer | 20367 | 1 | 0 | 2025/03/14 13:53:57 | -17.33 | -107.753
29 | 3 | 9 | Homer | 19880 | 1 | 0 | 2025/03/14 14:02:45 | -17.325 | -107.663
30 | 3 | 10 | Homer | 19705 | 1 | 0 | 2025/03/14 14:11:32 | -17.342 | -107.951
31 | 4 | 1 | Blacksburg | 19954 | 1 | 0 | 2025/03/14 14:20:19 | -10.498 | -50.908
32 | 4 | 2 | Blacksburg | 19963 | 1 | 0 | 2025/03/14 14:29:07 | -10.178 | -44.518
33 | 4 | 3 | Blacksburg | 20272 | 1 | 0 | 2025/03/14 14:37:54 | -10.06 | -43.475
34 | 4 | 4 | Blacksburg | 20632 | 1 | 0 | 2025/03/14 14:46:42 | -10.117 | -42.532
35 | 4 | 5 | Blacksburg | 19981 | 1 | 0 | 2025/03/14 14:55:29 | -10.071 | -42.567
36 | 4 | 6 | Blacksburg | 19553 | 1 | 0 | 2025/03/14 15:04:17 | -10.048 | -42.225
37 | 4 | 7 | Blacksburg | 19565 | 1 | 0 | 2025/03/14 15:13:04 | -10.009 | -42.175
38 | 4 | 8 | Blacksburg | 19924 | 1 | 0 | 2025/03/14 15:21:52 | -10.004 | -42.105
39 | 4 | 9 | Blacksburg | 19566 | 1 | 0 | 2025/03/14 15:30:40 | -10.096 | -42.021
40 | 4 | 10 | Blacksburg | 19989 | 1 | 0 | 2025/03/14 15:39:28 | -9.944 | -42.073
41 | 5 | 1 | Hawaii | 19972 | 1 | 0 | 2025/03/14 15:48:15 | -6.27 | -10.476
42 | 5 | 2 | Hawaii | 19944 | 1 | 0 | 2025/03/14 15:57:02 | -6.124 | -6.728
43 | 5 | 3 | Hawaii | 19959 | 1 | 0 | 2025/03/14 16:05:50 | -6.131 | -5.763
44 | 5 | 4 | Hawaii | 19939 | 1 | 0 | 2025/03/14 16:14:39 | -6.032 | -5.573
45 | 6 | 1 | CC4 17feb25 BOT | 19966 | 1 | 0 | 2025/03/14 16:23:27 | -7.551 | -22.036
46 | 6 | 2 | CC4 17feb25 BOT | 19971 | 1 | 0 | 2025/03/14 16:32:16 | -7.643 | -23.21
47 | 6 | 3 | CC4 17feb25 BOT | 19965 | 1 | 0 | 2025/03/14 16:41:04 | -7.636 | -23.819
48 | 6 | 4 | CC4 17feb25 BOT | 19838 | 1 | 0 | 2025/03/14 16:49:52 | -7.689 | -23.668
49 | 7 | 1 | C50 17feb25 20m | 19899 | 1 | 0 | 2025/03/14 16:58:40 | -7.558 | -23.929
50 | 7 | 2 | C50 17feb25 20m | 19915 | 1 | 0 | 2025/03/14 17:07:29 | -7.565 | -23.985
51 | 7 | 3 | C50 17feb25 20m | 20048 | 1 | 0 | 2025/03/14 17:16:17 | -7.658 | -23.639
52 | 7 | 4 | C50 17feb25 20m | 19926 | 1 | 0 | 2025/03/14 17:25:05 | -7.66 | -23.926
53 | 8 | 1 | C50 17feb25 6m | 19953 | 1 | 0 | 2025/03/14 17:33:54 | -7.706 | -24.015
54 | 8 | 2 | C50 17feb25 6m | 19919 | 1 | 0 | 2025/03/14 17:42:43 | -7.631 | -23.581
55 | 8 | 3 | C50 17feb25 6m | 19929 | 1 | 0 | 2025/03/14 17:51:31 | -7.628 | -23.795
56 | 8 | 4 | C50 17feb25 6m | 19937 | 1 | 0 | 2025/03/14 18:00:20 | -7.718 | -24.117
57 | 9 | 1 | Blacksburg | 19871 | 1 | 0 | 2025/03/14 18:09:08 | -9.987 | -38.838
58 | 9 | 2 | Blacksburg | 19960 | 1 | 0 | 2025/03/14 18:17:57 | -10.087 | -40.646
59 | 9 | 3 | Blacksburg | 19953 | 1 | 0 | 2025/03/14 18:26:45 | -10.137 | -41.022
60 | 9 | 4 | Blacksburg | 19935 | 1 | 0 | 2025/03/14 18:35:34 | -10.123 | -41.1
61 | 10 | 1 | CC4 17feb25 1.5m | 19938 | 1 | 0 | 2025/03/14 18:44:22 | -7.932 | -26.007
62 | 10 | 2 | CC4 17feb25 1.5m | 19915 | 1 | 0 | 2025/03/14 18:53:10 | -7.695 | -24.471
63 | 10 | 3 | CC4 17feb25 1.5m | 19922 | 1 | 0 | 2025/03/14 19:02:00 | -7.667 | -24.288
64 | 10 | 4 | CC4 17feb25 1.5m | 19890 | 1 | 0 | 2025/03/14 19:10:48 | -7.685 | -23.988
65 | 11 | 1 | CC3 15aug24 0.1m | 19961 | 1 | 0 | 2025/03/14 19:19:37 | -6.617 | -19.647
66 | 11 | 2 | CC3 15aug24 0.1m | 19925 | 1 | 0 | 2025/03/14 19:28:26 | -6.431 | -19.078
67 | 11 | 3 | CC3 15aug24 0.1m | 19805 | 1 | 0 | 2025/03/14 19:37:15 | -6.438 | -19.037
68 | 11 | 4 | CC3 15aug24 0.1m | 19941 | 1 | 0 | 2025/03/14 19:46:05 | -6.538 | -19.076
69 | 12 | 1 | CP1 17feb25 BOT | 19819 | 1 | 0 | 2025/03/14 19:54:54 | -8.695 | -27.748
70 | 12 | 2 | CP1 17feb25 BOT | 19803 | 1 | 0 | 2025/03/14 20:03:43 | -8.67 | -28.683
71 | 12 | 3 | CP1 17feb25 BOT | 19902 | 1 | 0 | 2025/03/14 20:12:33 | -8.803 | -29.025
72 | 12 | 4 | CP1 17feb25 BOT | 19880 | 1 | 0 | 2025/03/14 20:21:23 | -8.77 | -29.073
73 | 13 | 1 | CC5 22may24 13m | 19925 | 1 | 0 | 2025/03/14 20:30:13 | -7.872 | -25.87
74 | 13 | 2 | CC5 22may24 13m | 19923 | 1 | 0 | 2025/03/14 20:39:02 | -7.818 | -25.633
75 | 13 | 3 | CC5 22may24 13m | 19906 | 1 | 0 | 2025/03/14 20:47:51 | -7.918 | -25.576
76 | 13 | 4 | CC5 22may24 13m | 19815 | 1 | 0 | 2025/03/14 20:56:41 | -7.815 | -25.774
77 | 14 | 1 | CC2 22may24 BOT | 19975 | 1 | 0 | 2025/03/14 21:05:30 | -7.578 | -24.492
78 | 14 | 2 | CC2 22may24 BOT | 19873 | 1 | 0 | 2025/03/14 21:14:19 | -7.463 | -24.32
79 | 14 | 3 | CC2 22may24 BOT | 19910 | 1 | 0 | 2025/03/14 21:23:08 | -7.541 | -24.359
80 | 14 | 4 | CC2 22may24 BOT | 19948 | 1 | 0 | 2025/03/14 21:31:58 | -7.552 | -24.509
81 | 15 | 1 | CP2 17feb25 BOT | 19973 | 1 | 0 | 2025/03/14 21:40:48 | -7.648 | -24.283
82 | 15 | 2 | CP2 17feb25 BOT | 19888 | 1 | 0 | 2025/03/14 21:49:39 | -7.544 | -24.64
83 | 15 | 3 | CP2 17feb25 BOT | 19858 | 1 | 0 | 2025/03/14 21:58:29 | -7.47 | -24.338
84 | 15 | 4 | CP2 17feb25 BOT | 19852 | 1 | 0 | 2025/03/14 22:07:18 | -7.423 | -24.918
85 | 16 | 1 | CS1 15aug24 0.1m | 19921 | 1 | 0 | 2025/03/14 22:16:08 | -8.643 | -30.102
86 | 16 | 2 | CS1 15aug24 0.1m | 19916 | 1 | 0 | 2025/03/14 22:24:57 | -8.702 | -31.291
87 | 16 | 3 | CS1 15aug24 0.1m | 19993 | 1 | 0 | 2025/03/14 22:33:48 | -8.718 | -31.056
88 | 16 | 4 | CS1 15aug24 0.1m | 19967 | 1 | 0 | 2025/03/14 22:42:37 | -8.693 | -31.259
89 | 17 | 1 | CC2 17feb25 BOT | 19928 | 1 | 0 | 2025/03/14 22:51:27 | -7.483 | -24.989
90 | 17 | 2 | CC2 17feb25 BOT | 19913 | 1 | 0 | 2025/03/14 23:00:18 | -7.454 | -24.606
91 | 17 | 3 | CC2 17feb25 BOT | 19884 | 1 | 0 | 2025/03/14 23:09:07 | -7.439 | -24.546
92 | 17 | 4 | CC2 17feb25 BOT | 19837 | 1 | 0 | 2025/03/14 23:17:57 | -7.54 | -24.428
93 | 18 | 1 | Blacksburg | 19945 | 1 | 0 | 2025/03/14 23:26:47 | -9.758 | -39.313
94 | 18 | 2 | Blacksburg | 19893 | 1 | 0 | 2025/03/14 23:35:36 | -9.947 | -40.964
95 | 18 | 3 | Blacksburg | 19945 | 1 | 0 | 2025/03/14 23:44:27 | -9.907 | -41.287
96 | 18 | 4 | Blacksburg | 19943 | 1 | 0 | 2025/03/14 23:53:18 | -9.908 | -41.448
97 | 19 | 1 | CS1 17feb25 0.1m | 19918 | 1 | 0 | 2025/03/15 00:02:08 | -9.046 | -32.733
98 | 19 | 2 | CS1 17feb25 0.1m | 19907 | 1 | 0 | 2025/03/15 00:10:58 | -9.066 | -31.352
99 | 19 | 3 | CS1 17feb25 0.1m | 19904 | 1 | 0 | 2025/03/15 00:19:50 | -8.973 | -31.484
100 | 19 | 4 | CS1 17feb25 0.1m | 19957 | 1 | 0 | 2025/03/15 00:28:41 | -9.048 | -31.193
101 | 20 | 1 | CC2 17feb25 0.1m | 19953 | 1 | 0 | 2025/03/15 00:37:31 | -7.547 | -25.034
102 | 20 | 2 | CC2 17feb25 0.1m | 19922 | 1 | 0 | 2025/03/15 00:46:22 | -7.525 | -24.712
103 | 20 | 3 | CC2 17feb25 0.1m | 19930 | 1 | 0 | 2025/03/15 00:55:14 | -7.456 | -24.384
104 | 20 | 4 | CC2 17feb25 0.1m | 19916 | 1 | 0 | 2025/03/15 01:04:05 | -7.496 | -24.735
105 | 21 | 1 | CS1 22may24 0.1m | 19913 | 1 | 0 | 2025/03/15 01:12:56 | -9.005 | -31.431
106 | 21 | 2 | CS1 22may24 0.1m | 19936 | 1 | 0 | 2025/03/15 01:21:48 | -9.212 | -32.092
107 | 21 | 3 | CS1 22may24 0.1m | 19958 | 1 | 0 | 2025/03/15 01:30:38 | -9.208 | -32.556
108 | 21 | 4 | CS1 22may24 0.1m | 19921 | 1 | 0 | 2025/03/15 01:39:29 | -9.204 | -32.682
109 | 22 | 1 | CP1 17feb25 0.1m | 19956 | 1 | 0 | 2025/03/15 01:48:22 | -8.966 | -31.485
110 | 22 | 2 | CP1 17feb25 0.1m | 19960 | 1 | 0 | 2025/03/15 01:57:13 | -9.062 | -31.048
111 | 22 | 3 | CP1 17feb25 0.1m | 19944 | 1 | 0 | 2025/03/15 02:06:05 | -9.012 | -31.091
112 | 22 | 4 | CP1 17feb25 0.1m | 19938 | 1 | 0 | 2025/03/15 02:14:56 | -8.965 | -31.591
113 | 23 | 1 | CP2 17feb25 0.1m | 19932 | 1 | 0 | 2025/03/15 02:23:48 | -8.694 | -30.28
114 | 23 | 2 | CP2 17feb25 0.1m | 19944 | 1 | 0 | 2025/03/15 02:32:40 | -8.757 | -30.137
115 | 23 | 3 | CP2 17feb25 0.1m | 19929 | 1 | 0 | 2025/03/15 02:41:31 | -8.721 | -30.335
116 | 23 | 4 | CP2 17feb25 0.1m | 19965 | 1 | 0 | 2025/03/15 02:50:21 | -8.794 | -30.191
117 | 24 | 1 | CC3 17feb25 BOT | 19963 | 1 | 0 | 2025/03/15 02:59:13 | -7.458 | -25.066
118 | 24 | 2 | CC3 17feb25 BOT | 19943 | 1 | 0 | 2025/03/15 03:08:05 | -7.535 | -24.528
119 | 24 | 3 | CC3 17feb25 BOT | 19927 | 1 | 0 | 2025/03/15 03:16:56 | -7.37 | -23.994
120 | 24 | 4 | CC3 17feb25 BOT | 19924 | 1 | 0 | 2025/03/15 03:25:47 | -7.402 | -24.411
121 | 25 | 1 | CS1 15aug24 0.1m - DUP | 19946 | 1 | 0 | 2025/03/15 03:34:40 | -8.564 | -30.524
122 | 25 | 2 | CS1 15aug24 0.1m - DUP | 19922 | 1 | 0 | 2025/03/15 03:43:32 | -8.53 | -30.945
123 | 25 | 3 | CS1 15aug24 0.1m - DUP | 19763 | 1 | 0 | 2025/03/15 03:52:24 | -8.612 | -31.158
124 | 25 | 4 | CS1 15aug24 0.1m - DUP | 19969 | 1 | 0 | 2025/03/15 04:01:15 | -8.645 | -31.01
125 | 26 | 1 | Hawaii | 19944 | 1 | 0 | 2025/03/15 04:10:06 | -6.035 | -8.828
126 | 26 | 2 | Hawaii | 19951 | 1 | 0 | 2025/03/15 04:18:58 | -5.797 | -6.491
127 | 26 | 3 | Hawaii | 19921 | 1 | 0 | 2025/03/15 04:27:50 | -5.921 | -5.816
128 | 26 | 4 | Hawaii | 19894 | 1 | 0 | 2025/03/15 04:36:42 | -5.708 | -5.981
129 | 27 | 1 | Blacksburg | 19953 | 1 | 0 | 2025/03/15 04:45:33 | -9.51 | -37.189
130 | 27 | 2 | Blacksburg | 19941 | 1 | 0 | 2025/03/15 04:54:25 | -9.783 | -40.283
131 | 27 | 3 | Blacksburg | 19831 | 1 | 0 | 2025/03/15 05:03:17 | -9.836 | -41.103
132 | 27 | 4 | Blacksburg | 19923 | 1 | 0 | 2025/03/15 05:12:10 | -9.82 | -41.279

## 18O
Line | Sample | Injection | Sample Name | Correction Type | Good | Ignore | d(18_16) Picarro Raw | Stdev Raw | d(18_16) Memory Corrected | Stdev Memory Corrected | d(18_16) Memory Corrected Drifts | d(18_16) Drift Corrected | Drift Corrected Drifts | d(18_16) Drift Corrected High, Low, and Drift | d(18_16) Reference Defined | d18O Final Corrected Value | Post run corrections:
1 | 1 | 1 | Blacksburg | Drift | 1 | -1 | -10.405 | [9 empty] | (1) Memory correction |  |  | (1) Graphs for memory correction
2 | 1 | 2 | Blacksburg | Drift | 1 | -1 | -10.211 | [9 empty] | Combined Stdev | 0.084255
3 | 1 | 3 | Blacksburg | Drift | 1 | -1 | -10.116 | [9 empty] | Date:
4 | 1 | 4 | Blacksburg | Drift | 1 | -1 | -10.119 | [9 empty] | injection | 18O memory coef
5 | 1 | 5 | Blacksburg | Drift | 1 | -1 | -10.105 | [9 empty] | 1 | 0.91561
6 | 1 | 6 | Blacksburg | Drift | 1 | -1 | -10.045 | [9 empty] | 2 | 0.97208
7 | 1 | 7 | Blacksburg | Drift | 1 | 0 | -9.944 |  | -9.944 |  | -9.944 | -9.9544 | -9.9544 |  |  | -7.4081 | 3 | 0.99144
8 | 1 | 8 | Blacksburg | Drift | 1 | 0 | -10.057 |  | -10.057 |  | -10.057 | -10.0689 | -10.0689 |  |  | -7.5178 | 4 | 0.99171
9 | 1 | 9 | Blacksburg | Drift | 1 | 0 | -10.057 |  | -10.057 |  | -10.057 | -10.0703 | -10.0703 |  |  | -7.5193 | 5 | 0.99171
10 | 1 | 10 | Blacksburg | Drift | 1 | 0 | -9.999 | 0.05423 | -9.999 | 0.05423 | -9.999 | -10.0138 | -10.0138 |  |  | -7.4651 | 6 | 0.99171
11 | 2 | 1 | Myrtle | High Reference | 1 | 0 | -5.096 |  | -4.6823 |  |  | -4.6986 |  | -4.6986 | -2.2908 | -2.3717 | 7 | 0.99352
12 | 2 | 2 | Myrtle | High Reference | 1 | 0 | -4.803 |  | -4.6579 |  |  | -4.6757 |  | -4.6757 | -2.2908 | -2.3498 | 8 | 0.9951
13 | 2 | 3 | Myrtle | High Reference | 1 | 0 | -4.673 |  | -4.6274 |  |  | -4.6467 |  | -4.6467 | -2.2908 | -2.322 | 9 | 0.9951
14 | 2 | 4 | Myrtle | High Reference | 1 | 0 | -4.676 |  | -4.6318 |  |  | -4.6526 |  | -4.6526 | -2.2908 | -2.3276 | 10 | 1
15 | 2 | 5 | Myrtle | High Reference | 1 | 0 | -4.796 |  | -4.7528 |  |  | -4.7751 |  | -4.7751 | -2.2908 | -2.445
16 | 2 | 6 | Myrtle | High Reference | 1 | 0 | -4.642 |  | -4.5976 |  |  | -4.6213 |  | -4.6213 | -2.2908 | -2.2976 | (2) Drift correction
17 | 2 | 7 | Myrtle | High Reference | 1 | 0 | -4.651 |  | -4.6164 |  |  | -4.6416 |  | -4.6416 | -2.2908 | -2.3171 | Slope m: | 0.0014828
18 | 2 | 8 | Myrtle | High Reference | 1 | 0 | -4.644 |  | -4.6178 |  |  | -4.6444 |  | -4.6444 | -2.2908 | -2.3198 | Stdev Mem Corr. | 0.11575
19 | 2 | 9 | Myrtle | High Reference | 1 | 0 | -4.611 |  | -4.5846 |  |  | -4.6128 |  | -4.6128 | -2.2908 | -2.2894 | Stdev Drift Corr. | 0.092862
20 | 2 | 10 | Myrtle | High Reference | 1 | 0 | -4.572 | 0.15219 | -4.572 | 0.052978 |  | -4.6017 |  | -4.6017 | -2.2908 | -2.2788
21 | 3 | 1 | Homer | Low Reference | 1 | 0 | -16.374 |  | -17.3699 |  |  | -17.4011 |  | -17.4011 | -14.5399 | -14.544 | (3) Normalization to VSMOW scale
22 | 3 | 2 | Homer | Low Reference | 1 | 0 | -17.045 |  | -17.3932 |  |  | -17.4259 |  | -17.4259 | -14.5399 | -14.5678 | slope m | 0.95827
23 | 3 | 3 | Homer | Low Reference | 1 | 0 | -17.283 |  | -17.3918 |  |  | -17.4259 |  | -17.4259 | -14.5399 | -14.5678 | interception t | 2.1308
24 | 3 | 4 | Homer | Low Reference | 1 | 0 | -17.398 |  | -17.5044 |  |  | -17.54 |  | -17.54 | -14.5399 | -14.6772
25 | 3 | 5 | Homer | Low Reference | 1 | 0 | -17.308 |  | -17.4136 |  |  | -17.4507 |  | -17.4507 | -14.5399 | -14.5916 | (4) Quality control |  | offset
26 | 3 | 6 | Homer | Low Reference | 1 | 0 | -17.227 |  | -17.332 |  |  | -17.3705 |  | -17.3705 | -14.5399 | -14.5148 | QC Check Defined | -3.6108 | accepted ± 0.15
27 | 3 | 7 | Homer | Low Reference | 1 | 0 | -17.284 |  | -17.3663 |  |  | -17.4064 |  | -17.4064 | -14.5399 | -14.5491 | QC Check Vial 5 | -3.6972 | -0.086381
28 | 3 | 8 | Homer | Low Reference | 1 | 0 | -17.33 |  | -17.3925 |  |  | -17.434 |  | -17.434 | -14.5399 | -14.5756 | QC Check Vial 26 | -3.5862 | 0.024645
29 | 3 | 9 | Homer | Low Reference | 1 | 0 | -17.325 |  | -17.3875 |  |  | -17.4305 |  | -17.4305 | -14.5399 | -14.5722
30 | 3 | 10 | Homer | Low Reference | 1 | 0 | -17.342 | 0.30255 | -17.342 | 0.047494 |  | -17.3865 |  | -17.3865 | -14.5399 | -14.5301 | Drift Defined | -7.5615 | accepted ± 0.15
31 | 4 | 1 | Blacksburg | Drift | 1 | 0 | -10.498 |  | -9.9205 |  |  | -9.9664 |  | -9.9664 | -7.5615 | -7.4197 | Drift Average | -7.5405 | 0.020957
32 | 4 | 2 | Blacksburg | Drift | 1 | 0 | -10.178 |  | -9.978 |  |  | -10.0254 |  | -10.0254 | -7.5615 | -7.4762 | Drift Vial 4 | -7.4908 | 0.070736
33 | 4 | 3 | Blacksburg | Drift | 1 | 0 | -10.06 |  | -9.9977 |  |  | -10.0466 |  | -10.0466 | -7.5615 | -7.4965 | Drift Vial 9 | -7.6864 | -0.12495
34 | 4 | 4 | Blacksburg | Drift | 1 | 0 | -10.117 |  | -10.0571 |  |  | -10.1075 |  | -10.1075 | -7.5615 | -7.5548 | Drift Vial 18 | -7.5416 | 0.019857
35 | 4 | 5 | Blacksburg | Drift | 1 | 0 | -10.071 |  | -10.0107 |  |  | -10.0626 |  | -10.0626 | -7.5615 | -7.5118 | Drift Vial 27 | -7.5062 | 0.055242
36 | 4 | 6 | Blacksburg | Drift | 1 | 0 | -10.048 |  | -9.9875 |  |  | -10.0409 |  | -10.0409 | -7.5615 | -7.491
37 | 4 | 7 | Blacksburg | Drift | 1 | 0 | -10.009 |  | -9.9615 |  |  | -10.0164 |  | -10.0164 | -7.5615 | -7.4675 | Reference Defined Values |  | d18O
38 | 4 | 8 | Blacksburg | Drift | 1 | 0 | -10.004 |  | -9.968 |  |  | -10.0244 |  | -10.0244 | -7.5615 | -7.4752 | High Reference | Myrtle Beach | -2.2908
39 | 4 | 9 | Blacksburg | Drift | 1 | 0 | -10.096 |  | -10.0605 |  |  | -10.1183 |  | -10.1183 | -7.5615 | -7.5652 | Low Reference | Homer | -14.5399
40 | 4 | 10 | Blacksburg | Drift | 1 | 0 | -9.944 | 0.15337 | -9.944 | 0.045127 |  | -10.0033 |  | -10.0033 | -7.5615 | -7.455 | Drift | Blacksburg | -7.5615
41 | 5 | 1 | Hawaii | QC Check | 1 | 0 | -6.27 |  | -5.96 |  |  | -6.0208 |  |  |  | -3.6387 | QC Check | Hawaii | -3.6108
42 | 5 | 2 | Hawaii | QC Check | 1 | 0 | -6.124 |  | -6.0173 |  |  | -6.0796 |  |  |  | -3.6951 | Calibration Date:
43 | 5 | 3 | Hawaii | QC Check | 1 | 0 | -6.131 |  | -6.0984 |  |  | -6.1621 |  |  |  | -3.7742
44 | 5 | 4 | Hawaii | QC Check | 1 | 0 | -6.032 | 0.098147 | -5.9996 | 0.058212 |  | -6.0648 |  |  |  | -3.6809
45 | 6 | 1 | CC4 17feb25 BOT |  | 1 | 0 | -7.551 |  | -7.6792 |  |  | -7.7459 |  |  |  | -5.2918
46 | 6 | 2 | CC4 17feb25 BOT |  | 1 | 0 | -7.643 |  | -7.688 |  |  | -7.7562 |  |  |  | -5.3017 | Denali | -19.9287
47 | 6 | 3 | CC4 17feb25 BOT |  | 1 | 0 | -7.636 |  | -7.6497 |  |  | -7.7194 |  |  |  | -5.2664 | Alberta | -18.8392
48 | 6 | 4 | CC4 17feb25 BOT |  | 1 | 0 | -7.689 | 0.057523 | -7.7027 | 0.022347 |  | -7.7739 |  |  |  | -5.3187 | MT BHB | -18.431
49 | 7 | 1 | C50 17feb25 20m |  | 1 | 0 | -7.558 |  | -7.5469 |  |  | -7.6196 |  |  |  | -5.1708 | Anchorage | -16.4204
50 | 7 | 2 | C50 17feb25 20m |  | 1 | 0 | -7.565 |  | -7.5615 |  |  | -7.6357 |  |  |  | -5.1862 | Homer | -14.5399
51 | 7 | 3 | C50 17feb25 20m |  | 1 | 0 | -7.658 |  | -7.6577 |  |  | -7.7334 |  |  |  | -5.2798 | MT Creek | -13.3562
52 | 7 | 4 | C50 17feb25 20m |  | 1 | 0 | -7.66 | 0.05637 | -7.6598 | 0.060635 |  | -7.7369 |  |  |  | -5.2832 | Nickel Mines | -7.876
53 | 8 | 1 | C50 17feb25 6m |  | 1 | 0 | -7.706 |  | -7.7099 |  |  | -7.7885 |  |  |  | -5.3326 | Blacksburg | -7.5615
54 | 8 | 2 | C50 17feb25 6m |  | 1 | 0 | -7.631 |  | -7.6302 |  |  | -7.7103 |  |  |  | -5.2577 | Denton's Well | -6.87
55 | 8 | 3 | C50 17feb25 6m |  | 1 | 0 | -7.628 |  | -7.6277 |  |  | -7.7093 |  |  |  | -5.2567 | Hawaii | -3.6108
56 | 8 | 4 | C50 17feb25 6m |  | 1 | 0 | -7.718 | 0.047898 | -7.7185 | 0.049339 |  | -7.8015 |  |  |  | -5.3451 | Myrtle Beach | -2.2908
57 | 9 | 1 | Blacksburg | Drift | 1 | 0 | -9.987 |  | -10.1785 |  | -10.1785 | -10.263 | -10.263 |  |  | -7.7039 | Heavy Oxygen | 3.1753
58 | 9 | 2 | Blacksburg | Drift | 1 | 0 | -10.087 |  | -10.1531 |  | -10.1531 | -10.2391 | -10.2391 |  |  | -7.681 | Heavy Hydrogen | 11.1372
59 | 9 | 3 | Blacksburg | Drift | 1 | 0 | -10.137 |  | -10.1577 |  | -10.1577 | -10.2452 | -10.2452 |  |  | -7.6868 | Andres Heavy Water | 17.5664
60 | 9 | 4 | Blacksburg | Drift | 1 | 0 | -10.123 | 0.067693 | -10.1429 | 0.014936 | -10.1429 | -10.2319 | -10.2319 |  |  | -7.6741 | Picarro ZERO | 0.3
61 | 10 | 1 | CC4 17feb25 1.5m |  | 1 | 0 | -7.932 |  | -7.7471 |  |  | -7.8376 |  |  |  | -5.3797 | Picarro MID | -20.6
62 | 10 | 2 | CC4 17feb25 1.5m |  | 1 | 0 | -7.695 |  | -7.6272 |  |  | -7.7191 |  |  |  | -5.2662 | Picarro DEPL | -29.6
63 | 10 | 3 | CC4 17feb25 1.5m |  | 1 | 0 | -7.667 |  | -7.646 |  |  | -7.7394 |  |  |  | -5.2856
64 | 10 | 4 | CC4 17feb25 1.5m |  | 1 | 0 | -7.685 | 0.12537 | -7.6648 | 0.052835 |  | -7.7597 |  |  |  | -5.305
65 | 11 | 1 | CC3 15aug24 0.1m |  | 1 | 0 | -6.617 |  | -6.5269 |  |  | -6.6233 |  |  |  | -4.216
66 | 11 | 2 | CC3 15aug24 0.1m |  | 1 | 0 | -6.431 |  | -6.396 |  |  | -6.4939 |  |  |  | -4.092
67 | 11 | 3 | CC3 15aug24 0.1m |  | 1 | 0 | -6.438 |  | -6.4273 |  |  | -6.5267 |  |  |  | -4.1235
68 | 11 | 4 | CC3 15aug24 0.1m |  | 1 | 0 | -6.538 | 0.088683 | -6.5285 | 0.068199 |  | -6.6293 |  |  |  | -4.2218
69 | 12 | 1 | CP1 17feb25 BOT |  | 1 | 0 | -8.695 |  | -8.877 |  |  | -8.9793 |  |  |  | -6.4738
70 | 12 | 2 | CP1 17feb25 BOT |  | 1 | 0 | -8.67 |  | -8.7295 |  |  | -8.8333 |  |  |  | -6.3339
71 | 12 | 3 | CP1 17feb25 BOT |  | 1 | 0 | -8.803 |  | -8.8224 |  |  | -8.9277 |  |  |  | -6.4243
72 | 12 | 4 | CP1 17feb25 BOT |  | 1 | 0 | -8.77 | 0.062378 | -8.7885 | 0.061794 |  | -8.8953 |  |  |  | -6.3932
73 | 13 | 1 | CC5 22may24 13m |  | 1 | 0 | -7.872 |  | -7.7962 |  |  | -7.9045 |  |  |  | -5.4438
74 | 13 | 2 | CC5 22may24 13m |  | 1 | 0 | -7.818 |  | -7.7914 |  |  | -7.9011 |  |  |  | -5.4406
75 | 13 | 3 | CC5 22may24 13m |  | 1 | 0 | -7.918 |  | -7.9107 |  |  | -8.0219 |  |  |  | -5.5563
76 | 13 | 4 | CC5 22may24 13m |  | 1 | 0 | -7.815 | 0.049074 | -7.8071 | 0.056615 |  | -7.9198 |  |  |  | -5.4584
77 | 14 | 1 | CC2 22may24 BOT |  | 1 | 0 | -7.578 |  | -7.558 |  |  | -7.6722 |  |  |  | -5.2212
78 | 14 | 2 | CC2 22may24 BOT |  | 1 | 0 | -7.463 |  | -7.4532 |  |  | -7.5688 |  |  |  | -5.1221
79 | 14 | 3 | CC2 22may24 BOT |  | 1 | 0 | -7.541 |  | -7.5387 |  |  | -7.6558 |  |  |  | -5.2055
80 | 14 | 4 | CC2 22may24 BOT |  | 1 | 0 | -7.552 | 0.049494 | -7.5498 | 0.048479 |  | -7.6684 |  |  |  | -5.2176
81 | 15 | 1 | CP2 17feb25 BOT |  | 1 | 0 | -7.648 |  | -7.6561 |  |  | -7.7762 |  |  |  | -5.3209
82 | 15 | 2 | CP2 17feb25 BOT |  | 1 | 0 | -7.544 |  | -7.5438 |  |  | -7.6654 |  |  |  | -5.2147
83 | 15 | 3 | CP2 17feb25 BOT |  | 1 | 0 | -7.47 |  | -7.4693 |  |  | -7.5924 |  |  |  | -5.1447
84 | 15 | 4 | CP2 17feb25 BOT |  | 1 | 0 | -7.423 | 0.098086 | -7.4219 | 0.10206 |  | -7.5465 |  |  |  | -5.1007
85 | 16 | 1 | CS1 15aug24 0.1m |  | 1 | 0 | -8.643 |  | -8.7459 |  |  | -8.872 |  |  |  | -6.3709
86 | 16 | 2 | CS1 15aug24 0.1m |  | 1 | 0 | -8.702 |  | -8.7377 |  |  | -8.8652 |  |  |  | -6.3644
87 | 16 | 3 | CS1 15aug24 0.1m |  | 1 | 0 | -8.718 |  | -8.7291 |  |  | -8.8581 |  |  |  | -6.3576
88 | 16 | 4 | CS1 15aug24 0.1m |  | 1 | 0 | -8.693 | 0.032363 | -8.7035 | 0.018363 |  | -8.834 |  |  |  | -6.3345
89 | 17 | 1 | CC2 17feb25 BOT |  | 1 | 0 | -7.483 |  | -7.3809 |  |  | -7.5129 |  |  |  | -5.0685
90 | 17 | 2 | CC2 17feb25 BOT |  | 1 | 0 | -7.454 |  | -7.4194 |  |  | -7.5529 |  |  |  | -5.1068
91 | 17 | 3 | CC2 17feb25 BOT |  | 1 | 0 | -7.439 |  | -7.4283 |  |  | -7.5632 |  |  |  | -5.1168
92 | 17 | 4 | CC2 17feb25 BOT |  | 1 | 0 | -7.54 | 0.04458 | -7.5304 | 0.063858 |  | -7.6669 |  |  |  | -5.2161
93 | 18 | 1 | Blacksburg | Drift | 1 | 0 | -9.758 |  | -9.9452 |  | -9.9452 | -10.0831 | -10.0831 |  |  | -7.5315
94 | 18 | 2 | Blacksburg | Drift | 1 | 0 | -9.947 |  | -10.0142 |  | -10.0142 | -10.1536 | -10.1536 |  |  | -7.599
95 | 18 | 3 | Blacksburg | Drift | 1 | 0 | -9.907 |  | -9.9273 |  | -9.9273 | -10.0681 | -10.0681 |  |  | -7.5171
96 | 18 | 4 | Blacksburg | Drift | 1 | 0 | -9.908 | 0.083439 | -9.9276 | 0.041278 | -9.9276 | -10.07 | -10.07 |  |  | -7.5189
97 | 19 | 1 | CS1 17feb25 0.1m |  | 1 | 0 | -9.046 |  | -8.9733 |  |  | -9.1171 |  |  |  | -6.6058
98 | 19 | 2 | CS1 17feb25 0.1m |  | 1 | 0 | -9.066 |  | -9.0425 |  |  | -9.1878 |  |  |  | -6.6736
99 | 19 | 3 | CS1 17feb25 0.1m |  | 1 | 0 | -8.973 |  | -8.965 |  |  | -9.1118 |  |  |  | -6.6007
100 | 19 | 4 | CS1 17feb25 0.1m |  | 1 | 0 | -9.048 | 0.041161 | -9.0409 | 0.042027 |  | -9.1891 |  |  |  | -6.6748
101 | 20 | 1 | CC2 17feb25 0.1m |  | 1 | 0 | -7.547 |  | -7.4203 |  |  | -7.5701 |  |  |  | -5.1234
102 | 20 | 2 | CC2 17feb25 0.1m |  | 1 | 0 | -7.525 |  | -7.4825 |  |  | -7.6337 |  |  |  | -5.1843
103 | 20 | 3 | CC2 17feb25 0.1m |  | 1 | 0 | -7.456 |  | -7.4424 |  |  | -7.5951 |  |  |  | -5.1473
104 | 20 | 4 | CC2 17feb25 0.1m |  | 1 | 0 | -7.496 | 0.039336 | -7.4831 | 0.031036 |  | -7.6373 |  |  |  | -5.1878
105 | 21 | 1 | CS1 22may24 0.1m |  | 1 | 0 | -9.005 |  | -9.1323 |  |  | -9.288 |  |  |  | -6.7696
106 | 21 | 2 | CS1 22may24 0.1m |  | 1 | 0 | -9.212 |  | -9.2599 |  |  | -9.4171 |  |  |  | -6.8933
107 | 21 | 3 | CS1 22may24 0.1m |  | 1 | 0 | -9.208 |  | -9.2226 |  |  | -9.3813 |  |  |  | -6.859
108 | 21 | 4 | CS1 22may24 0.1m |  | 1 | 0 | -9.204 | 0.10155 | -9.2182 | 0.053966 |  | -9.3783 |  |  |  | -6.8561
109 | 22 | 1 | CP1 17feb25 0.1m |  | 1 | 0 | -8.966 |  | -8.9459 |  |  | -9.1075 |  |  |  | -6.5966
110 | 22 | 2 | CP1 17feb25 0.1m |  | 1 | 0 | -9.062 |  | -9.058 |  |  | -9.2211 |  |  |  | -6.7055
111 | 22 | 3 | CP1 17feb25 0.1m |  | 1 | 0 | -9.012 |  | -9.0104 |  |  | -9.1749 |  |  |  | -6.6612
112 | 22 | 4 | CP1 17feb25 0.1m |  | 1 | 0 | -8.965 | 0.046053 | -8.963 | 0.050463 |  | -9.1291 |  |  |  | -6.6173
113 | 23 | 1 | CP2 17feb25 0.1m |  | 1 | 0 | -8.694 |  | -8.6711 |  |  | -8.8387 |  |  |  | -6.339
114 | 23 | 2 | CP2 17feb25 0.1m |  | 1 | 0 | -8.757 |  | -8.7512 |  |  | -8.9202 |  |  |  | -6.4171
115 | 23 | 3 | CP2 17feb25 0.1m |  | 1 | 0 | -8.721 |  | -8.7189 |  |  | -8.8894 |  |  |  | -6.3876
116 | 23 | 4 | CP2 17feb25 0.1m |  | 1 | 0 | -8.794 | 0.043486 | -8.7926 | 0.051336 |  | -8.9646 |  |  |  | -6.4596
117 | 24 | 1 | CC3 17feb25 BOT |  | 1 | 0 | -7.458 |  | -7.3453 |  |  | -7.5188 |  |  |  | -5.0742
118 | 24 | 2 | CC3 17feb25 BOT |  | 1 | 0 | -7.535 |  | -7.4998 |  |  | -7.6748 |  |  |  | -5.2237
119 | 24 | 3 | CC3 17feb25 BOT |  | 1 | 0 | -7.37 |  | -7.3578 |  |  | -7.5343 |  |  |  | -5.089
120 | 24 | 4 | CC3 17feb25 BOT |  | 1 | 0 | -7.402 | 0.072311 | -7.3905 | 0.070299 |  | -7.5684 |  |  |  | -5.1217
121 | 25 | 1 | CS1 15aug24 0.1m - DUP |  | 1 | 0 | -8.564 |  | -8.6621 |  |  | -8.8415 |  |  |  | -6.3417
122 | 25 | 2 | CS1 15aug24 0.1m - DUP |  | 1 | 0 | -8.53 |  | -8.5615 |  |  | -8.7424 |  |  |  | -6.2467
123 | 25 | 3 | CS1 15aug24 0.1m - DUP |  | 1 | 0 | -8.612 |  | -8.6224 |  |  | -8.8047 |  |  |  | -6.3065
124 | 25 | 4 | CS1 15aug24 0.1m - DUP |  | 1 | 0 | -8.645 | 0.050875 | -8.6553 | 0.045941 |  | -8.8392 |  |  |  | -6.3395
125 | 26 | 1 | Hawaii |  | 1 | 0 | -6.035 |  | -5.8148 |  |  | -6.0001 |  |  |  | -3.6189
126 | 26 | 2 | Hawaii |  | 1 | 0 | -5.797 |  | -5.7175 |  |  | -5.9043 |  |  |  | -3.5271
127 | 26 | 3 | Hawaii |  | 1 | 0 | -5.921 |  | -5.8977 |  |  | -6.086 |  |  |  | -3.7012
128 | 26 | 4 | Hawaii |  | 1 | 0 | -5.708 | 0.14296 | -5.6836 | 0.097027 |  | -5.8734 |  |  |  | -3.4975
129 | 27 | 1 | Blacksburg | Drift | 1 | 0 | -9.51 |  | -9.8308 |  | -9.8308 | -10.0221 | -10.0221 |  |  | -7.473
130 | 27 | 2 | Blacksburg | Drift | 1 | 0 | -9.783 |  | -9.8968 |  | -9.8968 | -10.0895 | -10.0895 |  |  | -7.5377
131 | 27 | 3 | Blacksburg | Drift | 1 | 0 | -9.836 |  | -9.8713 |  | -9.8713 | -10.0656 | -10.0656 |  |  | -7.5147
132 | 27 | 4 | Blacksburg | Drift | 1 | 0 | -9.82 | 0.15312 | -9.8541 | 0.027829 | -9.8541 | -10.0498 | -10.0498 |  |  | -7.4996
Note on G2: Manually change first six injections to ignore "-1".
Note on G3: Manually change first six injections to ignore "-1".
Note on G4: Manually change first six injections to ignore "-1".
Note on G5: Manually change first six injections to ignore "-1".
Note on G6: Manually change first six injections to ignore "-1".
Note on G7: Manually change first six injections to ignore "-1".

## 2H
Line | Sample | Injection | Sample Name | Correction Type | Good | Ignore | dD_H) Picarro Raw | Stdev Raw | d(D_H) Memory Corrected | Stdev Memory Corrected | d(D_H) Memory Corrected Drift | d(D_H) Drift Corrected | Drift Corrected Drift | d(D_H) Drift Corrected High, Low, Drift | d(D_H) Reference Defined | dD Final Corrected Value | Post run corrections:
1 | 1 | 1 | Blacksburg | Drift | 1 | -1 | -46.647 | [9 empty] | (1) Memory correction |  |  | (1) Graphs for memory correction
2 | 1 | 2 | Blacksburg | Drift | 1 | -1 | -43.011 | [9 empty] | Combined Stdev | 0.17846
3 | 1 | 3 | Blacksburg | Drift | 1 | -1 | -42.545 | [9 empty] | Date:
4 | 1 | 4 | Blacksburg | Drift | 1 | -1 | -42.237 | [9 empty] | injection | 18O memory coef
5 | 1 | 5 | Blacksburg | Drift | 1 | -1 | -41.934 | [9 empty] | 1 | 0.84436
6 | 1 | 6 | Blacksburg | Drift | 1 | -1 | -42.269 | [9 empty] | 2 | 0.96009
7 | 1 | 7 | Blacksburg | Drift | 1 | 0 | -42.343 |  | -42.343 |  | -42.343 | -42.3582 | -42.3582 |  |  | -48.0049 | 3 | 0.97662
8 | 1 | 8 | Blacksburg | Drift | 1 | 0 | -42.167 |  | -42.167 |  | -42.167 | -42.1843 | -42.1843 |  |  | -47.8392 | 4 | 0.98737
9 | 1 | 9 | Blacksburg | Drift | 1 | 0 | -42.147 |  | -42.147 |  | -42.147 | -42.1665 | -42.1665 |  |  | -47.8222 | 5 | 0.99074
10 | 1 | 10 | Blacksburg | Drift | 1 | 0 | -42.018 | 0.13362 | -42.018 | 0.13362 | -42.018 | -42.0397 | -42.0397 |  |  | -47.7013 | 6 | 0.99464
11 | 2 | 1 | Myrtle | High Reference | 1 | 0 | -8.127 |  | -2.8521 |  |  | -2.8759 |  | -2.8759 | -10.453 | -10.3728 | 7 | 0.99658
12 | 2 | 2 | Myrtle | High Reference | 1 | 0 | -4.225 |  | -2.7169 |  |  | -2.7429 |  | -2.7429 | -10.453 | -10.246 | 8 | 0.99889
13 | 2 | 3 | Myrtle | High Reference | 1 | 0 | -3.581 |  | -2.6823 |  |  | -2.7105 |  | -2.7105 | -10.453 | -10.2151 | 9 | 0.99889
14 | 2 | 4 | Myrtle | High Reference | 1 | 0 | -3.305 |  | -2.816 |  |  | -2.8464 |  | -2.8464 | -10.453 | -10.3447 | 10 | 1
15 | 2 | 5 | Myrtle | High Reference | 1 | 0 | -2.96 |  | -2.5982 |  |  | -2.6307 |  | -2.6307 | -10.453 | -10.1391
16 | 2 | 6 | Myrtle | High Reference | 1 | 0 | -3.021 |  | -2.8122 |  |  | -2.8468 |  | -2.8468 | -10.453 | -10.3451 | (2) Drift correction
17 | 2 | 7 | Myrtle | High Reference | 1 | 0 | -2.744 |  | -2.6098 |  |  | -2.6466 |  | -2.6466 | -10.453 | -10.1542 | Slope m: | 0.0021664
18 | 2 | 8 | Myrtle | High Reference | 1 | 0 | -2.876 |  | -2.8324 |  |  | -2.8714 |  | -2.8714 | -10.453 | -10.3685 | Stdev Mem Corr. | 0.34703
19 | 2 | 9 | Myrtle | High Reference | 1 | 0 | -2.793 |  | -2.7493 |  |  | -2.7905 |  | -2.7905 | -10.453 | -10.2914 | Stdev Drift Corr. | 0.33202
20 | 2 | 10 | Myrtle | High Reference | 1 | 0 | -2.983 | 1.6319 | -2.983 | 0.11823 |  | -3.0263 |  | -3.0263 | -10.453 | -10.5162
21 | 3 | 1 | Homer | Low Reference | 1 | 0 | -93.746 |  | -107.8726 |  |  | -107.9181 |  | -107.9181 | -110.5169 | -110.4926 | (3) Normalization to VSMOW scale
22 | 3 | 2 | Homer | Low Reference | 1 | 0 | -103.773 |  | -107.795 |  |  | -107.8427 |  | -107.8427 | -110.5169 | -110.4208 | slope m | 0.95314
23 | 3 | 3 | Homer | Low Reference | 1 | 0 | -105.376 |  | -107.77 |  |  | -107.8198 |  | -107.8198 | -110.5169 | -110.3989 | interception t | -7.6317
24 | 3 | 4 | Homer | Low Reference | 1 | 0 | -106.349 |  | -107.6545 |  |  | -107.7065 |  | -107.7065 | -110.5169 | -110.2909
25 | 3 | 5 | Homer | Low Reference | 1 | 0 | -106.773 |  | -107.7344 |  |  | -107.7885 |  | -107.7885 | -110.5169 | -110.3691 | (4) Quality control |  | offset
26 | 3 | 6 | Homer | Low Reference | 1 | 0 | -107.201 |  | -107.7591 |  |  | -107.8154 |  | -107.8154 | -110.5169 | -110.3948 | QC Check Defined | -13.2294 | accepted ±1.5
27 | 3 | 7 | Homer | Low Reference | 1 | 0 | -107.375 |  | -107.7317 |  |  | -107.7902 |  | -107.7902 | -110.5169 | -110.3707 | QC Check Vial 5 | -12.7 | 0.5294
28 | 3 | 8 | Homer | Low Reference | 1 | 0 | -107.753 |  | -107.8696 |  |  | -107.9303 |  | -107.9303 | -110.5169 | -110.5042 | QC Check Vial 26 | -13.0827 | 0.1467
29 | 3 | 9 | Homer | Low Reference | 1 | 0 | -107.663 |  | -107.7795 |  |  | -107.8424 |  | -107.8424 | -110.5169 | -110.4204
30 | 3 | 10 | Homer | Low Reference | 1 | 0 | -107.951 | 4.289 | -107.951 | 0.085234 |  | -108.016 |  | -108.016 | -110.5169 | -110.5859 | Drift Defined | -47.446 | accepted ±1.5
31 | 4 | 1 | Blacksburg | Drift | 1 | 0 | -50.908 |  | -42.0297 |  |  | -42.0968 |  | -42.0968 | -47.446 | -47.7558 | Drift Average | -47.5993 | -0.15329
32 | 4 | 2 | Blacksburg | Drift | 1 | 0 | -44.518 |  | -41.9867 |  |  | -42.056 |  | -42.056 | -47.446 | -47.7169 | Drift Vial 4 | -47.7436 | -0.29758
33 | 4 | 3 | Blacksburg | Drift | 1 | 0 | -43.475 |  | -41.9675 |  |  | -42.039 |  | -42.039 | -47.446 | -47.7007 | Drift Vial 9 | -47.1091 | 0.33689
34 | 4 | 4 | Blacksburg | Drift | 1 | 0 | -42.532 |  | -41.7057 |  |  | -41.7794 |  | -41.7794 | -47.446 | -47.4532 | Drift Vial 18 | -47.5245 | -0.07848
35 | 4 | 5 | Blacksburg | Drift | 1 | 0 | -42.567 |  | -41.9614 |  |  | -42.0372 |  | -42.0372 | -47.446 | -47.6989 | Drift Vial 27 | -47.7774 | -0.33138
36 | 4 | 6 | Blacksburg | Drift | 1 | 0 | -42.225 |  | -41.873 |  |  | -41.951 |  | -41.951 | -47.446 | -47.6168
37 | 4 | 7 | Blacksburg | Drift | 1 | 0 | -42.175 |  | -41.9502 |  |  | -42.0304 |  | -42.0304 | -47.446 | -47.6925 | Reference Defined Values |  | d2H
38 | 4 | 8 | Blacksburg | Drift | 1 | 0 | -42.105 |  | -42.0317 |  |  | -42.114 |  | -42.114 | -47.446 | -47.7722 | High Reference | North Myrtle Beach, SC | -10.453
39 | 4 | 9 | Blacksburg | Drift | 1 | 0 | -42.021 |  | -41.9476 |  |  | -42.0321 |  | -42.0321 | -47.446 | -47.6941 | Low Reference | Homer, AK | -110.5169
40 | 4 | 10 | Blacksburg | Drift | 1 | 0 | -42.073 | 2.7337 | -42.073 | 0.10298 |  | -42.1597 |  | -42.1597 | -47.446 | -47.8157 | Drift | Blacksburg | -47.446
41 | 5 | 1 | Hawaii | QC Check | 1 | 0 | -10.476 |  | -5.5582 |  |  | -5.647 |  |  |  | -13.014 | QC Check | Hawaii | -13.2294
42 | 5 | 2 | Hawaii | QC Check | 1 | 0 | -6.728 |  | -5.3176 |  |  | -5.4085 |  |  |  | -12.7867 | Calibration Date:
43 | 5 | 3 | Hawaii | QC Check | 1 | 0 | -5.763 |  | -4.9141 |  |  | -5.0072 |  |  |  | -12.4042
44 | 5 | 4 | Hawaii | QC Check | 1 | 0 | -5.573 | 2.284 | -5.112 | 0.27629 |  | -5.2073 |  |  |  | -12.595
45 | 6 | 1 | CC4 17feb25 BOT |  | 1 | 0 | -22.036 |  | -24.5983 |  |  | -24.6958 |  |  |  | -31.1702
46 | 6 | 2 | CC4 17feb25 BOT |  | 1 | 0 | -23.21 |  | -23.9138 |  |  | -24.0135 |  |  |  | -30.5198 | Denali National Park | -153.7172
47 | 6 | 3 | CC4 17feb25 BOT |  | 1 | 0 | -23.819 |  | -24.2456 |  |  | -24.3474 |  |  |  | -30.8381 | Edmonton, Alberta | -146.8245
48 | 6 | 4 | CC4 17feb25 BOT |  | 1 | 0 | -23.668 | 0.80748 | -23.8965 | 0.33138 |  | -24.0005 |  |  |  | -30.5075 | MT BHB | -140.5778
49 | 7 | 1 | C50 17feb25 20m |  | 1 | 0 | -23.929 |  | -23.9696 |  |  | -24.0758 |  |  |  | -30.5792 | Anchorage, AK | -129.9028
50 | 7 | 2 | C50 17feb25 20m |  | 1 | 0 | -23.985 |  | -23.9976 |  |  | -24.106 |  |  |  | -30.608 | Homer, AK | -110.5169
51 | 7 | 3 | C50 17feb25 20m |  | 1 | 0 | -23.639 |  | -23.6383 |  |  | -23.7488 |  |  |  | -30.2676 | MT Creek | -97.7428
52 | 7 | 4 | C50 17feb25 20m |  | 1 | 0 | -23.926 | 0.15621 | -23.9293 | 0.16599 |  | -24.0419 |  |  |  | -30.5469 | Nickel Mines, PA | -48.7951
53 | 8 | 1 | C50 17feb25 6m |  | 1 | 0 | -24.015 |  | -24.0289 |  |  | -24.1437 |  |  |  | -30.6439 | Blacksburg | -47.446
54 | 8 | 2 | C50 17feb25 6m |  | 1 | 0 | -23.581 |  | -23.5672 |  |  | -23.6842 |  |  |  | -30.206 | Denton's Well | -41.771
55 | 8 | 3 | C50 17feb25 6m |  | 1 | 0 | -23.795 |  | -23.7919 |  |  | -23.9111 |  |  |  | -30.4223 | Hawaii | -13.2294
56 | 8 | 4 | C50 17feb25 6m |  | 1 | 0 | -24.117 | 0.23874 | -24.1194 | 0.24834 |  | -24.2407 |  |  |  | -30.7364 | North Myrtle Beach, SC | -10.453
57 | 9 | 1 | Blacksburg | Drift | 1 | 0 | -38.838 |  | -41.1292 |  | -41.1292 | -41.2527 | -41.2527 |  |  | -46.9512 | WQL Heavy Oxygen | -15.7881
58 | 9 | 2 | Blacksburg | Drift | 1 | 0 | -40.646 |  | -41.3056 |  | -41.3056 | -41.4312 | -41.4312 |  |  | -47.1214 | WQL Heavy Hydrogen | 4.0332
59 | 9 | 3 | Blacksburg | Drift | 1 | 0 | -41.022 |  | -41.4172 |  | -41.4172 | -41.5451 | -41.5451 |  |  | -47.2299 | Andres Heavy Water | 20.5382
60 | 9 | 4 | Blacksburg | Drift | 1 | 0 | -41.1 | 1.061 | -41.3145 | 0.11955 | -41.3145 | -41.4445 | -41.4445 |  |  | -47.134
61 | 10 | 1 | CC4 17feb25 1.5m |  | 1 | 0 | -26.007 |  | -23.6579 |  |  | -23.79 |  |  |  | -30.3069
62 | 10 | 2 | CC4 17feb25 1.5m |  | 1 | 0 | -24.471 |  | -23.8074 |  |  | -23.9417 |  |  |  | -30.4515
63 | 10 | 3 | CC4 17feb25 1.5m |  | 1 | 0 | -24.288 |  | -23.8949 |  |  | -24.0314 |  |  |  | -30.5369
64 | 10 | 4 | CC4 17feb25 1.5m |  | 1 | 0 | -23.988 | 0.90127 | -23.7719 | 0.098152 |  | -23.9105 |  |  |  | -30.4217
65 | 11 | 1 | CC3 15aug24 0.1m |  | 1 | 0 | -19.647 |  | -18.9714 |  |  | -19.1122 |  |  |  | -25.8482
66 | 11 | 2 | CC3 15aug24 0.1m |  | 1 | 0 | -19.078 |  | -18.8821 |  |  | -19.025 |  |  |  | -25.7652
67 | 11 | 3 | CC3 15aug24 0.1m |  | 1 | 0 | -19.037 |  | -18.9212 |  |  | -19.0664 |  |  |  | -25.8046
68 | 11 | 4 | CC3 15aug24 0.1m |  | 1 | 0 | -19.076 | 0.29228 | -19.014 | 0.05761 |  | -19.1613 |  |  |  | -25.895
69 | 12 | 1 | CP1 17feb25 BOT |  | 1 | 0 | -27.748 |  | -29.0977 |  |  | -29.2472 |  |  |  | -35.5083
70 | 12 | 2 | CP1 17feb25 BOT |  | 1 | 0 | -28.683 |  | -29.0664 |  |  | -29.218 |  |  |  | -35.4805
71 | 12 | 3 | CP1 17feb25 BOT |  | 1 | 0 | -29.025 |  | -29.2576 |  |  | -29.4114 |  |  |  | -35.6648
72 | 12 | 4 | CP1 17feb25 BOT |  | 1 | 0 | -29.073 | 0.61454 | -29.1993 | 0.088738 |  | -29.3552 |  |  |  | -35.6113
73 | 13 | 1 | CC5 22may24 13m |  | 1 | 0 | -25.87 |  | -25.3715 |  |  | -25.5296 |  |  |  | -31.9649
74 | 13 | 2 | CC5 22may24 13m |  | 1 | 0 | -25.633 |  | -25.4957 |  |  | -25.656 |  |  |  | -32.0854
75 | 13 | 3 | CC5 22may24 13m |  | 1 | 0 | -25.576 |  | -25.4942 |  |  | -25.6567 |  |  |  | -32.0861
76 | 13 | 4 | CC5 22may24 13m |  | 1 | 0 | -25.774 | 0.13359 | -25.7323 | 0.15094 |  | -25.897 |  |  |  | -32.3151
77 | 14 | 1 | CC2 22may24 BOT |  | 1 | 0 | -24.492 |  | -24.2925 |  |  | -24.4593 |  |  |  | -30.9448
78 | 14 | 2 | CC2 22may24 BOT |  | 1 | 0 | -24.32 |  | -24.262 |  |  | -24.431 |  |  |  | -30.9178
79 | 14 | 3 | CC2 22may24 BOT |  | 1 | 0 | -24.359 |  | -24.3259 |  |  | -24.4971 |  |  |  | -30.9808
80 | 14 | 4 | CC2 22may24 BOT |  | 1 | 0 | -24.509 | 0.094562 | -24.493 | 0.10314 |  | -24.6663 |  |  |  | -31.1421
81 | 15 | 1 | CP2 17feb25 BOT |  | 1 | 0 | -24.283 |  | -24.2478 |  |  | -24.4233 |  |  |  | -30.9105
82 | 15 | 2 | CP2 17feb25 BOT |  | 1 | 0 | -24.64 |  | -24.6452 |  |  | -24.8229 |  |  |  | -31.2913
83 | 15 | 3 | CP2 17feb25 BOT |  | 1 | 0 | -24.338 |  | -24.334 |  |  | -24.5138 |  |  |  | -30.9967
84 | 15 | 4 | CP2 17feb25 BOT |  | 1 | 0 | -24.918 | 0.29419 | -24.9232 | 0.30858 |  | -25.1051 |  |  |  | -31.5604
85 | 16 | 1 | CS1 15aug24 0.1m |  | 1 | 0 | -30.102 |  | -30.9089 |  |  | -31.093 |  |  |  | -37.2676
86 | 16 | 2 | CS1 15aug24 0.1m |  | 1 | 0 | -31.291 |  | -31.5453 |  |  | -31.7316 |  |  |  | -37.8763
87 | 16 | 3 | CS1 15aug24 0.1m |  | 1 | 0 | -31.056 |  | -31.1995 |  |  | -31.388 |  |  |  | -37.5488
88 | 16 | 4 | CS1 15aug24 0.1m |  | 1 | 0 | -31.259 | 0.55976 | -31.3391 | 0.26712 |  | -31.5297 |  |  |  | -37.6839
89 | 17 | 1 | CC2 17feb25 BOT |  | 1 | 0 | -24.989 |  | -24.0131 |  |  | -24.2059 |  |  |  | -30.7033
90 | 17 | 2 | CC2 17feb25 BOT |  | 1 | 0 | -24.606 |  | -24.3405 |  |  | -24.5355 |  |  |  | -31.0174
91 | 17 | 3 | CC2 17feb25 BOT |  | 1 | 0 | -24.546 |  | -24.389 |  |  | -24.5862 |  |  |  | -31.0657
92 | 17 | 4 | CC2 17feb25 BOT |  | 1 | 0 | -24.428 | 0.2427 | -24.3417 | 0.17347 |  | -24.541 |  |  |  | -31.0227
93 | 18 | 1 | Blacksburg | Drift | 1 | 0 | -39.313 |  | -41.6297 |  | -41.6297 | -41.8312 | -41.8312 |  |  | -47.5026
94 | 18 | 2 | Blacksburg | Drift | 1 | 0 | -40.964 |  | -41.6239 |  | -41.6239 | -41.8275 | -41.8275 |  |  | -47.4991
95 | 18 | 3 | Blacksburg | Drift | 1 | 0 | -41.287 |  | -41.6812 |  | -41.6812 | -41.887 | -41.887 |  |  | -47.5558
96 | 18 | 4 | Blacksburg | Drift | 1 | 0 | -41.448 | 0.98087 | -41.663 | 0.027273 | -41.663 | -41.8709 | -41.8709 |  |  | -47.5405
97 | 19 | 1 | CS1 17feb25 0.1m |  | 1 | 0 | -32.733 |  | -31.3766 |  |  | -31.5867 |  |  |  | -37.7382
98 | 19 | 2 | CS1 17feb25 0.1m |  | 1 | 0 | -31.352 |  | -30.9491 |  |  | -31.1614 |  |  |  | -37.3328
99 | 19 | 3 | CS1 17feb25 0.1m |  | 1 | 0 | -31.484 |  | -31.251 |  |  | -31.4655 |  |  |  | -37.6227
100 | 19 | 4 | CS1 17feb25 0.1m |  | 1 | 0 | -31.193 | 0.70511 | -31.0635 | 0.1906 |  | -31.2801 |  |  |  | -37.446
101 | 20 | 1 | CC2 17feb25 0.1m |  | 1 | 0 | -25.034 |  | -24.0754 |  |  | -24.2942 |  |  |  | -30.7874
102 | 20 | 2 | CC2 17feb25 0.1m |  | 1 | 0 | -24.712 |  | -24.4534 |  |  | -24.6743 |  |  |  | -31.1497
103 | 20 | 3 | CC2 17feb25 0.1m |  | 1 | 0 | -24.384 |  | -24.2248 |  |  | -24.4479 |  |  |  | -30.9339
104 | 20 | 4 | CC2 17feb25 0.1m |  | 1 | 0 | -24.735 | 0.26566 | -24.6534 | 0.25418 |  | -24.8787 |  |  |  | -31.3446
105 | 21 | 1 | CS1 22may24 0.1m |  | 1 | 0 | -31.431 |  | -32.4732 |  |  | -32.7007 |  |  |  | -38.7999
106 | 21 | 2 | CS1 22may24 0.1m |  | 1 | 0 | -32.092 |  | -32.3856 |  |  | -32.6152 |  |  |  | -38.7185
107 | 21 | 3 | CS1 22may24 0.1m |  | 1 | 0 | -32.556 |  | -32.7389 |  |  | -32.9707 |  |  |  | -39.0573
108 | 21 | 4 | CS1 22may24 0.1m |  | 1 | 0 | -32.682 | 0.56619 | -32.7824 | 0.19536 |  | -33.0163 |  |  |  | -39.1008
109 | 22 | 1 | CP1 17feb25 0.1m |  | 1 | 0 | -31.485 |  | -31.2987 |  |  | -31.5348 |  |  |  | -37.6887
110 | 22 | 2 | CP1 17feb25 0.1m |  | 1 | 0 | -31.048 |  | -30.9828 |  |  | -31.2211 |  |  |  | -37.3897
111 | 22 | 3 | CP1 17feb25 0.1m |  | 1 | 0 | -31.091 |  | -31.0538 |  |  | -31.2943 |  |  |  | -37.4595
112 | 22 | 4 | CP1 17feb25 0.1m |  | 1 | 0 | -31.591 | 0.27449 | -31.5772 | 0.26921 |  | -31.8199 |  |  |  | -37.9604
113 | 23 | 1 | CP2 17feb25 0.1m |  | 1 | 0 | -30.28 |  | -30.076 |  |  | -30.3208 |  |  |  | -36.5316
114 | 23 | 2 | CP2 17feb25 0.1m |  | 1 | 0 | -30.137 |  | -30.079 |  |  | -30.3259 |  |  |  | -36.5365
115 | 23 | 3 | CP2 17feb25 0.1m |  | 1 | 0 | -30.335 |  | -30.3056 |  |  | -30.5548 |  |  |  | -36.7546
116 | 23 | 4 | CP2 17feb25 0.1m |  | 1 | 0 | -30.191 | 0.088624 | -30.1733 | 0.10802 |  | -30.4246 |  |  |  | -36.6306
117 | 24 | 1 | CC3 17feb25 BOT |  | 1 | 0 | -25.066 |  | -24.2683 |  |  | -24.5218 |  |  |  | -31.0043
118 | 24 | 2 | CC3 17feb25 BOT |  | 1 | 0 | -24.528 |  | -24.302 |  |  | -24.5577 |  |  |  | -31.0385
119 | 24 | 3 | CC3 17feb25 BOT |  | 1 | 0 | -23.994 |  | -23.8491 |  |  | -24.1069 |  |  |  | -30.6089
120 | 24 | 4 | CC3 17feb25 BOT |  | 1 | 0 | -24.411 | 0.44162 | -24.338 | 0.22861 |  | -24.598 |  |  |  | -31.0769
121 | 25 | 1 | CS1 15aug24 0.1m - DUP |  | 1 | 0 | -30.524 |  | -31.4754 |  |  | -31.7376 |  |  |  | -37.882
122 | 25 | 2 | CS1 15aug24 0.1m - DUP |  | 1 | 0 | -30.945 |  | -31.2057 |  |  | -31.47 |  |  |  | -37.627
123 | 25 | 3 | CS1 15aug24 0.1m - DUP |  | 1 | 0 | -31.158 |  | -31.3157 |  |  | -31.5822 |  |  |  | -37.7339
124 | 25 | 4 | CS1 15aug24 0.1m - DUP |  | 1 | 0 | -31.01 | 0.27186 | -31.0933 | 0.1629 |  | -31.362 |  |  |  | -37.524
125 | 26 | 1 | Hawaii |  | 1 | 0 | -8.828 |  | -5.3755 |  |  | -5.6463 |  |  |  | -13.0134
126 | 26 | 2 | Hawaii |  | 1 | 0 | -6.491 |  | -5.5126 |  |  | -5.7855 |  |  |  | -13.1461
127 | 26 | 3 | Hawaii |  | 1 | 0 | -5.816 |  | -5.227 |  |  | -5.5021 |  |  |  | -12.8759
128 | 26 | 4 | Hawaii |  | 1 | 0 | -5.981 | 1.3959 | -5.6649 | 0.18733 |  | -5.9422 |  |  |  | -13.2954
129 | 27 | 1 | Blacksburg | Drift | 1 | 0 | -37.189 |  | -42.0463 |  | -42.0463 | -42.3258 | -42.3258 |  |  | -47.974
130 | 27 | 2 | Blacksburg | Drift | 1 | 0 | -40.283 |  | -41.6518 |  | -41.6518 | -41.9335 | -41.9335 |  |  | -47.6001
131 | 27 | 3 | Blacksburg | Drift | 1 | 0 | -41.103 |  | -41.9242 |  | -41.9242 | -42.208 | -42.208 |  |  | -47.8617
132 | 27 | 4 | Blacksburg | Drift | 1 | 0 | -41.279 | 1.8999 | -41.7248 | 0.181 | -41.7248 | -42.0108 | -42.0108 |  |  | -47.6738
Note on G2: Manually change first six injections to ignore "-1".
Note on G3: Manually change first six injections to ignore "-1".
Note on G4: Manually change first six injections to ignore "-1".
Note on G5: Manually change first six injections to ignore "-1".
Note on G6: Manually change first six injections to ignore "-1".
Note on G7: Manually change first six injections to ignore "-1".

## RESULTS
| Analyst: | DWH |  |  |  | Complete Sequence |  |  | Run Summary | 18O | 18O | 2H | 2H
| Analysis Date: | 2025-03-14 |  |  |  | Line | Inj Number | Identifier | Ignore | Ignore | d18O Final Value | d2H Final Value | Ignore
| Project: | CCR |  |  |  | 1 | 1 | Blacksburg | 0 | -1 |  |  | -1
[6 empty] | 2 | 2 | Blacksburg | 0 | -1 |  |  | -1
[6 empty] | 3 | 3 | Blacksburg | 0 | -1 |  |  | -1
Vial | Sample Name | d18O | 1 s.d. | d2H | 1 s.d. | 4 | 4 | Blacksburg | 0 | -1 |  |  | -1
|  | ‰ vs. VSMOW |  | ‰ vs. VSMOW |  | 5 | 5 | Blacksburg | 0 | -1 |  |  | -1
1 | Blacksburg | -7.4776 | 0.052715 | -47.8419 | 0.1248 | 6 | 6 | Blacksburg | 0 | -1 |  |  | -1
2 | Myrtle | -2.3319 | 0.048272 | -10.2993 | 0.11454 | 7 | 7 | Blacksburg | 0 | 0 | -7.4081 | -48.0049 | 0
3 | Homer | -14.569 | 0.0444 | -110.4248 | 0.083274 | 8 | 8 | Blacksburg | 0 | 0 | -7.5178 | -47.8392 | 0
4 | Blacksburg | -7.4913 | 0.04403 | -47.6917 | 0.099553 | 9 | 9 | Blacksburg | 0 | 0 | -7.5193 | -47.8222 | 0
5 | Hawaii | -3.6972 | 0.056619 | -12.7 | 0.26118 | 10 | 10 | Blacksburg | 0 | 0 | -7.4651 | -47.7013 | 0
6 | CC4 17feb25 BOT | -5.2947 | 0.021833 | -30.7589 | 0.31402 | 11 | 1 | Myrtle | 0 | 0 | -2.3717 | -10.3728 | 0
7 | C50 17feb25 20m | -5.23 | 0.059806 | -30.5004 | 0.15723 | 12 | 2 | Myrtle | 0 | 0 | -2.3498 | -10.246 | 0
8 | C50 17feb25 6m | -5.298 | 0.047427 | -30.5022 | 0.2374 | 13 | 3 | Myrtle | 0 | 0 | -2.322 | -10.2151 | 0
9 | Blacksburg | -7.6864 | 0.012725 | -47.1091 | 0.11588 | 14 | 4 | Myrtle | 0 | 0 | -2.3276 | -10.3447 | 0
10 | CC4 17feb25 1.5m | -5.3091 | 0.04963 | -30.4292 | 0.095084 | 15 | 5 | Myrtle | 0 | 0 | -2.445 | -10.1391 | 0
11 | CC3 15aug24 0.1m | -4.1633 | 0.065504 | -25.8282 | 0.055962 | 16 | 6 | Myrtle | 0 | 0 | -2.2976 | -10.3451 | 0
12 | CP1 17feb25 BOT | -6.4063 | 0.058578 | -35.5662 | 0.086522 | 17 | 7 | Myrtle | 0 | 0 | -2.3171 | -10.1542 | 0
13 | CC5 22may24 13m | -5.4748 | 0.054914 | -32.1129 | 0.14634 | 18 | 8 | Myrtle | 0 | 0 | -2.3198 | -10.3685 | 0
14 | CC2 22may24 BOT | -5.1916 | 0.046789 | -30.9963 | 0.10054 | 19 | 9 | Myrtle | 0 | 0 | -2.2894 | -10.2914 | 0
15 | CP2 17feb25 BOT | -5.1952 | 0.095994 | -31.1897 | 0.29604 | 20 | 10 | Myrtle | 0 | 0 | -2.2788 | -10.5162 | 0
16 | CS1 15aug24 0.1m | -6.3569 | 0.015854 | -37.5941 | 0.25583 | 21 | 1 | Homer | 0 | 0 | -14.544 | -110.4926 | 0
17 | CC2 17feb25 BOT | -5.127 | 0.062894 | -30.9523 | 0.1674 | 22 | 2 | Homer | 0 | 0 | -14.5678 | -110.4208 | 0
18 | Blacksburg | -7.5416 | 0.03879 | -47.5245 | 0.028033 | 23 | 3 | Homer | 0 | 0 | -14.5678 | -110.3989 | 0
19 | CS1 17feb25 0.1m | -6.6387 | 0.041014 | -37.5349 | 0.18053 | 24 | 4 | Homer | 0 | 0 | -14.6772 | -110.2909 | 0
20 | CC2 17feb25 0.1m | -5.1607 | 0.030906 | -31.0539 | 0.24432 | 25 | 5 | Homer | 0 | 0 | -14.5916 | -110.3691 | 0
21 | CS1 22may24 0.1m | -6.8445 | 0.052704 | -38.9191 | 0.18847 | 26 | 6 | Homer | 0 | 0 | -14.5148 | -110.3948 | 0
22 | CP1 17feb25 0.1m | -6.6452 | 0.048409 | -37.6246 | 0.25777 | 27 | 7 | Homer | 0 | 0 | -14.5491 | -110.3707 | 0
23 | CP2 17feb25 0.1m | -6.4009 | 0.050736 | -36.6133 | 0.10463 | 28 | 8 | Homer | 0 | 0 | -14.5756 | -110.5042 | 0
24 | CC3 17feb25 BOT | -5.1272 | 0.067369 | -30.9322 | 0.21755 | 29 | 9 | Homer | 0 | 0 | -14.5722 | -110.4204 | 0
25 | CS1 15aug24 0.1m - DUP | -6.3086 | 0.04427 | -37.6917 | 0.15308 | 30 | 10 | Homer | 0 | 0 | -14.5301 | -110.5859 | 0
26 | Hawaii | -3.5862 | 0.092474 | -13.0827 | 0.17964 | 31 | 1 | Blacksburg | 0 | 0 | -7.4197 | -47.7558 | 0
27 | Blacksburg | -7.5062 | 0.027104 | -47.7774 | 0.17121 | 32 | 2 | Blacksburg | 0 | 0 | -7.4762 | -47.7169 | 0
[6 empty] | 33 | 3 | Blacksburg | 0 | 0 | -7.4965 | -47.7007 | 0
[6 empty] | 34 | 4 | Blacksburg | 0 | 0 | -7.5548 | -47.4532 | 0
[6 empty] | 35 | 5 | Blacksburg | 0 | 0 | -7.5118 | -47.6989 | 0
| Quality Assurance Verification |  |  |  |  | 36 | 6 | Blacksburg | 0 | 0 | -7.491 | -47.6168 | 0
| Metric |  | Reviewer 1 | Reviewer 2 |  | 37 | 7 | Blacksburg | 0 | 0 | -7.4675 | -47.6925 | 0
| d18O Cal Curve R^2 | 0.999+ |  |  |  | 38 | 8 | Blacksburg | 0 | 0 | -7.4752 | -47.7722 | 0
| d18O QC checks | < 0.15 |  |  |  | 39 | 9 | Blacksburg | 0 | 0 | -7.5652 | -47.6941 | 0
| dD Cal Curve R^2 | 0.999+ |  |  |  | 40 | 10 | Blacksburg | 0 | 0 | -7.455 | -47.8157 | 0
| dD QC checks | < 1.5 |  |  |  | 41 | 1 | Hawaii | 0 | 0 | -3.6387 | -13.014 | 0
| Initials |  |  |  |  | 42 | 2 | Hawaii | 0 | 0 | -3.6951 | -12.7867 | 0
| Date |  |  |  |  | 43 | 3 | Hawaii | 0 | 0 | -3.7742 | -12.4042 | 0
| Comments: |  |  |  |  | 44 | 4 | Hawaii | 0 | 0 | -3.6809 | -12.595 | 0
[6 empty] | 45 | 1 | CC4 17feb25 BOT | 0 | 0 | -5.2918 | -31.1702 | 0
[6 empty] | 46 | 2 | CC4 17feb25 BOT | 0 | 0 | -5.3017 | -30.5198 | 0
[6 empty] | 47 | 3 | CC4 17feb25 BOT | 0 | 0 | -5.2664 | -30.8381 | 0
[6 empty] | 48 | 4 | CC4 17feb25 BOT | 0 | 0 | -5.3187 | -30.5075 | 0
[6 empty] | 49 | 1 | C50 17feb25 20m | 0 | 0 | -5.1708 | -30.5792 | 0
[6 empty] | 50 | 2 | C50 17feb25 20m | 0 | 0 | -5.1862 | -30.608 | 0
[6 empty] | 51 | 3 | C50 17feb25 20m | 0 | 0 | -5.2798 | -30.2676 | 0
[6 empty] | 52 | 4 | C50 17feb25 20m | 0 | 0 | -5.2832 | -30.5469 | 0
[6 empty] | 53 | 1 | C50 17feb25 6m | 0 | 0 | -5.3326 | -30.6439 | 0
[6 empty] | 54 | 2 | C50 17feb25 6m | 0 | 0 | -5.2577 | -30.206 | 0
[6 empty] | 55 | 3 | C50 17feb25 6m | 0 | 0 | -5.2567 | -30.4223 | 0
[6 empty] | 56 | 4 | C50 17feb25 6m | 0 | 0 | -5.3451 | -30.7364 | 0
[6 empty] | 57 | 1 | Blacksburg | 0 | 0 | -7.7039 | -46.9512 | 0
[6 empty] | 58 | 2 | Blacksburg | 0 | 0 | -7.681 | -47.1214 | 0
[6 empty] | 59 | 3 | Blacksburg | 0 | 0 | -7.6868 | -47.2299 | 0
[6 empty] | 60 | 4 | Blacksburg | 0 | 0 | -7.6741 | -47.134 | 0
[6 empty] | 61 | 1 | CC4 17feb25 1.5m | 0 | 0 | -5.3797 | -30.3069 | 0
[6 empty] | 62 | 2 | CC4 17feb25 1.5m | 0 | 0 | -5.2662 | -30.4515 | 0
[6 empty] | 63 | 3 | CC4 17feb25 1.5m | 0 | 0 | -5.2856 | -30.5369 | 0
[6 empty] | 64 | 4 | CC4 17feb25 1.5m | 0 | 0 | -5.305 | -30.4217 | 0
[6 empty] | 65 | 1 | CC3 15aug24 0.1m | 0 | 0 | -4.216 | -25.8482 | 0
[6 empty] | 66 | 2 | CC3 15aug24 0.1m | 0 | 0 | -4.092 | -25.7652 | 0
[6 empty] | 67 | 3 | CC3 15aug24 0.1m | 0 | 0 | -4.1235 | -25.8046 | 0
[6 empty] | 68 | 4 | CC3 15aug24 0.1m | 0 | 0 | -4.2218 | -25.895 | 0
[6 empty] | 69 | 1 | CP1 17feb25 BOT | 0 | 0 | -6.4738 | -35.5083 | 0
[6 empty] | 70 | 2 | CP1 17feb25 BOT | 0 | 0 | -6.3339 | -35.4805 | 0
[6 empty] | 71 | 3 | CP1 17feb25 BOT | 0 | 0 | -6.4243 | -35.6648 | 0
[6 empty] | 72 | 4 | CP1 17feb25 BOT | 0 | 0 | -6.3932 | -35.6113 | 0
[6 empty] | 73 | 1 | CC5 22may24 13m | 0 | 0 | -5.4438 | -31.9649 | 0
[6 empty] | 74 | 2 | CC5 22may24 13m | 0 | 0 | -5.4406 | -32.0854 | 0
[6 empty] | 75 | 3 | CC5 22may24 13m | 0 | 0 | -5.5563 | -32.0861 | 0
[6 empty] | 76 | 4 | CC5 22may24 13m | 0 | 0 | -5.4584 | -32.3151 | 0
[6 empty] | 77 | 1 | CC2 22may24 BOT | 0 | 0 | -5.2212 | -30.9448 | 0
[6 empty] | 78 | 2 | CC2 22may24 BOT | 0 | 0 | -5.1221 | -30.9178 | 0
[6 empty] | 79 | 3 | CC2 22may24 BOT | 0 | 0 | -5.2055 | -30.9808 | 0
[6 empty] | 80 | 4 | CC2 22may24 BOT | 0 | 0 | -5.2176 | -31.1421 | 0
[6 empty] | 81 | 1 | CP2 17feb25 BOT | 0 | 0 | -5.3209 | -30.9105 | 0
[6 empty] | 82 | 2 | CP2 17feb25 BOT | 0 | 0 | -5.2147 | -31.2913 | 0
[6 empty] | 83 | 3 | CP2 17feb25 BOT | 0 | 0 | -5.1447 | -30.9967 | 0
[6 empty] | 84 | 4 | CP2 17feb25 BOT | 0 | 0 | -5.1007 | -31.5604 | 0
[6 empty] | 85 | 1 | CS1 15aug24 0.1m | 0 | 0 | -6.3709 | -37.2676 | 0
[6 empty] | 86 | 2 | CS1 15aug24 0.1m | 0 | 0 | -6.3644 | -37.8763 | 0
[6 empty] | 87 | 3 | CS1 15aug24 0.1m | 0 | 0 | -6.3576 | -37.5488 | 0
[6 empty] | 88 | 4 | CS1 15aug24 0.1m | 0 | 0 | -6.3345 | -37.6839 | 0
[6 empty] | 89 | 1 | CC2 17feb25 BOT | 0 | 0 | -5.0685 | -30.7033 | 0
[6 empty] | 90 | 2 | CC2 17feb25 BOT | 0 | 0 | -5.1068 | -31.0174 | 0
[6 empty] | 91 | 3 | CC2 17feb25 BOT | 0 | 0 | -5.1168 | -31.0657 | 0
[6 empty] | 92 | 4 | CC2 17feb25 BOT | 0 | 0 | -5.2161 | -31.0227 | 0
[6 empty] | 93 | 1 | Blacksburg | 0 | 0 | -7.5315 | -47.5026 | 0
[6 empty] | 94 | 2 | Blacksburg | 0 | 0 | -7.599 | -47.4991 | 0
[6 empty] | 95 | 3 | Blacksburg | 0 | 0 | -7.5171 | -47.5558 | 0
[6 empty] | 96 | 4 | Blacksburg | 0 | 0 | -7.5189 | -47.5405 | 0
[6 empty] | 97 | 1 | CS1 17feb25 0.1m | 0 | 0 | -6.6058 | -37.7382 | 0
[6 empty] | 98 | 2 | CS1 17feb25 0.1m | 0 | 0 | -6.6736 | -37.3328 | 0
[6 empty] | 99 | 3 | CS1 17feb25 0.1m | 0 | 0 | -6.6007 | -37.6227 | 0
[6 empty] | 100 | 4 | CS1 17feb25 0.1m | 0 | 0 | -6.6748 | -37.446 | 0
[6 empty] | 101 | 1 | CC2 17feb25 0.1m | 0 | 0 | -5.1234 | -30.7874 | 0
[6 empty] | 102 | 2 | CC2 17feb25 0.1m | 0 | 0 | -5.1843 | -31.1497 | 0
[6 empty] | 103 | 3 | CC2 17feb25 0.1m | 0 | 0 | -5.1473 | -30.9339 | 0
[6 empty] | 104 | 4 | CC2 17feb25 0.1m | 0 | 0 | -5.1878 | -31.3446 | 0
[6 empty] | 105 | 1 | CS1 22may24 0.1m | 0 | 0 | -6.7696 | -38.7999 | 0
[6 empty] | 106 | 2 | CS1 22may24 0.1m | 0 | 0 | -6.8933 | -38.7185 | 0
[6 empty] | 107 | 3 | CS1 22may24 0.1m | 0 | 0 | -6.859 | -39.0573 | 0
[6 empty] | 108 | 4 | CS1 22may24 0.1m | 0 | 0 | -6.8561 | -39.1008 | 0
[6 empty] | 109 | 1 | CP1 17feb25 0.1m | 0 | 0 | -6.5966 | -37.6887 | 0
[6 empty] | 110 | 2 | CP1 17feb25 0.1m | 0 | 0 | -6.7055 | -37.3897 | 0
[6 empty] | 111 | 3 | CP1 17feb25 0.1m | 0 | 0 | -6.6612 | -37.4595 | 0
[6 empty] | 112 | 4 | CP1 17feb25 0.1m | 0 | 0 | -6.6173 | -37.9604 | 0
[6 empty] | 113 | 1 | CP2 17feb25 0.1m | 0 | 0 | -6.339 | -36.5316 | 0
[6 empty] | 114 | 2 | CP2 17feb25 0.1m | 0 | 0 | -6.4171 | -36.5365 | 0
[6 empty] | 115 | 3 | CP2 17feb25 0.1m | 0 | 0 | -6.3876 | -36.7546 | 0
[6 empty] | 116 | 4 | CP2 17feb25 0.1m | 0 | 0 | -6.4596 | -36.6306 | 0
[6 empty] | 117 | 1 | CC3 17feb25 BOT | 0 | 0 | -5.0742 | -31.0043 | 0
[6 empty] | 118 | 2 | CC3 17feb25 BOT | 0 | 0 | -5.2237 | -31.0385 | 0
[6 empty] | 119 | 3 | CC3 17feb25 BOT | 0 | 0 | -5.089 | -30.6089 | 0
[6 empty] | 120 | 4 | CC3 17feb25 BOT | 0 | 0 | -5.1217 | -31.0769 | 0
[6 empty] | 121 | 1 | CS1 15aug24 0.1m - DUP | 0 | 0 | -6.3417 | -37.882 | 0
[6 empty] | 122 | 2 | CS1 15aug24 0.1m - DUP | 0 | 0 | -6.2467 | -37.627 | 0
[6 empty] | 123 | 3 | CS1 15aug24 0.1m - DUP | 0 | 0 | -6.3065 | -37.7339 | 0
[6 empty] | 124 | 4 | CS1 15aug24 0.1m - DUP | 0 | 0 | -6.3395 | -37.524 | 0
[6 empty] | 125 | 1 | Hawaii | 0 | 0 | -3.6189 | -13.0134 | 0
[6 empty] | 126 | 2 | Hawaii | 0 | 0 | -3.5271 | -13.1461 | 0
[6 empty] | 127 | 3 | Hawaii | 0 | 0 | -3.7012 | -12.8759 | 0
[6 empty] | 128 | 4 | Hawaii | 0 | 0 | -3.4975 | -13.2954 | 0
[6 empty] | 129 | 1 | Blacksburg | 0 | 0 | -7.473 | -47.974 | 0
[6 empty] | 130 | 2 | Blacksburg | 0 | 0 | -7.5377 | -47.6001 | 0
[6 empty] | 131 | 3 | Blacksburg | 0 | 0 | -7.5147 | -47.8617 | 0
[6 empty] | 132 | 4 | Blacksburg | 0 | 0 | -7.4996 | -47.6738 | 0