## Instructions
Picarro Isotope Analyzer
Data Processor 3.0 (KAP)
How to Use this Spreadsheet:
Step 1: Retrieve raw result file from Piccaro CPU (C:Piccaro\IsotopeData). Add initials and project name to end
of file name. Copy file into your folder on Water Server.
Step 2: On syncing computer, move raw file into Raw Data folder on Google Drive:
Instruments and Analyses --> Picarro Water Isotope Analyzer --> Raw Data
Step 3: Save this spreadsheet in the Processed Data folder on Google Drive:
Instruments and Analyses --> Picarro Water Isotope Analyzer --> Processed Data
Use file name: YYYYMMDD_Initials_(anything else to identify ex. project, run number, etc.)
Step 4: Enter run information and sample positions in the "Tray Configuration" tab. (orange cells)
(It is important to use the same names for the standards throughout this spreadsheet)
Step 5: Copy text from raw result file and paste into "Picarro Output" tab.
Step 6: In Run Summary tab, check that H2O Mean is within "good" bounds. (between 17,000 - 23,000)
Step 7: In 18O tab, verify that the reference defined values match the references used in run (cells U39 - 42). Run Solver and
verify stdev memory corrected are reduced (column K values less than corresponding column I values). Verify stdev of
drift corrected reference is lower than memory corrected. Verify normalization R2 is 0.99+. Verify QC checks are within
0.15 tolerance.
Step 8: In dD tab, verify that the reference defined values match the references used in run (cells U39 - 42. Run Solver and
verify stdev memory corrected are reduced (column K values less than corresponding column I values). Verify stdev of
drift corrected reference is lower than memory corrected. Verify normalization R2 is 0.99+. Verify QC checks are within
1.5 tolerance.
Step 9: Fill in Quality Assurance Verification chart on RESULTS tab. Dr. Scott or Kelly Peeler will fill in as reviewer 2.
Step 10: Add reference results to Water Isotope Analyzer run log.

## Tray Configuration
| Analyst |  | DWH
| PI |  | Durelle Scott
| Project |  | CCR
| Sampling Date |  | 2024 - 2025
| Sampling Location |  | Carvins Cove Reservoir
| Analysis Date |  | 2025-03-18
Notes
Rear Rack
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
1 | Blacksburg | Myrtle | Homer | Blacksburg | Hawaii | CC4 30sep24 BOT | CS1 28oct24 0.1m | CC2 28oct24 BOT | Blacksburg | 9
10 | CC4 30sep24 1.5m | CC2 30sep24 BOT | CCR rain 30sep24 | CC4 30sep24 0.1m | CC3 30sep24 BOT | CC4 16sep24 0.1m | C50 17feb25 9m | CC3 16sep24 1.5m | Blacksburg | 18
19 | CC3 30sep24 0.1m | CC3 16sep24 BOT | CC3 17feb25 0.1m | CC3 16sep24 0.1m | CS2 17feb25 0.1m | C50 17feb25 0.1m | C50 17feb25 9m - DUP | Hawaii | Blacksburg | 27
28 | [9 empty] | 36
37 | [9 empty] | 45
46 | [9 empty] | 54
Rear Rack
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
1 | Drift | High reference | Low Reference | Drift | QC Check | Sample | Sample | Sample | Drift | 9
10 | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Drift | 18
19 | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Drift | 27
28 | [9 empty] | 36
37 | [9 empty] | 45
46 | [9 empty] | 54
Standards
|  | Standard ID | Actual d(18_16) | Actual d(D_H)
| #1 | Denali | -19.9287 | -153.7172 | [6 empty] | -3.37
| #2 | Alberta | -18.8392 | -146.8245 | [6 empty] | -20.92
| #3 | MT BHB | -18.431 | -140.5778 | [6 empty] | -17.35
| #4 | Anchorage | -16.4204 | -129.9028 | [6 empty] | -7.39
| #5 | Homer | -14.5399 | -110.5169
| #6 | MT Creek | -13.3562 | -97.7428
| #7 | Nickel Mines | -7.876 | -48.7951
| #8 | Blacksburg | -7.5615 | -47.446
| #9 | Denton's Well | -6.9583 | -42.6011
| #10 | Hawaii | -3.6108 | -13.2294
| #11 | Myrtle Beach | -2.2908 | -10.453
| #12 | Heavy Oxygen | 3.1753 | -15.7881
| #13 | Heavy Hydrogen | 11.1372 | 4.0332
| #14 | Andres Heavy Water | 17.5664 | 20.5382
| #15 | Picarro ZERO | 0.3 | 1.8
| #16 | Picarro MID | -20.6 | -159
| #17 | Picarro DEPL | -29.6 | -235
#18
#19
#20
|  | DI | -6 | -36
Note on C34: Nate: The Standard ID must match the Standard ID used above in the trey protocol. For example, if you are using the Boulder Colorado Standard, and you labeled it "B" above, it must be labled "B" in this collum.

## Timing
| Enter start time and date --> | 2019-01-10 09:30
Tray Position | Sample ID | Start Time | # Injections
1 | Blacksburg | 2019-01-10 10:58 | 10
2 | Myrtle | 2019-01-10 12:27 | 10
3 | Homer | 2019-01-10 13:55 | 10
4 | Blacksburg | 2019-01-10 15:23 | 10
5 | Hawaii | 2019-01-10 15:59 | 4
6 | CC4 30sep24 BOT | 2019-01-10 16:34 | 4
7 | CS1 28oct24 0.1m | 2019-01-10 17:09 | 4
8 | CC2 28oct24 BOT | 2019-01-10 17:45 | 4
9 | Blacksburg | 2019-01-10 18:20 | 4
10 | CC4 30sep24 1.5m | 2019-01-10 18:55 | 4
11 | CC2 30sep24 BOT | 2019-01-10 19:31 | 4
12 | CCR rain 30sep24 | 2019-01-10 20:06 | 4
13 | CC4 30sep24 0.1m | 2019-01-10 20:41 | 4
14 | CC3 30sep24 BOT | 2019-01-10 21:17 | 4
15 | CC4 16sep24 0.1m | 2019-01-10 21:52 | 4
16 | C50 17feb25 9m | 2019-01-10 22:27 | 4
17 | CC3 16sep24 1.5m | 2019-01-10 23:03 | 4
18 | Blacksburg | 2019-01-10 23:38 | 4
19 | CC3 30sep24 0.1m | 2019-01-11 00:13 | 4
20 | CC3 16sep24 BOT | 2019-01-11 00:49 | 4
21 | CC3 17feb25 0.1m | 2019-01-11 01:24 | 4
22 | CC3 16sep24 0.1m | 2019-01-11 01:59 | 4
23 | CS2 17feb25 0.1m | 2019-01-11 02:35 | 4
24 | C50 17feb25 0.1m | 2019-01-11 03:10 | 4
25 | C50 17feb25 9m - DUP | 2019-01-11 03:45 | 4
26 | Hawaii | 2019-01-11 04:21 | 4
27 | Blacksburg | 2019-01-11 04:56 | 4
Formulas:
Start Time = [Start Time, row above]+"0:08:50"*10

## Picarro Output
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | days_Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | BASELINE_Mean | BASELINE_SD | DAS Temp | Tray | Sample | Job | Method
1 | P-11620 | 2025/03/18 09:22:35 | MT1-Rear-01 | 1 | -12.126 | -44.641 | 19394 | -1 | 1 |  |  | H2O | 67.4091 | 0.174 | 0.369 | 79 | -0.043 | 0.003 | 25.99 | 844 | 0 | 30.563 | MT1-Rear | 1 | 66 | BSE
2 | P-11620 | 2025/03/18 09:31:22 | MT1-Rear-01 | 2 | -11.886 | -40.155 | 19697 | -1 | 1 |  |  | H2O | 67.4152 | 0.196 | 0.41 | 80 | -0.038 | 0.001 | 26.51 | 844 | 0 | 29.75 | MT1-Rear | 1 | 66 | BSE
3 | P-11620 | 2025/03/18 09:40:08 | MT1-Rear-01 | 3 | -11.941 | -39.673 | 19665 | -1 | 1 |  |  | H2O | 67.4213 | 0.139 | 0.469 | 82 | -0.005 | -0.032 | 26.96 | 844 | 0 | 29.938 | MT1-Rear | 1 | 66 | BSE
4 | P-11620 | 2025/03/18 09:48:54 | MT1-Rear-01 | 4 | -11.861 | -39.483 | 19730 | 0 | 1 |  |  | H2O | 67.4274 | 0.179 | 0.299 | 82 | 0.007 | 0.026 | 27.04 | 844 | 0 | 30.063 | MT1-Rear | 1 | 66 | BSE
5 | P-11620 | 2025/03/18 09:57:40 | MT1-Rear-01 | 5 | -11.832 | -39.619 | 19645 | 0 | 1 |  |  | H2O | 67.4335 | 0.176 | 0.472 | 82 | 0.008 | 0.038 | 27.13 | 844 | 0 | 29.813 | MT1-Rear | 1 | 66 | BSE
6 | P-11620 | 2025/03/18 10:06:26 | MT1-Rear-01 | 6 | -11.779 | -39.477 | 19726 | 0 | 1 |  |  | H2O | 67.4396 | 0.152 | 0.417 | 85 | -0.012 | -0.088 | 28.18 | 844 | 0 | 29.75 | MT1-Rear | 1 | 66 | BSE
7 | P-11620 | 2025/03/18 10:15:13 | MT1-Rear-01 | 7 | -11.738 | -39.419 | 19707 | 0 | 1 |  |  | H2O | 67.4457 | 0.116 | 0.37 | 83 | -0.003 | 0.026 | 27.34 | 844 | 0 | 30.125 | MT1-Rear | 1 | 66 | BSE
8 | P-11620 | 2025/03/18 10:23:58 | MT1-Rear-01 | 8 | -11.762 | -39.291 | 19506 | 0 | 1 |  |  | H2O | 67.4517 | 0.119 | 0.477 | 91 | -0.009 | 0.026 | 27.49 | 844 | 0 | 30.125 | MT1-Rear | 1 | 66 | BSE
9 | P-11620 | 2025/03/18 10:32:44 | MT1-Rear-01 | 9 | -11.728 | -39.3 | 19482 | 0 | 1 |  |  | H2O | 67.4578 | 0.168 | 0.495 | 92 | -0.009 | -0.033 | 27.76 | 844 | 0 | 29.875 | MT1-Rear | 1 | 66 | BSE
10 | P-11620 | 2025/03/18 10:41:30 | MT1-Rear-01 | 10 | -11.704 | -39.214 | 19464 | 0 | 1 |  |  | H2O | 67.4639 | 0.161 | 0.179 | 83 | 0.019 | -0.007 | 27.46 | 844 | 0 | 29.875 | MT1-Rear | 1 | 66 | BSE
11 | P-11621 | 2025/03/18 10:50:16 | MT1-Rear-02 | 1 | -6.799 | -5.618 | 19736 | -1 | 1 |  |  | H2O | 67.47 | 0.164 | 0.367 | 90 | -0.038 | -0.058 | 27.08 | 844 | 0 | 30.125 | MT1-Rear | 2 | 66 | BSE
12 | P-11621 | 2025/03/18 10:59:02 | MT1-Rear-02 | 2 | -6.518 | -1.901 | 19672 | -1 | 1 |  |  | H2O | 67.4761 | 0.211 | 0.305 | 95 | -0.05 | -0.045 | 28.7 | 844 | 0 | 30.375 | MT1-Rear | 2 | 66 | BSE
13 | P-11621 | 2025/03/18 11:07:49 | MT1-Rear-02 | 3 | -6.484 | -0.952 | 19650 | -1 | 1 |  |  | H2O | 67.4822 | 0.206 | 0.461 | 93 | -0.045 | -0.07 | 28.14 | 844 | 0 | 29.688 | MT1-Rear | 2 | 66 | BSE
14 | P-11621 | 2025/03/18 11:16:36 | MT1-Rear-02 | 4 | -6.364 | -0.878 | 19676 | 0 | 1 |  |  | H2O | 67.4883 | 0.129 | 0.348 | 93 | 0.002 | -0.051 | 27.95 | 844 | 0 | 30 | MT1-Rear | 2 | 66 | BSE
15 | P-11621 | 2025/03/18 11:25:22 | MT1-Rear-02 | 5 | -6.353 | -0.407 | 19631 | 0 | 1 |  |  | H2O | 67.4943 | 0.136 | 0.277 | 85 | -0.003 | 0.016 | 28.05 | 844 | 0 | 30.25 | MT1-Rear | 2 | 66 | BSE
16 | P-11621 | 2025/03/18 11:34:08 | MT1-Rear-02 | 6 | -6.367 | -0.221 | 19645 | 0 | 1 |  |  | H2O | 67.5004 | 0.217 | 0.399 | 85 | -0.035 | -0.013 | 28.1 | 844 | 0 | 29.688 | MT1-Rear | 2 | 66 | BSE
17 | P-11621 | 2025/03/18 11:42:55 | MT1-Rear-02 | 7 | -6.305 | 0.036 | 19789 | 0 | 1 |  |  | H2O | 67.5065 | 0.15 | 0.42 | 83 | 0.001 | -0.06 | 27.23 | 844 | 0 | 29.688 | MT1-Rear | 2 | 66 | BSE
18 | P-11621 | 2025/03/18 11:51:42 | MT1-Rear-02 | 8 | -6.307 | -0.114 | 19619 | 0 | 1 |  |  | H2O | 67.5126 | 0.118 | 0.54 | 83 | 0.003 | -0.096 | 27.47 | 844 | 0 | 29.563 | MT1-Rear | 2 | 66 | BSE
19 | P-11621 | 2025/03/18 12:00:28 | MT1-Rear-02 | 9 | -6.288 | -0.081 | 19629 | 0 | 1 |  |  | H2O | 67.5187 | 0.109 | 0.35 | 86 | -0.005 | -0.035 | 28.34 | 844 | 0 | 30.313 | MT1-Rear | 2 | 66 | BSE
20 | P-11621 | 2025/03/18 12:09:15 | MT1-Rear-02 | 10 | -6.319 | 0.024 | 19607 | 0 | 1 |  |  | H2O | 67.5248 | 0.191 | 0.248 | 83 | -0.036 | -0.043 | 27.48 | 844 | 0 | 29.875 | MT1-Rear | 2 | 66 | BSE
21 | P-11622 | 2025/03/18 12:18:02 | MT1-Rear-03 | 1 | -17.794 | -89.866 | 19728 | -1 | 1 |  |  | H2O | 67.5309 | 0.179 | 0.427 | 84 | -0.003 | 0.07 | 27.56 | 844 | 0 | 29.688 | MT1-Rear | 3 | 66 | BSE
22 | P-11622 | 2025/03/18 12:26:49 | MT1-Rear-03 | 2 | -18.781 | -100.037 | 19627 | -1 | 1 |  |  | H2O | 67.537 | 0.171 | 0.413 | 84 | 0.003 | 0.094 | 27.78 | 844 | 0 | 30.125 | MT1-Rear | 3 | 66 | BSE
23 | P-11622 | 2025/03/18 12:35:35 | MT1-Rear-03 | 3 | -18.842 | -102.607 | 19664 | -1 | 1 |  |  | H2O | 67.5431 | 0.166 | 0.471 | 84 | -0.003 | 0.073 | 27.79 | 844 | 0 | 30.125 | MT1-Rear | 3 | 66 | BSE
24 | P-11622 | 2025/03/18 12:44:21 | MT1-Rear-03 | 4 | -18.95 | -103.387 | 19743 | 0 | 1 |  |  | H2O | 67.5492 | 0.089 | 0.377 | 86 | -0.01 | 0.085 | 28.48 | 844 | 0 | 29.625 | MT1-Rear | 3 | 66 | BSE
25 | P-11622 | 2025/03/18 12:53:08 | MT1-Rear-03 | 5 | -19.008 | -103.884 | 19779 | 0 | 1 |  |  | H2O | 67.5553 | 0.183 | 0.512 | 95 | -0.03 | 0.055 | 28.48 | 844 | 0 | 29.125 | MT1-Rear | 3 | 66 | BSE
26 | P-11622 | 2025/03/18 13:01:56 | MT1-Rear-03 | 6 | -19.03 | -104.221 | 19744 | 0 | 1 |  |  | H2O | 67.5614 | 0.119 | 0.251 | 90 | 0.003 | 0.042 | 27.21 | 844 | 0 | 30.5 | MT1-Rear | 3 | 66 | BSE
27 | P-11622 | 2025/03/18 13:10:43 | MT1-Rear-03 | 7 | -19.064 | -104.099 | 19477 | 0 | 1 |  |  | H2O | 67.5675 | 0.147 | 0.361 | 91 | -0.005 | 0.051 | 27.3 | 843 | 0 | 29.875 | MT1-Rear | 3 | 66 | BSE
28 | P-11622 | 2025/03/18 13:19:30 | MT1-Rear-03 | 8 | -19.123 | -104.686 | 19778 | 0 | 1 |  |  | H2O | 67.5737 | 0.117 | 0.387 | 87 | -0.02 | 0.094 | 28.63 | 844 | 0 | 30.188 | MT1-Rear | 3 | 66 | BSE
29 | P-11622 | 2025/03/18 13:28:17 | MT1-Rear-03 | 9 | -19.069 | -104.564 | 19755 | 0 | 1 |  |  | H2O | 67.5798 | 0.138 | 0.383 | 84 | -0.012 | 0.049 | 27.59 | 844 | 0 | 30.063 | MT1-Rear | 3 | 66 | BSE
30 | P-11622 | 2025/03/18 13:37:04 | MT1-Rear-03 | 10 | -19.166 | -104.324 | 19774 | 0 | 1 |  |  | H2O | 67.5859 | 0.13 | 0.429 | 89 | -0.007 | 0.106 | 29.25 | 844 | 0 | 30.063 | MT1-Rear | 3 | 66 | BSE
31 | P-11623 | 2025/03/18 13:45:52 | MT1-Rear-04 | 1 | -12.472 | -47.313 | 19788 | -1 | 1 |  |  | H2O | 67.592 | 0.152 | 0.374 | 86 | -0.009 | 0.008 | 28.46 | 844 | 0 | 29.938 | MT1-Rear | 4 | 66 | BSE
32 | P-11623 | 2025/03/18 13:54:39 | MT1-Rear-04 | 2 | -12.045 | -41.765 | 19765 | -1 | 1 |  |  | H2O | 67.5981 | 0.164 | 0.469 | 85 | -0.041 | 0.044 | 28.22 | 844 | 0 | 30.188 | MT1-Rear | 4 | 66 | BSE
33 | P-11623 | 2025/03/18 14:03:26 | MT1-Rear-04 | 3 | -12.017 | -40.47 | 19724 | -1 | 1 |  |  | H2O | 67.6042 | 0.207 | 0.442 | 85 | -0.06 | 0.068 | 28.15 | 844 | 0 | 30.625 | MT1-Rear | 4 | 66 | BSE
34 | P-11623 | 2025/03/18 14:12:12 | MT1-Rear-04 | 4 | -11.969 | -39.508 | 19762 | 0 | 1 |  |  | H2O | 67.6103 | 0.149 | 0.339 | 88 | -0.033 | 0.009 | 28.92 | 844 | 0 | 29.688 | MT1-Rear | 4 | 66 | BSE
35 | P-11623 | 2025/03/18 14:20:59 | MT1-Rear-04 | 5 | -11.982 | -39.504 | 19762 | 0 | 1 |  |  | H2O | 67.6164 | 0.2 | 0.36 | 86 | -0.028 | 0.06 | 28.51 | 844 | 0 | 30.125 | MT1-Rear | 4 | 66 | BSE
36 | P-11623 | 2025/03/18 14:29:46 | MT1-Rear-04 | 6 | -11.915 | -39.33 | 19784 | 0 | 1 |  |  | H2O | 67.6224 | 0.12 | 0.285 | 94 | -0.003 | -0.028 | 28.41 | 844 | 0 | 30.188 | MT1-Rear | 4 | 66 | BSE
37 | P-11623 | 2025/03/18 14:38:33 | MT1-Rear-04 | 7 | -11.859 | -39.365 | 19698 | 0 | 1 |  |  | H2O | 67.6285 | 0.134 | 0.497 | 96 | -0.011 | 0.093 | 28.96 | 844 | 0 | 29.938 | MT1-Rear | 4 | 66 | BSE
38 | P-11623 | 2025/03/18 14:47:21 | MT1-Rear-04 | 8 | -12.005 | -39.03 | 20636 | 0 | 1 |  |  | H2O | 67.6346 | 0.17 | 0.231 | 81 | -0.019 | 0.048 | 29.61 | 844 | 0 | 30 | MT1-Rear | 4 | 66 | BSE
39 | P-11623 | 2025/03/18 14:56:08 | MT1-Rear-04 | 9 | -12.025 | -39.03 | 20305 | 0 | 1 |  |  | H2O | 67.6407 | 0.158 | 0.315 | 89 | -0.014 | 0.086 | 29.44 | 844 | 0 | 30.813 | MT1-Rear | 4 | 66 | BSE
40 | P-11623 | 2025/03/18 15:04:56 | MT1-Rear-04 | 10 | -11.902 | -39.267 | 19743 | 0 | 1 |  |  | H2O | 67.6468 | 0.092 | 0.401 | 85 | -0.018 | -0.016 | 28.11 | 844 | 0 | 30.188 | MT1-Rear | 4 | 66 | BSE
41 | P-11624 | 2025/03/18 15:13:43 | MT1-Rear-05 | 1 | -8.225 | -7.524 | 19743 | -1 | 1 |  |  | H2O | 67.6529 | 0.13 | 0.472 | 91 | 0.016 | 0.037 | 29.98 | 844 | 0 | 30.375 | MT1-Rear | 5 | 67 | BSE
42 | P-11624 | 2025/03/18 15:22:31 | MT1-Rear-05 | 2 | -7.972 | -3.784 | 19775 | -1 | 1 |  |  | H2O | 67.659 | 0.228 | 0.277 | 86 | -0.02 | 0.014 | 28.52 | 844 | 0 | 30.125 | MT1-Rear | 5 | 67 | BSE
43 | P-11624 | 2025/03/18 15:31:19 | MT1-Rear-05 | 3 | -8.024 | -3.107 | 19786 | -1 | 1 |  |  | H2O | 67.6651 | 0.147 | 0.279 | 89 | -0.034 | 0.011 | 29.27 | 844 | 0 | 30.438 | MT1-Rear | 5 | 67 | BSE
44 | P-11624 | 2025/03/18 15:40:07 | MT1-Rear-05 | 4 | -8.007 | -2.853 | 19769 | 0 | 1 |  |  | H2O | 67.6712 | 0.168 | 0.27 | 89 | -0.019 | 0.011 | 29.48 | 844 | 0 | 30.188 | MT1-Rear | 5 | 67 | BSE
45 | P-11625 | 2025/03/18 15:48:55 | MT1-Rear-06 | 1 | -9.347 | -19.206 | 19759 | -1 | 1 |  |  | H2O | 67.6774 | 0.215 | 0.397 | 90 | 0.006 | 0.03 | 29.85 | 844 | 0 | 30.125 | MT1-Rear | 6 | 67 | BSE
46 | P-11625 | 2025/03/18 15:57:42 | MT1-Rear-06 | 2 | -9.467 | -21.041 | 19796 | -1 | 1 |  |  | H2O | 67.6835 | 0.165 | 0.442 | 85 | -0.017 | -0.119 | 28.17 | 844 | 0 | 30.5 | MT1-Rear | 6 | 67 | BSE
47 | P-11625 | 2025/03/18 16:06:30 | MT1-Rear-06 | 3 | -9.438 | -21.457 | 19816 | -1 | 1 |  |  | H2O | 67.6896 | 0.115 | 0.298 | 89 | -0.016 | 0.034 | 29.35 | 844 | 0 | 30.563 | MT1-Rear | 6 | 67 | BSE
48 | P-11625 | 2025/03/18 16:15:17 | MT1-Rear-06 | 4 | -9.604 | -21.152 | 19749 | 0 | 1 |  |  | H2O | 67.6957 | 0.127 | 0.169 | 88 | -0.014 | 0.028 | 29 | 844 | 0 | 30.5 | MT1-Rear | 6 | 67 | BSE
49 | P-11626 | 2025/03/18 16:24:05 | MT1-Rear-07 | 1 | -10.505 | -24.933 | 19815 | -1 | 1 |  |  | H2O | 67.7018 | 0.133 | 0.489 | 89 | -0.016 | 0.114 | 29.43 | 844 | 0 | 30.375 | MT1-Rear | 7 | 67 | BSE
50 | P-11626 | 2025/03/18 16:32:53 | MT1-Rear-07 | 2 | -10.537 | -25.204 | 19821 | -1 | 1 |  |  | H2O | 67.7079 | 0.119 | 0.472 | 89 | -0.016 | 0.121 | 29.23 | 844 | 0 | 30.5 | MT1-Rear | 7 | 67 | BSE
51 | P-11626 | 2025/03/18 16:41:42 | MT1-Rear-07 | 3 | -10.547 | -25.172 | 19827 | -1 | 1 |  |  | H2O | 67.714 | 0.191 | 0.442 | 89 | -0.011 | 0.122 | 29.35 | 844 | 0 | 30.188 | MT1-Rear | 7 | 67 | BSE
52 | P-11626 | 2025/03/18 16:50:30 | MT1-Rear-07 | 4 | -10.657 | -25.305 | 19817 | 0 | 1 |  |  | H2O | 67.7201 | 0.141 | 0.419 | 88 | -0.006 | -0.007 | 28.98 | 844 | 0 | 30.375 | MT1-Rear | 7 | 67 | BSE
53 | P-11627 | 2025/03/18 16:59:19 | MT1-Rear-08 | 1 | -8.785 | -17.007 | 19851 | -1 | 1 |  |  | H2O | 67.7263 | 0.125 | 0.507 | 89 | 0.003 | -0.108 | 29.45 | 844 | 0 | 30.625 | MT1-Rear | 8 | 67 | BSE
54 | P-11627 | 2025/03/18 17:08:08 | MT1-Rear-08 | 2 | -8.659 | -16.323 | 19801 | -1 | 1 |  |  | H2O | 67.7324 | 0.148 | 0.444 | 88 | -0.005 | 0.056 | 28.9 | 844 | 0 | 30.5 | MT1-Rear | 8 | 67 | BSE
55 | P-11627 | 2025/03/18 17:16:56 | MT1-Rear-08 | 3 | -8.634 | -15.961 | 19801 | -1 | 1 |  |  | H2O | 67.7385 | 0.171 | 0.458 | 89 | -0.028 | 0.111 | 29.52 | 844 | 0 | 30.5 | MT1-Rear | 8 | 67 | BSE
56 | P-11627 | 2025/03/18 17:25:45 | MT1-Rear-08 | 4 | -8.721 | -15.666 | 19811 | 0 | 1 |  |  | H2O | 67.7446 | 0.131 | 0.426 | 84 | -0.03 | -0.018 | 27.76 | 844 | 0 | 30.563 | MT1-Rear | 8 | 67 | BSE
57 | P-11628 | 2025/03/18 17:34:34 | MT1-Rear-09 | 1 | -11.807 | -35.762 | 19919 | -1 | 1 |  |  | H2O | 67.7507 | 0.099 | 0.435 | 89 | 0.002 | 0.048 | 29.48 | 844 | 0 | 30.438 | MT1-Rear | 9 | 67 | BSE
58 | P-11628 | 2025/03/18 17:43:22 | MT1-Rear-09 | 2 | -12.075 | -37.791 | 19789 | -1 | 1 |  |  | H2O | 67.7569 | 0.146 | 0.346 | 84 | -0.014 | -0.048 | 27.58 | 844 | 0 | 31.25 | MT1-Rear | 9 | 67 | BSE
59 | P-11628 | 2025/03/18 17:52:11 | MT1-Rear-09 | 3 | -12.056 | -38.299 | 19777 | -1 | 1 |  |  | H2O | 67.763 | 0.126 | 0.283 | 89 | 0.001 | 0.065 | 29.24 | 844 | 0 | 30.5 | MT1-Rear | 9 | 67 | BSE
60 | P-11628 | 2025/03/18 18:00:59 | MT1-Rear-09 | 4 | -12.091 | -38.387 | 19780 | 0 | 1 |  |  | H2O | 67.7691 | 0.209 | 0.354 | 88 | -0.032 | 0.095 | 28.97 | 844 | 0 | 30.75 | MT1-Rear | 9 | 67 | BSE
61 | P-11629 | 2025/03/18 18:09:49 | MT1-Rear-10 | 1 | -8.467 | -17.567 | 19837 | -1 | 1 |  |  | H2O | 67.7752 | 0.17 | 0.308 | 87 | -0.039 | -0.038 | 28.58 | 844 | 0 | 30.625 | MT1-Rear | 10 | 67 | BSE
62 | P-11629 | 2025/03/18 18:18:38 | MT1-Rear-10 | 2 | -8.454 | -15.925 | 19779 | -1 | 1 |  |  | H2O | 67.7814 | 0.157 | 0.354 | 88 | -0.035 | 0.074 | 29.08 | 844 | 0 | 30.75 | MT1-Rear | 10 | 67 | BSE
63 | P-11629 | 2025/03/18 18:27:27 | MT1-Rear-10 | 3 | -8.405 | -15.606 | 19769 | -1 | 1 |  |  | H2O | 67.7875 | 0.209 | 0.329 | 85 | -0.049 | 0.064 | 28.15 | 844 | 0 | 30.875 | MT1-Rear | 10 | 67 | BSE
64 | P-11629 | 2025/03/18 18:36:16 | MT1-Rear-10 | 4 | -8.392 | -15.4 | 19696 | 0 | 1 |  |  | H2O | 67.7936 | 0.146 | 0.303 | 89 | -0.015 | -0.056 | 29.51 | 844 | 0 | 31.063 | MT1-Rear | 10 | 67 | BSE
65 | P-11630 | 2025/03/18 18:45:05 | MT1-Rear-11 | 1 | -8.341 | -15.108 | 19840 | -1 | 1 |  |  | H2O | 67.7998 | 0.139 | 0.36 | 92 | -0.003 | 0.082 | 30.34 | 844 | 0 | 30.625 | MT1-Rear | 11 | 67 | BSE
66 | P-11630 | 2025/03/18 18:53:54 | MT1-Rear-11 | 2 | -8.256 | -15.153 | 19781 | -1 | 1 |  |  | H2O | 67.8059 | 0.121 | 0.389 | 88 | -0.022 | -0.09 | 28.96 | 844 | 0 | 30.813 | MT1-Rear | 11 | 67 | BSE
67 | P-11630 | 2025/03/18 19:02:42 | MT1-Rear-11 | 3 | -8.385 | -15.136 | 19860 | -1 | 1 |  |  | H2O | 67.812 | 0.096 | 0.321 | 88 | -0.011 | -0.01 | 29.2 | 844 | 0 | 30.438 | MT1-Rear | 11 | 67 | BSE
68 | P-11630 | 2025/03/18 19:11:32 | MT1-Rear-11 | 4 | -8.311 | -15.312 | 19834 | 0 | 1 |  |  | H2O | 67.8181 | 0.199 | 0.442 | 91 | -0.017 | 0.022 | 29.88 | 844 | 0 | 30.688 | MT1-Rear | 11 | 67 | BSE
69 | P-11631 | 2025/03/18 19:20:21 | MT1-Rear-12 | 1 | -8.645 | -10.36 | 19939 | -1 | 1 |  |  | H2O | 67.8243 | 0.181 | 0.353 | 88 | 0.038 | -0.042 | 28.93 | 844 | 0 | 30.875 | MT1-Rear | 12 | 67 | BSE
70 | P-11631 | 2025/03/18 19:29:10 | MT1-Rear-12 | 2 | -8.797 | -9.523 | 19859 | -1 | 1 |  |  | H2O | 67.8304 | 0.163 | 0.466 | 92 | 0.007 | 0.132 | 30.21 | 844 | 0 | 30.688 | MT1-Rear | 12 | 67 | BSE
71 | P-11631 | 2025/03/18 19:37:59 | MT1-Rear-12 | 3 | -8.676 | -9.593 | 19824 | -1 | 1 |  |  | H2O | 67.8365 | 0.221 | 0.305 | 87 | -0.047 | -0.036 | 28.64 | 844 | 0 | 30.563 | MT1-Rear | 12 | 67 | BSE
72 | P-11631 | 2025/03/18 19:46:49 | MT1-Rear-12 | 4 | -8.685 | -9.387 | 19666 | 0 | 1 |  |  | H2O | 67.8426 | 0.131 | 0.306 | 87 | -0.007 | -0.016 | 28.8 | 844 | 0 | 30.5 | MT1-Rear | 12 | 67 | BSE
73 | P-11632 | 2025/03/18 19:55:39 | MT1-Rear-13 | 1 | -8.213 | -14.401 | 19784 | -1 | 1 |  |  | H2O | 67.8487 | 0.121 | 0.39 | 87 | -0.033 | 0.07 | 28.85 | 844 | 0 | 30.5 | MT1-Rear | 13 | 67 | BSE
74 | P-11632 | 2025/03/18 20:04:29 | MT1-Rear-13 | 2 | -8.152 | -15.056 | 19719 | -1 | 1 |  |  | H2O | 67.8548 | 0.075 | 0.548 | 87 | -0.005 | 0.066 | 28.63 | 844 | 0 | 30.688 | MT1-Rear | 13 | 67 | BSE
75 | P-11632 | 2025/03/18 20:13:18 | MT1-Rear-13 | 3 | -8.207 | -15.1 | 19825 | -1 | 1 |  |  | H2O | 67.861 | 0.109 | 0.353 | 88 | -0.011 | 0.01 | 29.2 | 844 | 0 | 30.688 | MT1-Rear | 13 | 67 | BSE
76 | P-11632 | 2025/03/18 20:22:08 | MT1-Rear-13 | 4 | -8.127 | -15.131 | 19719 | 0 | 1 |  |  | H2O | 67.8671 | 0.211 | 0.23 | 86 | -0.027 | -0.008 | 28.36 | 844 | 0 | 30.375 | MT1-Rear | 13 | 67 | BSE
77 | P-11633 | 2025/03/18 20:30:57 | MT1-Rear-14 | 1 | -8.071 | -15.265 | 19815 | -1 | 1 |  |  | H2O | 67.8732 | 0.124 | 0.359 | 86 | -0.019 | 0.045 | 28.27 | 844 | 0 | 31.125 | MT1-Rear | 14 | 67 | BSE
78 | P-11633 | 2025/03/18 20:39:46 | MT1-Rear-14 | 2 | -8.099 | -15.379 | 19812 | -1 | 1 |  |  | H2O | 67.8793 | 0.171 | 0.275 | 87 | -0.013 | -0.016 | 28.7 | 844 | 0 | 30.563 | MT1-Rear | 14 | 67 | BSE
79 | P-11633 | 2025/03/18 20:48:36 | MT1-Rear-14 | 3 | -8.14 | -15.311 | 19776 | -1 | 1 |  |  | H2O | 67.8855 | 0.169 | 0.395 | 88 | -0.037 | 0.036 | 29.05 | 844 | 0 | 30.125 | MT1-Rear | 14 | 67 | BSE
80 | P-11633 | 2025/03/18 20:57:26 | MT1-Rear-14 | 4 | -8.176 | -15.297 | 19780 | 0 | 1 |  |  | H2O | 67.8917 | 0.185 | 0.265 | 94 | 0.015 | 0.008 | 28.48 | 844 | 0 | 30.188 | MT1-Rear | 14 | 67 | BSE
81 | P-11634 | 2025/03/18 21:06:16 | MT1-Rear-15 | 1 | -7.954 | -15.183 | 19832 | -1 | 1 |  |  | H2O | 67.8979 | 0.21 | 0.446 | 89 | -0.001 | -0.006 | 29.33 | 844 | 0 | 30.188 | MT1-Rear | 15 | 67 | BSE
82 | P-11634 | 2025/03/18 21:15:06 | MT1-Rear-15 | 2 | -7.942 | -15.692 | 19802 | -1 | 1 |  |  | H2O | 67.904 | 0.112 | 0.344 | 89 | -0.002 | -0.057 | 29.19 | 844 | 0 | 29.938 | MT1-Rear | 15 | 67 | BSE
83 | P-11634 | 2025/03/18 21:23:56 | MT1-Rear-15 | 3 | -8.026 | -15.485 | 19794 | -1 | 1 |  |  | H2O | 67.9101 | 0.187 | 0.414 | 87 | -0.013 | -0.018 | 28.66 | 844 | 0 | 30.125 | MT1-Rear | 15 | 67 | BSE
84 | P-11634 | 2025/03/18 21:32:47 | MT1-Rear-15 | 4 | -7.948 | -15.222 | 19813 | 0 | 1 |  |  | H2O | 67.9162 | 0.222 | 0.33 | 91 | 0.008 | -0.01 | 30.01 | 844 | 0 | 30.125 | MT1-Rear | 15 | 67 | BSE
85 | P-11635 | 2025/03/18 21:41:37 | MT1-Rear-16 | 1 | -9.228 | -20.842 | 19817 | -1 | 1 |  |  | H2O | 67.9224 | 0.203 | 0.322 | 90 | -0.042 | 0.061 | 29.71 | 844 | 0 | 30.25 | MT1-Rear | 16 | 67 | BSE
86 | P-11635 | 2025/03/18 21:50:27 | MT1-Rear-16 | 2 | -9.4 | -21.661 | 19742 | -1 | 1 |  |  | H2O | 67.9285 | 0.129 | 0.292 | 86 | -0.027 | 0.004 | 28.25 | 844 | 0 | 30.313 | MT1-Rear | 16 | 67 | BSE
87 | P-11635 | 2025/03/18 21:59:17 | MT1-Rear-16 | 3 | -9.327 | -21.416 | 19565 | -1 | 1 |  |  | H2O | 67.9346 | 0.134 | 0.536 | 86 | -0.018 | 0.108 | 28.52 | 844 | 0 | 30.563 | MT1-Rear | 16 | 67 | BSE
88 | P-11635 | 2025/03/18 22:08:07 | MT1-Rear-16 | 4 | -9.388 | -21.397 | 19807 | 0 | 1 |  |  | H2O | 67.9407 | 0.132 | 0.388 | 91 | -0.011 | 0.078 | 30.14 | 844 | 0 | 30.063 | MT1-Rear | 16 | 67 | BSE
89 | P-11636 | 2025/03/18 22:16:57 | MT1-Rear-17 | 1 | -7.923 | -15.995 | 19812 | -1 | 1 |  |  | H2O | 67.9468 | 0.141 | 0.4 | 88 | -0.028 | -0.016 | 28.91 | 844 | 0 | 30.75 | MT1-Rear | 17 | 67 | BSE
90 | P-11636 | 2025/03/18 22:25:47 | MT1-Rear-17 | 2 | -7.912 | -15.733 | 19731 | -1 | 1 |  |  | H2O | 67.953 | 0.169 | 0.459 | 86 | -0.029 | -0.059 | 28.36 | 844 | 0 | 30.313 | MT1-Rear | 17 | 67 | BSE
91 | P-11636 | 2025/03/18 22:34:38 | MT1-Rear-17 | 3 | -7.931 | -15.437 | 19667 | -1 | 1 |  |  | H2O | 67.9591 | 0.171 | 0.31 | 85 | -0.04 | 0.068 | 27.9 | 844 | 0 | 30.438 | MT1-Rear | 17 | 67 | BSE
92 | P-11636 | 2025/03/18 22:43:29 | MT1-Rear-17 | 4 | -7.89 | -15.378 | 19746 | 0 | 1 |  |  | H2O | 67.9654 | 0.15 | 0.432 | 86 | -0.02 | 0.049 | 28.52 | 844 | 0 | 30.125 | MT1-Rear | 17 | 67 | BSE
93 | P-11637 | 2025/03/18 22:52:19 | MT1-Rear-18 | 1 | -11.504 | -36.159 | 19752 | -1 | 1 |  |  | H2O | 67.9715 | 0.089 | 0.398 | 91 | -0.016 | 0.079 | 29.99 | 844 | 0 | 29.688 | MT1-Rear | 18 | 67 | BSE
94 | P-11637 | 2025/03/18 23:01:10 | MT1-Rear-18 | 2 | -11.722 | -38.092 | 19707 | -1 | 1 |  |  | H2O | 67.9776 | 0.137 | 0.391 | 87 | -0.019 | 0.086 | 28.6 | 844 | 0 | 30.188 | MT1-Rear | 18 | 67 | BSE
95 | P-11637 | 2025/03/18 23:10:01 | MT1-Rear-18 | 3 | -11.781 | -38.563 | 19651 | -1 | 1 |  |  | H2O | 67.9837 | 0.168 | 0.451 | 89 | -0.027 | 0.115 | 29.32 | 844 | 0 | 30.063 | MT1-Rear | 18 | 67 | BSE
96 | P-11637 | 2025/03/18 23:18:51 | MT1-Rear-18 | 4 | -11.696 | -38.782 | 19651 | 0 | 1 |  |  | H2O | 67.9899 | 0.134 | 0.493 | 96 | -0.014 | 0.103 | 28.78 | 844 | 0 | 29.875 | MT1-Rear | 18 | 67 | BSE
97 | P-11638 | 2025/03/18 23:27:40 | MT1-Rear-19 | 1 | -8.303 | -18.172 | 19796 | -1 | 1 |  |  | H2O | 67.9961 | 0.142 | 0.449 | 91 | -0.021 | 0.095 | 29.95 | 844 | 0 | 30.125 | MT1-Rear | 19 | 67 | BSE
98 | P-11638 | 2025/03/18 23:36:32 | MT1-Rear-19 | 2 | -8.09 | -16.159 | 19577 | -1 | 1 |  |  | H2O | 68.0022 | 0.137 | 0.318 | 90 | -0.024 | -0.035 | 29.64 | 844 | 0 | 30 | MT1-Rear | 19 | 67 | BSE
99 | P-11638 | 2025/03/18 23:45:23 | MT1-Rear-19 | 3 | -8.091 | -15.891 | 19789 | -1 | 1 |  |  | H2O | 68.0083 | 0.15 | 0.303 | 85 | -0.023 | -0.023 | 27.98 | 844 | 0 | 30.125 | MT1-Rear | 19 | 67 | BSE
100 | P-11638 | 2025/03/18 23:54:13 | MT1-Rear-19 | 4 | -7.984 | -15.533 | 19702 | 0 | 1 |  |  | H2O | 68.0144 | 0.209 | 0.277 | 86 | -0.008 | 0.002 | 28.44 | 844 | 0 | 30.125 | MT1-Rear | 19 | 67 | BSE
101 | P-11639 | 2025/03/19 00:03:05 | MT1-Rear-20 | 1 | -7.932 | -15.66 | 19500 | -1 | 1 |  |  | H2O | 68.0207 | 0.159 | 0.313 | 89 | -0.018 | 0.081 | 29.34 | 844 | 0 | 31.25 | MT1-Rear | 20 | 67 | BSE
102 | P-11639 | 2025/03/19 00:11:55 | MT1-Rear-20 | 2 | -7.877 | -15.205 | 19780 | -1 | 1 |  |  | H2O | 68.0268 | 0.136 | 0.469 | 88 | -0.024 | -0.067 | 28.96 | 844 | 0 | 30 | MT1-Rear | 20 | 67 | BSE
103 | P-11639 | 2025/03/19 00:20:45 | MT1-Rear-20 | 3 | -7.88 | -15.697 | 19692 | -1 | 1 |  |  | H2O | 68.0329 | 0.159 | 0.357 | 90 | -0.022 | 0.07 | 29.79 | 844 | 0 | 30.25 | MT1-Rear | 20 | 67 | BSE
104 | P-11639 | 2025/03/19 00:29:36 | MT1-Rear-20 | 4 | -7.783 | -15.559 | 19760 | 0 | 1 |  |  | H2O | 68.039 | 0.174 | 0.44 | 88 | -0.034 | 0.044 | 29.2 | 844 | 0 | 30.125 | MT1-Rear | 20 | 67 | BSE
105 | P-11640 | 2025/03/19 00:38:28 | MT1-Rear-21 | 1 | -9.274 | -21.249 | 19702 | -1 | 1 |  |  | H2O | 68.0451 | 0.213 | 0.308 | 86 | -0.048 | -0.003 | 28.29 | 844 | 0 | 30.188 | MT1-Rear | 21 | 67 | BSE
106 | P-11640 | 2025/03/19 00:47:20 | MT1-Rear-21 | 2 | -9.339 | -22.143 | 19780 | -1 | 1 |  |  | H2O | 68.0513 | 0.185 | 0.362 | 94 | -0.039 | 0.015 | 28.22 | 844 | 0 | 30.125 | MT1-Rear | 21 | 67 | BSE
107 | P-11640 | 2025/03/19 00:56:10 | MT1-Rear-21 | 3 | -9.32 | -22.31 | 19731 | -1 | 1 |  |  | H2O | 68.0575 | 0.141 | 0.338 | 87 | -0.023 | -0.026 | 28.72 | 844 | 0 | 30.188 | MT1-Rear | 21 | 67 | BSE
108 | P-11640 | 2025/03/19 01:05:01 | MT1-Rear-21 | 4 | -9.4 | -22.11 | 19642 | 0 | 1 |  |  | H2O | 68.0636 | 0.164 | 0.436 | 84 | -0.02 | -0.012 | 27.72 | 844 | 0 | 30.25 | MT1-Rear | 21 | 67 | BSE
109 | P-11641 | 2025/03/19 01:13:51 | MT1-Rear-22 | 1 | -7.898 | -16.009 | 19740 | -1 | 1 |  |  | H2O | 68.0697 | 0.109 | 0.286 | 88 | -0.017 | -0.04 | 28.97 | 844 | 0 | 29.875 | MT1-Rear | 22 | 67 | BSE
110 | P-11641 | 2025/03/19 01:22:42 | MT1-Rear-22 | 2 | -7.834 | -15.64 | 19668 | -1 | 1 |  |  | H2O | 68.0759 | 0.149 | 0.411 | 89 | 0.008 | 0.011 | 29.26 | 844 | 0 | 30 | MT1-Rear | 22 | 67 | BSE
111 | P-11641 | 2025/03/19 01:31:34 | MT1-Rear-22 | 3 | -7.81 | -15.91 | 19644 | -1 | 1 |  |  | H2O | 68.082 | 0.176 | 0.425 | 94 | 0.007 | 0.03 | 28.46 | 844 | 0 | 30.125 | MT1-Rear | 22 | 67 | BSE
112 | P-11641 | 2025/03/19 01:40:24 | MT1-Rear-22 | 4 | -7.799 | -15.415 | 19603 | 0 | 1 |  |  | H2O | 68.0882 | 0.163 | 0.422 | 86 | -0.013 | -0.049 | 28.34 | 844 | 0 | 30.063 | MT1-Rear | 22 | 67 | BSE
113 | P-11642 | 2025/03/19 01:49:16 | MT1-Rear-23 | 1 | -10.638 | -26.981 | 19742 | -1 | 1 |  |  | H2O | 68.0943 | 0.142 | 0.282 | 89 | -0.039 | 0.042 | 29.4 | 844 | 0 | 29.938 | MT1-Rear | 23 | 67 | BSE
114 | P-11642 | 2025/03/19 01:58:07 | MT1-Rear-23 | 2 | -10.771 | -28.336 | 19694 | -1 | 1 |  |  | H2O | 68.1005 | 0.108 | 0.262 | 91 | -0.005 | 0.06 | 30 | 844 | 0 | 30 | MT1-Rear | 23 | 67 | BSE
115 | P-11642 | 2025/03/19 02:06:59 | MT1-Rear-23 | 3 | -10.778 | -28.57 | 19760 | -1 | 1 |  |  | H2O | 68.1066 | 0.118 | 0.367 | 88 | -0.03 | -0.01 | 29.1 | 844 | 0 | 30.063 | MT1-Rear | 23 | 67 | BSE
116 | P-11642 | 2025/03/19 02:15:49 | MT1-Rear-23 | 4 | -10.844 | -28.85 | 19735 | 0 | 1 |  |  | H2O | 68.1127 | 0.144 | 0.357 | 86 | -0.003 | 0.053 | 28.39 | 844 | 0 | 30.125 | MT1-Rear | 23 | 67 | BSE
117 | P-11643 | 2025/03/19 02:24:40 | MT1-Rear-24 | 1 | -9.36 | -22.807 | 19712 | -1 | 1 |  |  | H2O | 68.119 | 0.17 | 0.359 | 87 | -0.036 | -0.02 | 28.72 | 844 | 0 | 30.063 | MT1-Rear | 24 | 67 | BSE
118 | P-11643 | 2025/03/19 02:33:32 | MT1-Rear-24 | 2 | -9.177 | -22.153 | 19661 | -1 | 1 |  |  | H2O | 68.1251 | 0.193 | 0.275 | 86 | -0.047 | -0.004 | 28.43 | 844 | 0 | 30.25 | MT1-Rear | 24 | 67 | BSE
119 | P-11643 | 2025/03/19 02:42:23 | MT1-Rear-24 | 3 | -9.236 | -21.759 | 19724 | -1 | 1 |  |  | H2O | 68.1312 | 0.106 | 0.298 | 87 | -0.006 | -0.019 | 28.75 | 844 | 0 | 29.875 | MT1-Rear | 24 | 67 | BSE
120 | P-11643 | 2025/03/19 02:51:16 | MT1-Rear-24 | 4 | -9.235 | -21.935 | 19676 | 0 | 1 |  |  | H2O | 68.1374 | 0.103 | 0.417 | 90 | -0.006 | 0.118 | 29.61 | 844 | 0 | 29.688 | MT1-Rear | 24 | 67 | BSE
121 | P-11644 | 2025/03/19 03:00:08 | MT1-Rear-25 | 1 | -9.142 | -21.884 | 19721 | -1 | 1 |  |  | H2O | 68.1435 | 0.104 | 0.527 | 85 | -0.024 | -0.068 | 28.01 | 844 | 0 | 30.063 | MT1-Rear | 25 | 67 | BSE
122 | P-11644 | 2025/03/19 03:09:00 | MT1-Rear-25 | 2 | -9.234 | -21.563 | 19766 | -1 | 1 |  |  | H2O | 68.1498 | 0.147 | 0.409 | 87 | -0.035 | 0.027 | 28.75 | 844 | 0 | 30 | MT1-Rear | 25 | 67 | BSE
123 | P-11644 | 2025/03/19 03:17:52 | MT1-Rear-25 | 3 | -9.207 | -21.798 | 19734 | -1 | 1 |  |  | H2O | 68.1559 | 0.085 | 0.397 | 84 | -0.005 | -0.077 | 27.75 | 844 | 0 | 29.938 | MT1-Rear | 25 | 67 | BSE
124 | P-11644 | 2025/03/19 03:26:44 | MT1-Rear-25 | 4 | -9.144 | -22.079 | 19665 | 0 | 1 |  |  | H2O | 68.162 | 0.113 | 0.442 | 85 | -0.023 | -0.09 | 27.9 | 844 | 0 | 29.75 | MT1-Rear | 25 | 67 | BSE
125 | P-11645 | 2025/03/19 03:35:35 | MT1-Rear-26 | 1 | -7.719 | -5.378 | 19730 | -1 | 1 |  |  | H2O | 68.1681 | 0.165 | 0.313 | 85 | -0.018 | -0.046 | 27.97 | 844 | 0 | 29.938 | MT1-Rear | 26 | 67 | BSE
126 | P-11645 | 2025/03/19 03:44:27 | MT1-Rear-26 | 2 | -7.623 | -3.309 | 19763 | -1 | 1 |  |  | H2O | 68.1744 | 0.156 | 0.439 | 91 | 0.02 | -0.029 | 30.04 | 844 | 0 | 30.063 | MT1-Rear | 26 | 67 | BSE
127 | P-11645 | 2025/03/19 03:53:18 | MT1-Rear-26 | 3 | -7.731 | -3.102 | 19661 | -1 | 1 |  |  | H2O | 68.1805 | 0.094 | 0.51 | 92 | 0.005 | 0.105 | 30.37 | 844 | 0 | 30.25 | MT1-Rear | 26 | 67 | BSE
128 | P-11645 | 2025/03/19 04:02:10 | MT1-Rear-26 | 4 | -7.603 | -2.927 | 19689 | 0 | 1 |  |  | H2O | 68.1866 | 0.15 | 0.338 | 86 | -0.037 | -0.018 | 28.4 | 844 | 0 | 29.813 | MT1-Rear | 26 | 67 | BSE
129 | P-11646 | 2025/03/19 04:11:02 | MT1-Rear-27 | 1 | -11.383 | -34.482 | 19725 | -1 | 1 |  |  | H2O | 68.1928 | 0.163 | 0.396 | 83 | -0.028 | 0.041 | 27.44 | 844 | 0 | 29.75 | MT1-Rear | 27 | 67 | BSE
130 | P-11646 | 2025/03/19 04:19:54 | MT1-Rear-27 | 2 | -11.508 | -38.191 | 19699 | -1 | 1 |  |  | H2O | 68.199 | 0.152 | 0.393 | 86 | -0.033 | 0.076 | 28.53 | 844 | 0 | 30.188 | MT1-Rear | 27 | 67 | BSE
131 | P-11646 | 2025/03/19 04:28:46 | MT1-Rear-27 | 3 | -11.657 | -38.484 | 19717 | -1 | 1 |  |  | H2O | 68.2051 | 0.155 | 0.49 | 86 | -0.004 | -0.06 | 28.26 | 844 | 0 | 30.188 | MT1-Rear | 27 | 67 | BSE
132 | P-11646 | 2025/03/19 04:37:37 | MT1-Rear-27 | 4 | -11.622 | -38.573 | 19687 | 0 | 1 |  |  | H2O | 68.2113 | 0.135 | 0.408 | 87 | -0.01 | 0.011 | 28.87 | 844 | 0 | 30 | MT1-Rear | 27 | 67 | BSE

## Run Summary
Line | Sample | Inj Nr | Sample Name | H2O_Mean | Good | Ignore | Time Code | d(18_16)Mean | d(D_H)Mean
1 | 1 | 1 | Blacksburg | 19394 | 1 | 0 | 2025/03/18 09:22:35 | -12.126 | -44.641 |  | Min | Max
2 | 1 | 2 | Blacksburg | 19697 | 1 | 0 | 2025/03/18 09:31:22 | -11.886 | -40.155 | 0 | 17000 | 23000
3 | 1 | 3 | Blacksburg | 19665 | 1 | 0 | 2025/03/18 09:40:08 | -11.941 | -39.673 | 500 | 17000 | 23000
4 | 1 | 4 | Blacksburg | 19730 | 1 | 0 | 2025/03/18 09:48:54 | -11.861 | -39.483
5 | 1 | 5 | Blacksburg | 19645 | 1 | 0 | 2025/03/18 09:57:40 | -11.832 | -39.619
6 | 1 | 6 | Blacksburg | 19726 | 1 | 0 | 2025/03/18 10:06:26 | -11.779 | -39.477
7 | 1 | 7 | Blacksburg | 19707 | 1 | 0 | 2025/03/18 10:15:13 | -11.738 | -39.419
8 | 1 | 8 | Blacksburg | 19506 | 1 | 0 | 2025/03/18 10:23:58 | -11.762 | -39.291
9 | 1 | 9 | Blacksburg | 19482 | 1 | 0 | 2025/03/18 10:32:44 | -11.728 | -39.3
10 | 1 | 10 | Blacksburg | 19464 | 1 | 0 | 2025/03/18 10:41:30 | -11.704 | -39.214
11 | 2 | 1 | Myrtle | 19736 | 1 | 0 | 2025/03/18 10:50:16 | -6.799 | -5.618
12 | 2 | 2 | Myrtle | 19672 | 1 | 0 | 2025/03/18 10:59:02 | -6.518 | -1.901
13 | 2 | 3 | Myrtle | 19650 | 1 | 0 | 2025/03/18 11:07:49 | -6.484 | -0.952
14 | 2 | 4 | Myrtle | 19676 | 1 | 0 | 2025/03/18 11:16:36 | -6.364 | -0.878
15 | 2 | 5 | Myrtle | 19631 | 1 | 0 | 2025/03/18 11:25:22 | -6.353 | -0.407
16 | 2 | 6 | Myrtle | 19645 | 1 | 0 | 2025/03/18 11:34:08 | -6.367 | -0.221
17 | 2 | 7 | Myrtle | 19789 | 1 | 0 | 2025/03/18 11:42:55 | -6.305 | 0.036 | Concentration of water vapor should be about 20,000 ppmV (roughly +/- 1000; but not outside 17,000 and 23,000 ppmV).
18 | 2 | 8 | Myrtle | 19619 | 1 | 0 | 2025/03/18 11:51:42 | -6.307 | -0.114 | Large scatter may indicate a problem with the syringe or the septa.
19 | 2 | 9 | Myrtle | 19629 | 1 | 0 | 2025/03/18 12:00:28 | -6.288 | -0.081
20 | 2 | 10 | Myrtle | 19607 | 1 | 0 | 2025/03/18 12:09:15 | -6.319 | 0.024
21 | 3 | 1 | Homer | 19728 | 1 | 0 | 2025/03/18 12:18:02 | -17.794 | -89.866
22 | 3 | 2 | Homer | 19627 | 1 | 0 | 2025/03/18 12:26:49 | -18.781 | -100.037
23 | 3 | 3 | Homer | 19664 | 1 | 0 | 2025/03/18 12:35:35 | -18.842 | -102.607
24 | 3 | 4 | Homer | 19743 | 1 | 0 | 2025/03/18 12:44:21 | -18.95 | -103.387
25 | 3 | 5 | Homer | 19779 | 1 | 0 | 2025/03/18 12:53:08 | -19.008 | -103.884
26 | 3 | 6 | Homer | 19744 | 1 | 0 | 2025/03/18 13:01:56 | -19.03 | -104.221
27 | 3 | 7 | Homer | 19477 | 1 | 0 | 2025/03/18 13:10:43 | -19.064 | -104.099
28 | 3 | 8 | Homer | 19778 | 1 | 0 | 2025/03/18 13:19:30 | -19.123 | -104.686
29 | 3 | 9 | Homer | 19755 | 1 | 0 | 2025/03/18 13:28:17 | -19.069 | -104.564
30 | 3 | 10 | Homer | 19774 | 1 | 0 | 2025/03/18 13:37:04 | -19.166 | -104.324
31 | 4 | 1 | Blacksburg | 19788 | 1 | 0 | 2025/03/18 13:45:52 | -12.472 | -47.313
32 | 4 | 2 | Blacksburg | 19765 | 1 | 0 | 2025/03/18 13:54:39 | -12.045 | -41.765
33 | 4 | 3 | Blacksburg | 19724 | 1 | 0 | 2025/03/18 14:03:26 | -12.017 | -40.47
34 | 4 | 4 | Blacksburg | 19762 | 1 | 0 | 2025/03/18 14:12:12 | -11.969 | -39.508
35 | 4 | 5 | Blacksburg | 19762 | 1 | 0 | 2025/03/18 14:20:59 | -11.982 | -39.504
36 | 4 | 6 | Blacksburg | 19784 | 1 | 0 | 2025/03/18 14:29:46 | -11.915 | -39.33
37 | 4 | 7 | Blacksburg | 19698 | 1 | 0 | 2025/03/18 14:38:33 | -11.859 | -39.365
38 | 4 | 8 | Blacksburg | 20636 | 1 | 0 | 2025/03/18 14:47:21 | -12.005 | -39.03
39 | 4 | 9 | Blacksburg | 20305 | 1 | 0 | 2025/03/18 14:56:08 | -12.025 | -39.03
40 | 4 | 10 | Blacksburg | 19743 | 1 | 0 | 2025/03/18 15:04:56 | -11.902 | -39.267
41 | 5 | 1 | Hawaii | 19743 | 1 | 0 | 2025/03/18 15:13:43 | -8.225 | -7.524
42 | 5 | 2 | Hawaii | 19775 | 1 | 0 | 2025/03/18 15:22:31 | -7.972 | -3.784
43 | 5 | 3 | Hawaii | 19786 | 1 | 0 | 2025/03/18 15:31:19 | -8.024 | -3.107
44 | 5 | 4 | Hawaii | 19769 | 1 | 0 | 2025/03/18 15:40:07 | -8.007 | -2.853
45 | 6 | 1 | CC4 30sep24 BOT | 19759 | 1 | 0 | 2025/03/18 15:48:55 | -9.347 | -19.206
46 | 6 | 2 | CC4 30sep24 BOT | 19796 | 1 | 0 | 2025/03/18 15:57:42 | -9.467 | -21.041
47 | 6 | 3 | CC4 30sep24 BOT | 19816 | 1 | 0 | 2025/03/18 16:06:30 | -9.438 | -21.457
48 | 6 | 4 | CC4 30sep24 BOT | 19749 | 1 | 0 | 2025/03/18 16:15:17 | -9.604 | -21.152
49 | 7 | 1 | CS1 28oct24 0.1m | 19815 | 1 | 0 | 2025/03/18 16:24:05 | -10.505 | -24.933
50 | 7 | 2 | CS1 28oct24 0.1m | 19821 | 1 | 0 | 2025/03/18 16:32:53 | -10.537 | -25.204
51 | 7 | 3 | CS1 28oct24 0.1m | 19827 | 1 | 0 | 2025/03/18 16:41:42 | -10.547 | -25.172
52 | 7 | 4 | CS1 28oct24 0.1m | 19817 | 1 | 0 | 2025/03/18 16:50:30 | -10.657 | -25.305
53 | 8 | 1 | CC2 28oct24 BOT | 19851 | 1 | 0 | 2025/03/18 16:59:19 | -8.785 | -17.007
54 | 8 | 2 | CC2 28oct24 BOT | 19801 | 1 | 0 | 2025/03/18 17:08:08 | -8.659 | -16.323
55 | 8 | 3 | CC2 28oct24 BOT | 19801 | 1 | 0 | 2025/03/18 17:16:56 | -8.634 | -15.961
56 | 8 | 4 | CC2 28oct24 BOT | 19811 | 1 | 0 | 2025/03/18 17:25:45 | -8.721 | -15.666
57 | 9 | 1 | Blacksburg | 19919 | 1 | 0 | 2025/03/18 17:34:34 | -11.807 | -35.762
58 | 9 | 2 | Blacksburg | 19789 | 1 | 0 | 2025/03/18 17:43:22 | -12.075 | -37.791
59 | 9 | 3 | Blacksburg | 19777 | 1 | 0 | 2025/03/18 17:52:11 | -12.056 | -38.299
60 | 9 | 4 | Blacksburg | 19780 | 1 | 0 | 2025/03/18 18:00:59 | -12.091 | -38.387
61 | 10 | 1 | CC4 30sep24 1.5m | 19837 | 1 | 0 | 2025/03/18 18:09:49 | -8.467 | -17.567
62 | 10 | 2 | CC4 30sep24 1.5m | 19779 | 1 | 0 | 2025/03/18 18:18:38 | -8.454 | -15.925
63 | 10 | 3 | CC4 30sep24 1.5m | 19769 | 1 | 0 | 2025/03/18 18:27:27 | -8.405 | -15.606
64 | 10 | 4 | CC4 30sep24 1.5m | 19696 | 1 | 0 | 2025/03/18 18:36:16 | -8.392 | -15.4
65 | 11 | 1 | CC2 30sep24 BOT | 19840 | 1 | 0 | 2025/03/18 18:45:05 | -8.341 | -15.108
66 | 11 | 2 | CC2 30sep24 BOT | 19781 | 1 | 0 | 2025/03/18 18:53:54 | -8.256 | -15.153
67 | 11 | 3 | CC2 30sep24 BOT | 19860 | 1 | 0 | 2025/03/18 19:02:42 | -8.385 | -15.136
68 | 11 | 4 | CC2 30sep24 BOT | 19834 | 1 | 0 | 2025/03/18 19:11:32 | -8.311 | -15.312
69 | 12 | 1 | CCR rain 30sep24 | 19939 | 1 | 0 | 2025/03/18 19:20:21 | -8.645 | -10.36
70 | 12 | 2 | CCR rain 30sep24 | 19859 | 1 | 0 | 2025/03/18 19:29:10 | -8.797 | -9.523
71 | 12 | 3 | CCR rain 30sep24 | 19824 | 1 | 0 | 2025/03/18 19:37:59 | -8.676 | -9.593
72 | 12 | 4 | CCR rain 30sep24 | 19666 | 1 | 0 | 2025/03/18 19:46:49 | -8.685 | -9.387
73 | 13 | 1 | CC4 30sep24 0.1m | 19784 | 1 | 0 | 2025/03/18 19:55:39 | -8.213 | -14.401
74 | 13 | 2 | CC4 30sep24 0.1m | 19719 | 1 | 0 | 2025/03/18 20:04:29 | -8.152 | -15.056
75 | 13 | 3 | CC4 30sep24 0.1m | 19825 | 1 | 0 | 2025/03/18 20:13:18 | -8.207 | -15.1
76 | 13 | 4 | CC4 30sep24 0.1m | 19719 | 1 | 0 | 2025/03/18 20:22:08 | -8.127 | -15.131
77 | 14 | 1 | CC3 30sep24 BOT | 19815 | 1 | 0 | 2025/03/18 20:30:57 | -8.071 | -15.265
78 | 14 | 2 | CC3 30sep24 BOT | 19812 | 1 | 0 | 2025/03/18 20:39:46 | -8.099 | -15.379
79 | 14 | 3 | CC3 30sep24 BOT | 19776 | 1 | 0 | 2025/03/18 20:48:36 | -8.14 | -15.311
80 | 14 | 4 | CC3 30sep24 BOT | 19780 | 1 | 0 | 2025/03/18 20:57:26 | -8.176 | -15.297
81 | 15 | 1 | CC4 16sep24 0.1m | 19832 | 1 | 0 | 2025/03/18 21:06:16 | -7.954 | -15.183
82 | 15 | 2 | CC4 16sep24 0.1m | 19802 | 1 | 0 | 2025/03/18 21:15:06 | -7.942 | -15.692
83 | 15 | 3 | CC4 16sep24 0.1m | 19794 | 1 | 0 | 2025/03/18 21:23:56 | -8.026 | -15.485
84 | 15 | 4 | CC4 16sep24 0.1m | 19813 | 1 | 0 | 2025/03/18 21:32:47 | -7.948 | -15.222
85 | 16 | 1 | C50 17feb25 9m | 19817 | 1 | 0 | 2025/03/18 21:41:37 | -9.228 | -20.842
86 | 16 | 2 | C50 17feb25 9m | 19742 | 1 | 0 | 2025/03/18 21:50:27 | -9.4 | -21.661
87 | 16 | 3 | C50 17feb25 9m | 19565 | 1 | 0 | 2025/03/18 21:59:17 | -9.327 | -21.416
88 | 16 | 4 | C50 17feb25 9m | 19807 | 1 | 0 | 2025/03/18 22:08:07 | -9.388 | -21.397
89 | 17 | 1 | CC3 16sep24 1.5m | 19812 | 1 | 0 | 2025/03/18 22:16:57 | -7.923 | -15.995
90 | 17 | 2 | CC3 16sep24 1.5m | 19731 | 1 | 0 | 2025/03/18 22:25:47 | -7.912 | -15.733
91 | 17 | 3 | CC3 16sep24 1.5m | 19667 | 1 | 0 | 2025/03/18 22:34:38 | -7.931 | -15.437
92 | 17 | 4 | CC3 16sep24 1.5m | 19746 | 1 | 0 | 2025/03/18 22:43:29 | -7.89 | -15.378
93 | 18 | 1 | Blacksburg | 19752 | 1 | 0 | 2025/03/18 22:52:19 | -11.504 | -36.159
94 | 18 | 2 | Blacksburg | 19707 | 1 | 0 | 2025/03/18 23:01:10 | -11.722 | -38.092
95 | 18 | 3 | Blacksburg | 19651 | 1 | 0 | 2025/03/18 23:10:01 | -11.781 | -38.563
96 | 18 | 4 | Blacksburg | 19651 | 1 | 0 | 2025/03/18 23:18:51 | -11.696 | -38.782
97 | 19 | 1 | CC3 30sep24 0.1m | 19796 | 1 | 0 | 2025/03/18 23:27:40 | -8.303 | -18.172
98 | 19 | 2 | CC3 30sep24 0.1m | 19577 | 1 | 0 | 2025/03/18 23:36:32 | -8.09 | -16.159
99 | 19 | 3 | CC3 30sep24 0.1m | 19789 | 1 | 0 | 2025/03/18 23:45:23 | -8.091 | -15.891
100 | 19 | 4 | CC3 30sep24 0.1m | 19702 | 1 | 0 | 2025/03/18 23:54:13 | -7.984 | -15.533
101 | 20 | 1 | CC3 16sep24 BOT | 19500 | 1 | 0 | 2025/03/19 00:03:05 | -7.932 | -15.66
102 | 20 | 2 | CC3 16sep24 BOT | 19780 | 1 | 0 | 2025/03/19 00:11:55 | -7.877 | -15.205
103 | 20 | 3 | CC3 16sep24 BOT | 19692 | 1 | 0 | 2025/03/19 00:20:45 | -7.88 | -15.697
104 | 20 | 4 | CC3 16sep24 BOT | 19760 | 1 | 0 | 2025/03/19 00:29:36 | -7.783 | -15.559
105 | 21 | 1 | CC3 17feb25 0.1m | 19702 | 1 | 0 | 2025/03/19 00:38:28 | -9.274 | -21.249
106 | 21 | 2 | CC3 17feb25 0.1m | 19780 | 1 | 0 | 2025/03/19 00:47:20 | -9.339 | -22.143
107 | 21 | 3 | CC3 17feb25 0.1m | 19731 | 1 | 0 | 2025/03/19 00:56:10 | -9.32 | -22.31
108 | 21 | 4 | CC3 17feb25 0.1m | 19642 | 1 | 0 | 2025/03/19 01:05:01 | -9.4 | -22.11
109 | 22 | 1 | CC3 16sep24 0.1m | 19740 | 1 | 0 | 2025/03/19 01:13:51 | -7.898 | -16.009
110 | 22 | 2 | CC3 16sep24 0.1m | 19668 | 1 | 0 | 2025/03/19 01:22:42 | -7.834 | -15.64
111 | 22 | 3 | CC3 16sep24 0.1m | 19644 | 1 | 0 | 2025/03/19 01:31:34 | -7.81 | -15.91
112 | 22 | 4 | CC3 16sep24 0.1m | 19603 | 1 | 0 | 2025/03/19 01:40:24 | -7.799 | -15.415
113 | 23 | 1 | CS2 17feb25 0.1m | 19742 | 1 | 0 | 2025/03/19 01:49:16 | -10.638 | -26.981
114 | 23 | 2 | CS2 17feb25 0.1m | 19694 | 1 | 0 | 2025/03/19 01:58:07 | -10.771 | -28.336
115 | 23 | 3 | CS2 17feb25 0.1m | 19760 | 1 | 0 | 2025/03/19 02:06:59 | -10.778 | -28.57
116 | 23 | 4 | CS2 17feb25 0.1m | 19735 | 1 | 0 | 2025/03/19 02:15:49 | -10.844 | -28.85
117 | 24 | 1 | C50 17feb25 0.1m | 19712 | 1 | 0 | 2025/03/19 02:24:40 | -9.36 | -22.807
118 | 24 | 2 | C50 17feb25 0.1m | 19661 | 1 | 0 | 2025/03/19 02:33:32 | -9.177 | -22.153
119 | 24 | 3 | C50 17feb25 0.1m | 19724 | 1 | 0 | 2025/03/19 02:42:23 | -9.236 | -21.759
120 | 24 | 4 | C50 17feb25 0.1m | 19676 | 1 | 0 | 2025/03/19 02:51:16 | -9.235 | -21.935
121 | 25 | 1 | C50 17feb25 9m - DUP | 19721 | 1 | 0 | 2025/03/19 03:00:08 | -9.142 | -21.884
122 | 25 | 2 | C50 17feb25 9m - DUP | 19766 | 1 | 0 | 2025/03/19 03:09:00 | -9.234 | -21.563
123 | 25 | 3 | C50 17feb25 9m - DUP | 19734 | 1 | 0 | 2025/03/19 03:17:52 | -9.207 | -21.798
124 | 25 | 4 | C50 17feb25 9m - DUP | 19665 | 1 | 0 | 2025/03/19 03:26:44 | -9.144 | -22.079
125 | 26 | 1 | Hawaii | 19730 | 1 | 0 | 2025/03/19 03:35:35 | -7.719 | -5.378
126 | 26 | 2 | Hawaii | 19763 | 1 | 0 | 2025/03/19 03:44:27 | -7.623 | -3.309
127 | 26 | 3 | Hawaii | 19661 | 1 | 0 | 2025/03/19 03:53:18 | -7.731 | -3.102
128 | 26 | 4 | Hawaii | 19689 | 1 | 0 | 2025/03/19 04:02:10 | -7.603 | -2.927
129 | 27 | 1 | Blacksburg | 19725 | 1 | 0 | 2025/03/19 04:11:02 | -11.383 | -34.482
130 | 27 | 2 | Blacksburg | 19699 | 1 | 0 | 2025/03/19 04:19:54 | -11.508 | -38.191
131 | 27 | 3 | Blacksburg | 19717 | 1 | 0 | 2025/03/19 04:28:46 | -11.657 | -38.484
132 | 27 | 4 | Blacksburg | 19687 | 1 | 0 | 2025/03/19 04:37:37 | -11.622 | -38.573

## 18O
Line | Sample | Injection | Sample Name | Correction Type | Good | Ignore | d(18_16) Picarro Raw | Stdev Raw | d(18_16) Memory Corrected | Stdev Memory Corrected | d(18_16) Memory Corrected Drifts | d(18_16) Drift Corrected | Drift Corrected Drifts | d(18_16) Drift Corrected High, Low, and Drift | d(18_16) Reference Defined | d18O Final Corrected Value | Post run corrections:
1 | 1 | 1 | Blacksburg | Drift | 1 | -1 | -12.126 | [9 empty] | (1) Memory correction |  |  | (1) Graphs for memory correction
2 | 1 | 2 | Blacksburg | Drift | 1 | -1 | -11.886 | [9 empty] | Combined Stdev | 0.088955
3 | 1 | 3 | Blacksburg | Drift | 1 | -1 | -11.941 | [9 empty] | Date:
4 | 1 | 4 | Blacksburg | Drift | 1 | -1 | -11.861 | [9 empty] | injection | 18O memory coef
5 | 1 | 5 | Blacksburg | Drift | 1 | -1 | -11.832 | [9 empty] | 1 | 0.89101
6 | 1 | 6 | Blacksburg | Drift | 1 | -1 | -11.779 | [9 empty] | 2 | 0.97089
7 | 1 | 7 | Blacksburg | Drift | 1 | 0 | -11.738 |  | -11.738 |  | -11.738 | -11.7423 | -11.7423 |  |  | -7.4524 | 3 | 0.97588
8 | 1 | 8 | Blacksburg | Drift | 1 | 0 | -11.762 |  | -11.762 |  | -11.762 | -11.7669 | -11.7669 |  |  | -7.4759 | 4 | 0.9859
9 | 1 | 9 | Blacksburg | Drift | 1 | 0 | -11.728 |  | -11.728 |  | -11.728 | -11.7335 | -11.7335 |  |  | -7.4441 | 5 | 0.98884
10 | 1 | 10 | Blacksburg | Drift | 1 | 0 | -11.704 | 0.024028 | -11.704 | 0.024028 | -11.704 | -11.7102 | -11.7102 |  |  | -7.4218 | 6 | 0.99171
11 | 2 | 1 | Myrtle | High Reference | 1 | 0 | -6.799 |  | -6.2644 |  |  | -6.2712 |  | -6.2712 | -2.2908 | -2.2317 | 7 | 0.99472
12 | 2 | 2 | Myrtle | High Reference | 1 | 0 | -6.518 |  | -6.367 |  |  | -6.3744 |  | -6.3744 | -2.2908 | -2.3302 | 8 | 0.99472
13 | 2 | 3 | Myrtle | High Reference | 1 | 0 | -6.484 |  | -6.3581 |  |  | -6.3661 |  | -6.3661 | -2.2908 | -2.3223 | 9 | 0.99472
14 | 2 | 4 | Myrtle | High Reference | 1 | 0 | -6.364 |  | -6.2887 |  |  | -6.2973 |  | -6.2973 | -2.2908 | -2.2567 | 10 | 1
15 | 2 | 5 | Myrtle | High Reference | 1 | 0 | -6.353 |  | -6.2933 |  |  | -6.3025 |  | -6.3025 | -2.2908 | -2.2616
16 | 2 | 6 | Myrtle | High Reference | 1 | 0 | -6.367 |  | -6.3228 |  |  | -6.3326 |  | -6.3326 | -2.2908 | -2.2904 | (2) Drift correction
17 | 2 | 7 | Myrtle | High Reference | 1 | 0 | -6.305 |  | -6.2765 |  |  | -6.2869 |  | -6.2869 | -2.2908 | -2.2468 | Slope m: | 0.00061543
18 | 2 | 8 | Myrtle | High Reference | 1 | 0 | -6.307 |  | -6.2785 |  |  | -6.2896 |  | -6.2896 | -2.2908 | -2.2493 | Stdev Mem Corr. | 0.18578
19 | 2 | 9 | Myrtle | High Reference | 1 | 0 | -6.288 |  | -6.2594 |  |  | -6.2711 |  | -6.2711 | -2.2908 | -2.2316 | Stdev Drift Corr. | 0.18355
20 | 2 | 10 | Myrtle | High Reference | 1 | 0 | -6.319 | 0.15651 | -6.319 | 0.037678 |  | -6.3313 |  | -6.3313 | -2.2908 | -2.2891
21 | 3 | 1 | Homer | Low Reference | 1 | 0 | -17.794 |  | -19.0447 |  |  | -19.0576 |  | -19.0576 | -14.5399 | -14.4329 | (3) Normalization to VSMOW scale
22 | 3 | 2 | Homer | Low Reference | 1 | 0 | -18.781 |  | -19.1438 |  |  | -19.1574 |  | -19.1574 | -14.5399 | -14.5281 | slope m | 0.95423
23 | 3 | 3 | Homer | Low Reference | 1 | 0 | -18.842 |  | -19.1441 |  |  | -19.1582 |  | -19.1582 | -14.5399 | -14.5289 | interception t | 3.7524
24 | 3 | 4 | Homer | Low Reference | 1 | 0 | -18.95 |  | -19.1282 |  |  | -19.1429 |  | -19.1429 | -14.5399 | -14.5143
25 | 3 | 5 | Homer | Low Reference | 1 | 0 | -19.008 |  | -19.1496 |  |  | -19.165 |  | -19.165 | -14.5399 | -14.5354 | (4) Quality control |  | offset
26 | 3 | 6 | Homer | Low Reference | 1 | 0 | -19.03 |  | -19.1353 |  |  | -19.1513 |  | -19.1513 | -14.5399 | -14.5224 | QC Check Defined | -3.6108 | accepted ± 0.15
27 | 3 | 7 | Homer | Low Reference | 1 | 0 | -19.064 |  | -19.1314 |  |  | -19.148 |  | -19.148 | -14.5399 | -14.5191 | QC Check Vial 5 | -3.8025 | -0.19163
28 | 3 | 8 | Homer | Low Reference | 1 | 0 | -19.123 |  | -19.1907 |  |  | -19.2079 |  | -19.2079 | -14.5399 | -14.5763 | QC Check Vial 26 | -3.5789 | 0.031886
29 | 3 | 9 | Homer | Low Reference | 1 | 0 | -19.069 |  | -19.1364 |  |  | -19.1542 |  | -19.1542 | -14.5399 | -14.5251
30 | 3 | 10 | Homer | Low Reference | 1 | 0 | -19.166 | 0.40066 | -19.166 | 0.037435 |  | -19.1845 |  | -19.1845 | -14.5399 | -14.554 | Drift Defined | -7.5615 | accepted ± 0.15
31 | 4 | 1 | Blacksburg | Drift | 1 | 0 | -12.472 |  | -11.7424 |  |  | -11.7615 |  | -11.7615 | -7.5615 | -7.4707 | Drift Average | -7.6134 | -0.051904
32 | 4 | 2 | Blacksburg | Drift | 1 | 0 | -12.045 |  | -11.8377 |  |  | -11.8574 |  | -11.8574 | -7.5615 | -7.5622 | Drift Vial 4 | -7.6438 | -0.08236
33 | 4 | 3 | Blacksburg | Drift | 1 | 0 | -12.017 |  | -11.8446 |  |  | -11.8649 |  | -11.8649 | -7.5615 | -7.5694 | Drift Vial 9 | -7.8737 | -0.31219
34 | 4 | 4 | Blacksburg | Drift | 1 | 0 | -11.969 |  | -11.8675 |  |  | -11.8884 |  | -11.8884 | -7.5615 | -7.5919 | Drift Vial 18 | -7.6002 | -0.038718
35 | 4 | 5 | Blacksburg | Drift | 1 | 0 | -11.982 |  | -11.9018 |  |  | -11.9234 |  | -11.9234 | -7.5615 | -7.6252 | Drift Vial 27 | -7.5007 | 0.060814
36 | 4 | 6 | Blacksburg | Drift | 1 | 0 | -11.915 |  | -11.8549 |  |  | -11.8771 |  | -11.8771 | -7.5615 | -7.581
37 | 4 | 7 | Blacksburg | Drift | 1 | 0 | -11.859 |  | -11.8204 |  |  | -11.8432 |  | -11.8432 | -7.5615 | -7.5487 | Reference Defined Values |  | d18O
38 | 4 | 8 | Blacksburg | Drift | 1 | 0 | -12.005 |  | -11.9672 |  |  | -11.9905 |  | -11.9905 | -7.5615 | -7.6893 | High Reference | Myrtle Beach | -2.2908
39 | 4 | 9 | Blacksburg | Drift | 1 | 0 | -12.025 |  | -11.9873 |  |  | -12.0113 |  | -12.0113 | -7.5615 | -7.7091 | Low Reference | Homer | -14.5399
40 | 4 | 10 | Blacksburg | Drift | 1 | 0 | -11.902 | 0.17 | -11.902 | 0.071358 |  | -11.9266 |  | -11.9266 | -7.5615 | -7.6283 | Drift | Blacksburg | -7.5615
41 | 5 | 1 | Hawaii | QC Check | 1 | 0 | -8.225 |  | -7.8242 |  |  | -7.8495 |  |  |  | -3.7378 | QC Check | Hawaii | -3.6108
42 | 5 | 2 | Hawaii | QC Check | 1 | 0 | -7.972 |  | -7.8576 |  |  | -7.8834 |  |  |  | -3.7702 | Calibration Date:
43 | 5 | 3 | Hawaii | QC Check | 1 | 0 | -8.024 |  | -7.9305 |  |  | -7.9569 |  |  |  | -3.8403
44 | 5 | 4 | Hawaii | QC Check | 1 | 0 | -8.007 | 0.11407 | -7.9521 | 0.060164 |  | -7.9791 |  |  |  | -3.8615
45 | 6 | 1 | CC4 30sep24 BOT |  | 1 | 0 | -9.347 |  | -9.493 |  |  | -9.5207 |  |  |  | -5.3326
46 | 6 | 2 | CC4 30sep24 BOT |  | 1 | 0 | -9.467 |  | -9.5095 |  |  | -9.5378 |  |  |  | -5.3489 | Denali | -19.9287
47 | 6 | 3 | CC4 30sep24 BOT |  | 1 | 0 | -9.438 |  | -9.4725 |  |  | -9.5014 |  |  |  | -5.3141 | Alberta | -18.8392
48 | 6 | 4 | CC4 30sep24 BOT |  | 1 | 0 | -9.604 | 0.10642 | -9.6265 | 0.069094 |  | -9.6561 |  |  |  | -5.4617 | MT BHB | -18.431
49 | 7 | 1 | CS1 28oct24 0.1m |  | 1 | 0 | -10.505 |  | -10.6032 |  |  | -10.6334 |  |  |  | -6.3943 | Anchorage | -16.4204
50 | 7 | 2 | CS1 28oct24 0.1m |  | 1 | 0 | -10.537 |  | -10.5642 |  |  | -10.5949 |  |  |  | -6.3576 | Homer | -14.5399
51 | 7 | 3 | CS1 28oct24 0.1m |  | 1 | 0 | -10.547 |  | -10.5697 |  |  | -10.6011 |  |  |  | -6.3635 | MT Creek | -13.3562
52 | 7 | 4 | CS1 28oct24 0.1m |  | 1 | 0 | -10.657 | 0.066139 | -10.6719 | 0.049506 |  | -10.7039 |  |  |  | -6.4615 | Nickel Mines | -7.876
53 | 8 | 1 | CC2 28oct24 BOT |  | 1 | 0 | -8.785 |  | -8.581 |  |  | -8.6136 |  |  |  | -4.4669 | Blacksburg | -7.5615
54 | 8 | 2 | CC2 28oct24 BOT |  | 1 | 0 | -8.659 |  | -8.6008 |  |  | -8.6341 |  |  |  | -4.4865 | Denton's Well | -6.87
55 | 8 | 3 | CC2 28oct24 BOT |  | 1 | 0 | -8.634 |  | -8.5852 |  |  | -8.6191 |  |  |  | -4.4721 | Hawaii | -3.6108
56 | 8 | 4 | CC2 28oct24 BOT |  | 1 | 0 | -8.721 | 0.067584 | -8.6937 | 0.053039 |  | -8.7282 |  |  |  | -4.5763 | Myrtle Beach | -2.2908
57 | 9 | 1 | Blacksburg | Drift | 1 | 0 | -11.807 |  | -12.1433 |  | -12.1433 | -12.1784 | -12.1784 |  |  | -7.8686 | Heavy Oxygen | 3.1753
58 | 9 | 2 | Blacksburg | Drift | 1 | 0 | -12.075 |  | -12.1726 |  | -12.1726 | -12.2083 | -12.2083 |  |  | -7.8971 | Heavy Hydrogen | 11.1372
59 | 9 | 3 | Blacksburg | Drift | 1 | 0 | -12.056 |  | -12.1364 |  | -12.1364 | -12.1728 | -12.1728 |  |  | -7.8632 | Andres Heavy Water | 17.5664
60 | 9 | 4 | Blacksburg | Drift | 1 | 0 | -12.091 | 0.13426 | -12.1385 | 0.016852 | -12.1385 | -12.1755 | -12.1755 |  |  | -7.8658 | Picarro ZERO | 0.3
61 | 10 | 1 | CC4 30sep24 1.5m |  | 1 | 0 | -8.467 |  | -8.072 |  |  | -8.1096 |  |  |  | -3.986 | Picarro MID | -20.6
62 | 10 | 2 | CC4 30sep24 1.5m |  | 1 | 0 | -8.454 |  | -8.3481 |  |  | -8.3863 |  |  |  | -4.25 | Picarro DEPL | -29.6
63 | 10 | 3 | CC4 30sep24 1.5m |  | 1 | 0 | -8.405 |  | -8.3161 |  |  | -8.3549 |  |  |  | -4.22
64 | 10 | 4 | CC4 30sep24 1.5m |  | 1 | 0 | -8.392 | 0.036574 | -8.3398 | 0.13203 |  | -8.3792 |  |  |  | -4.2433
65 | 11 | 1 | CC2 30sep24 BOT |  | 1 | 0 | -8.341 |  | -8.3354 |  |  | -8.3754 |  |  |  | -4.2397
66 | 11 | 2 | CC2 30sep24 BOT |  | 1 | 0 | -8.256 |  | -8.252 |  |  | -8.2927 |  |  |  | -4.1607
67 | 11 | 3 | CC2 30sep24 BOT |  | 1 | 0 | -8.385 |  | -8.3848 |  |  | -8.4261 |  |  |  | -4.288
68 | 11 | 4 | CC2 30sep24 BOT |  | 1 | 0 | -8.311 | 0.054163 | -8.3099 | 0.055262 |  | -8.3517 |  |  |  | -4.217
69 | 12 | 1 | CCR rain 30sep24 |  | 1 | 0 | -8.645 |  | -8.6814 |  |  | -8.7239 |  |  |  | -4.5722
70 | 12 | 2 | CCR rain 30sep24 |  | 1 | 0 | -8.797 |  | -8.8111 |  |  | -8.8542 |  |  |  | -4.6966
71 | 12 | 3 | CCR rain 30sep24 |  | 1 | 0 | -8.676 |  | -8.6848 |  |  | -8.7285 |  |  |  | -4.5766
72 | 12 | 4 | CCR rain 30sep24 |  | 1 | 0 | -8.685 | 0.066415 | -8.6903 | 0.062934 |  | -8.7346 |  |  |  | -4.5824
73 | 13 | 1 | CC4 30sep24 0.1m |  | 1 | 0 | -8.213 |  | -8.1616 |  |  | -8.2065 |  |  |  | -4.0785
74 | 13 | 2 | CC4 30sep24 0.1m |  | 1 | 0 | -8.152 |  | -8.1365 |  |  | -8.182 |  |  |  | -4.0551
75 | 13 | 3 | CC4 30sep24 0.1m |  | 1 | 0 | -8.207 |  | -8.1955 |  |  | -8.2416 |  |  |  | -4.112
76 | 13 | 4 | CC4 30sep24 0.1m |  | 1 | 0 | -8.127 | 0.042035 | -8.1191 | 0.03315 |  | -8.1659 |  |  |  | -4.0397
77 | 14 | 1 | CC3 30sep24 BOT |  | 1 | 0 | -8.071 |  | -8.0649 |  |  | -8.1123 |  |  |  | -3.9886
78 | 14 | 2 | CC3 30sep24 BOT |  | 1 | 0 | -8.099 |  | -8.0982 |  |  | -8.1462 |  |  |  | -4.0209
79 | 14 | 3 | CC3 30sep24 BOT |  | 1 | 0 | -8.14 |  | -8.1403 |  |  | -8.1889 |  |  |  | -4.0617
80 | 14 | 4 | CC3 30sep24 BOT |  | 1 | 0 | -8.176 | 0.046076 | -8.1767 | 0.048781 |  | -8.2259 |  |  |  | -4.097
81 | 15 | 1 | CC4 16sep24 0.1m |  | 1 | 0 | -7.954 |  | -7.9298 |  |  | -7.9797 |  |  |  | -3.862
82 | 15 | 2 | CC4 16sep24 0.1m |  | 1 | 0 | -7.942 |  | -7.9352 |  |  | -7.9857 |  |  |  | -3.8677
83 | 15 | 3 | CC4 16sep24 0.1m |  | 1 | 0 | -8.026 |  | -8.0224 |  |  | -8.0735 |  |  |  | -3.9515
84 | 15 | 4 | CC4 16sep24 0.1m |  | 1 | 0 | -7.948 | 0.039306 | -7.9448 | 0.04334 |  | -7.9965 |  |  |  | -3.8781
85 | 16 | 1 | C50 17feb25 9m |  | 1 | 0 | -9.228 |  | -9.3675 |  |  | -9.4198 |  |  |  | -5.2363
86 | 16 | 2 | C50 17feb25 9m |  | 1 | 0 | -9.4 |  | -9.4423 |  |  | -9.4952 |  |  |  | -5.3082
87 | 16 | 3 | C50 17feb25 9m |  | 1 | 0 | -9.327 |  | -9.3603 |  |  | -9.4138 |  |  |  | -5.2305
88 | 16 | 4 | C50 17feb25 9m |  | 1 | 0 | -9.388 | 0.078623 | -9.4083 | 0.038182 |  | -9.4625 |  |  |  | -5.277
89 | 17 | 1 | CC3 16sep24 1.5m |  | 1 | 0 | -7.923 |  | -7.7633 |  |  | -7.8181 |  |  |  | -3.7079
90 | 17 | 2 | CC3 16sep24 1.5m |  | 1 | 0 | -7.912 |  | -7.869 |  |  | -7.9244 |  |  |  | -3.8093
91 | 17 | 3 | CC3 16sep24 1.5m |  | 1 | 0 | -7.931 |  | -7.8959 |  |  | -7.9519 |  |  |  | -3.8355
92 | 17 | 4 | CC3 16sep24 1.5m |  | 1 | 0 | -7.89 | 0.017795 | -7.8689 | 0.058682 |  | -7.9255 |  |  |  | -3.8103
93 | 18 | 1 | Blacksburg | Drift | 1 | 0 | -11.504 |  | -11.8979 |  | -11.8979 | -11.9551 | -11.9551 |  |  | -7.6555
94 | 18 | 2 | Blacksburg | Drift | 1 | 0 | -11.722 |  | -11.8336 |  | -11.8336 | -11.8914 | -11.8914 |  |  | -7.5947
95 | 18 | 3 | Blacksburg | Drift | 1 | 0 | -11.781 |  | -11.8749 |  | -11.8749 | -11.9333 | -11.9333 |  |  | -7.6347
96 | 18 | 4 | Blacksburg | Drift | 1 | 0 | -11.696 | 0.1199 | -11.7497 | 0.065222 | -11.7497 | -11.8088 | -11.8088 |  |  | -7.5159
97 | 19 | 1 | CC3 30sep24 0.1m |  | 1 | 0 | -8.303 |  | -7.9332 |  |  | -7.9929 |  |  |  | -3.8746
98 | 19 | 2 | CC3 30sep24 0.1m |  | 1 | 0 | -8.09 |  | -7.985 |  |  | -8.0453 |  |  |  | -3.9247
99 | 19 | 3 | CC3 30sep24 0.1m |  | 1 | 0 | -8.091 |  | -8.004 |  |  | -8.065 |  |  |  | -3.9434
100 | 19 | 4 | CC3 30sep24 0.1m |  | 1 | 0 | -7.984 | 0.13378 | -7.9316 | 0.036697 |  | -7.9932 |  |  |  | -3.8749
101 | 20 | 1 | CC3 16sep24 BOT |  | 1 | 0 | -7.932 |  | -7.9263 |  |  | -7.9885 |  |  |  | -3.8704
102 | 20 | 2 | CC3 16sep24 BOT |  | 1 | 0 | -7.877 |  | -7.8739 |  |  | -7.9367 |  |  |  | -3.821
103 | 20 | 3 | CC3 16sep24 BOT |  | 1 | 0 | -7.88 |  | -7.8775 |  |  | -7.9409 |  |  |  | -3.825
104 | 20 | 4 | CC3 16sep24 BOT |  | 1 | 0 | -7.783 | 0.062038 | -7.7802 | 0.061081 |  | -7.8442 |  |  |  | -3.7327
105 | 21 | 1 | CC3 17feb25 0.1m |  | 1 | 0 | -9.274 |  | -9.4365 |  |  | -9.5011 |  |  |  | -5.3138
106 | 21 | 2 | CC3 17feb25 0.1m |  | 1 | 0 | -9.339 |  | -9.3843 |  |  | -9.4495 |  |  |  | -5.2646
107 | 21 | 3 | CC3 17feb25 0.1m |  | 1 | 0 | -9.32 |  | -9.3571 |  |  | -9.4229 |  |  |  | -5.2392
108 | 21 | 4 | CC3 17feb25 0.1m |  | 1 | 0 | -9.4 | 0.052201 | -9.4228 | 0.036248 |  | -9.4893 |  |  |  | -5.3025
109 | 22 | 1 | CC3 16sep24 0.1m |  | 1 | 0 | -7.898 |  | -7.7343 |  |  | -7.8014 |  |  |  | -3.6919
110 | 22 | 2 | CC3 16sep24 0.1m |  | 1 | 0 | -7.834 |  | -7.7884 |  |  | -7.8561 |  |  |  | -3.7441
111 | 22 | 3 | CC3 16sep24 0.1m |  | 1 | 0 | -7.81 |  | -7.7716 |  |  | -7.84 |  |  |  | -3.7287
112 | 22 | 4 | CC3 16sep24 0.1m |  | 1 | 0 | -7.799 | 0.044312 | -7.7764 | 0.023355 |  | -7.8453 |  |  |  | -3.7339
113 | 23 | 1 | CS2 17feb25 0.1m |  | 1 | 0 | -10.638 |  | -10.9474 |  |  | -11.017 |  |  |  | -6.7603
114 | 23 | 2 | CS2 17feb25 0.1m |  | 1 | 0 | -10.771 |  | -10.8575 |  |  | -10.9277 |  |  |  | -6.6751
115 | 23 | 3 | CS2 17feb25 0.1m |  | 1 | 0 | -10.778 |  | -10.8499 |  |  | -10.9206 |  |  |  | -6.6684
116 | 23 | 4 | CS2 17feb25 0.1m |  | 1 | 0 | -10.844 | 0.086342 | -10.8869 | 0.044309 |  | -10.9583 |  |  |  | -6.7044
117 | 24 | 1 | C50 17feb25 0.1m |  | 1 | 0 | -9.36 |  | -9.1983 |  |  | -9.2703 |  |  |  | -5.0935
118 | 24 | 2 | C50 17feb25 0.1m |  | 1 | 0 | -9.177 |  | -9.1285 |  |  | -9.2011 |  |  |  | -5.0275
119 | 24 | 3 | C50 17feb25 0.1m |  | 1 | 0 | -9.236 |  | -9.1972 |  |  | -9.2704 |  |  |  | -5.0937
120 | 24 | 4 | C50 17feb25 0.1m |  | 1 | 0 | -9.235 | 0.077102 | -9.2123 | 0.037695 |  | -9.2862 |  |  |  | -5.1087
121 | 25 | 1 | C50 17feb25 9m - DUP |  | 1 | 0 | -9.142 |  | -9.1319 |  |  | -9.2063 |  |  |  | -5.0325
122 | 25 | 2 | C50 17feb25 9m - DUP |  | 1 | 0 | -9.234 |  | -9.234 |  |  | -9.3091 |  |  |  | -5.1306
123 | 25 | 3 | C50 17feb25 9m - DUP |  | 1 | 0 | -9.207 |  | -9.2063 |  |  | -9.282 |  |  |  | -5.1048
124 | 25 | 4 | C50 17feb25 9m - DUP |  | 1 | 0 | -9.144 | 0.04609 | -9.1427 | 0.049351 |  | -9.219 |  |  |  | -5.0447
125 | 26 | 1 | Hawaii |  | 1 | 0 | -7.719 |  | -7.5637 |  |  | -7.6406 |  |  |  | -3.5385
126 | 26 | 2 | Hawaii |  | 1 | 0 | -7.623 |  | -7.5787 |  |  | -7.6563 |  |  |  | -3.5534
127 | 26 | 3 | Hawaii |  | 1 | 0 | -7.731 |  | -7.6969 |  |  | -7.7751 |  |  |  | -3.6668
128 | 26 | 4 | Hawaii |  | 1 | 0 | -7.603 | 0.065361 | -7.5813 | 0.061669 |  | -7.66 |  |  |  | -3.557
129 | 27 | 1 | Blacksburg | Drift | 1 | 0 | -11.383 |  | -11.795 |  | -11.795 | -11.8744 | -11.8744 |  |  | -7.5785
130 | 27 | 2 | Blacksburg | Drift | 1 | 0 | -11.508 |  | -11.6217 |  | -11.6217 | -11.7017 | -11.7017 |  |  | -7.4137
131 | 27 | 3 | Blacksburg | Drift | 1 | 0 | -11.657 |  | -11.7548 |  | -11.7548 | -11.8354 | -11.8354 |  |  | -7.5413
132 | 27 | 4 | Blacksburg | Drift | 1 | 0 | -11.622 | 0.12391 | -11.6787 | 0.077423 | -11.6787 | -11.7599 | -11.7599 |  |  | -7.4693
Note on G2: Manually change first six injections to ignore "-1".
Note on G3: Manually change first six injections to ignore "-1".
Note on G4: Manually change first six injections to ignore "-1".
Note on G5: Manually change first six injections to ignore "-1".
Note on G6: Manually change first six injections to ignore "-1".
Note on G7: Manually change first six injections to ignore "-1".

## 2H
Line | Sample | Injection | Sample Name | Correction Type | Good | Ignore | dD_H) Picarro Raw | Stdev Raw | d(D_H) Memory Corrected | Stdev Memory Corrected | d(D_H) Memory Corrected Drift | d(D_H) Drift Corrected | Drift Corrected Drift | d(D_H) Drift Corrected High, Low, Drift | d(D_H) Reference Defined | dD Final Corrected Value | Post run corrections:
1 | 1 | 1 | Blacksburg | Drift | 1 | -1 | -44.641 | [9 empty] | (1) Memory correction |  |  | (1) Graphs for memory correction
2 | 1 | 2 | Blacksburg | Drift | 1 | -1 | -40.155 | [9 empty] | Combined Stdev | 0.37022
3 | 1 | 3 | Blacksburg | Drift | 1 | -1 | -39.673 | [9 empty] | Date:
4 | 1 | 4 | Blacksburg | Drift | 1 | -1 | -39.483 | [9 empty] | injection | 18O memory coef
5 | 1 | 5 | Blacksburg | Drift | 1 | -1 | -39.619 | [9 empty] | 1 | 0.84172
6 | 1 | 6 | Blacksburg | Drift | 1 | -1 | -39.477 | [9 empty] | 2 | 0.95537
7 | 1 | 7 | Blacksburg | Drift | 1 | 0 | -39.419 |  | -39.419 |  | -39.419 | -39.4138 | -39.4138 |  |  | -48.0374 | 3 | 0.98024
8 | 1 | 8 | Blacksburg | Drift | 1 | 0 | -39.291 |  | -39.291 |  | -39.291 | -39.2851 | -39.2851 |  |  | -47.9139 | 4 | 0.98933
9 | 1 | 9 | Blacksburg | Drift | 1 | 0 | -39.3 |  | -39.3 |  | -39.3 | -39.2933 | -39.2933 |  |  | -47.9218 | 5 | 0.99361
10 | 1 | 10 | Blacksburg | Drift | 1 | 0 | -39.214 | 0.084644 | -39.214 | 0.084644 | -39.214 | -39.2066 | -39.2066 |  |  | -47.8386 | 6 | 0.99672
11 | 2 | 1 | Myrtle | High Reference | 1 | 0 | -5.618 |  | -0.30051 |  |  | -0.29237 |  | -0.29237 | -10.453 | -10.501 | 7 | 0.99672
12 | 2 | 2 | Myrtle | High Reference | 1 | 0 | -1.901 |  | -0.23565 |  |  | -0.22677 |  | -0.22677 | -10.453 | -10.4381 | 8 | 1
13 | 2 | 3 | Myrtle | High Reference | 1 | 0 | -0.952 |  | -0.19601 |  |  | -0.1864 |  | -0.1864 | -10.453 | -10.3994 | 9 | 1
14 | 2 | 4 | Myrtle | High Reference | 1 | 0 | -0.878 |  | -0.4688 |  |  | -0.45845 |  | -0.45845 | -10.453 | -10.6604 | 10 | 1
15 | 2 | 5 | Myrtle | High Reference | 1 | 0 | -0.407 |  | -0.15918 |  |  | -0.14809 |  | -0.14809 | -10.453 | -10.3626
16 | 2 | 6 | Myrtle | High Reference | 1 | 0 | -0.221 |  | -0.093088 |  |  | -0.081255 |  | -0.081255 | -10.453 | -10.2985 | (2) Drift correction
17 | 2 | 7 | Myrtle | High Reference | 1 | 0 | 0.036 |  | 0.16475 |  |  | 0.17733 |  | 0.17733 | -10.453 | -10.0504 | Slope m: | -0.0007396
18 | 2 | 8 | Myrtle | High Reference | 1 | 0 | -0.114 |  | -0.114 |  |  | -0.10069 |  | -0.10069 | -10.453 | -10.3171 | Stdev Mem Corr. | 0.30324
19 | 2 | 9 | Myrtle | High Reference | 1 | 0 | -0.081 |  | -0.081 |  |  | -0.066948 |  | -0.066948 | -10.453 | -10.2848 | Stdev Drift Corr. | 0.30128
20 | 2 | 10 | Myrtle | High Reference | 1 | 0 | 0.024 | 1.7277 | 0.024 | 0.17394 |  | 0.038792 |  | 0.038792 | -10.453 | -10.1833
21 | 3 | 1 | Homer | Low Reference | 1 | 0 | -89.866 |  | -104.0936 |  |  | -104.078 |  | -104.078 | -110.5169 | -110.0816 | (3) Normalization to VSMOW scale
22 | 3 | 2 | Homer | Low Reference | 1 | 0 | -100.037 |  | -104.5029 |  |  | -104.4866 |  | -104.4866 | -110.5169 | -110.4737 | slope m | 0.95948
23 | 3 | 3 | Homer | Low Reference | 1 | 0 | -102.607 |  | -104.6348 |  |  | -104.6178 |  | -104.6178 | -110.5169 | -110.5995 | interception t | -10.2205
24 | 3 | 4 | Homer | Low Reference | 1 | 0 | -103.387 |  | -104.4908 |  |  | -104.4731 |  | -104.4731 | -110.5169 | -110.4607
25 | 3 | 5 | Homer | Low Reference | 1 | 0 | -103.884 |  | -104.5475 |  |  | -104.5291 |  | -104.5291 | -110.5169 | -110.5144 | (4) Quality control |  | offset
26 | 3 | 6 | Homer | Low Reference | 1 | 0 | -104.221 |  | -104.563 |  |  | -104.5437 |  | -104.5437 | -110.5169 | -110.5285 | QC Check Defined | -13.2294 | accepted ±1.5
27 | 3 | 7 | Homer | Low Reference | 1 | 0 | -104.099 |  | -104.4406 |  |  | -104.4206 |  | -104.4206 | -110.5169 | -110.4103 | QC Check Vial 5 | -12.4828 | 0.74659
28 | 3 | 8 | Homer | Low Reference | 1 | 0 | -104.686 |  | -104.686 |  |  | -104.6653 |  | -104.6653 | -110.5169 | -110.6451 | QC Check Vial 26 | -12.6867 | 0.5427
29 | 3 | 9 | Homer | Low Reference | 1 | 0 | -104.564 |  | -104.564 |  |  | -104.5426 |  | -104.5426 | -110.5169 | -110.5273
30 | 3 | 10 | Homer | Low Reference | 1 | 0 | -104.324 | 4.536 | -104.324 | 0.17015 |  | -104.3018 |  | -104.3018 | -110.5169 | -110.2963 | Drift Defined | -47.446 | accepted ±1.5
31 | 4 | 1 | Blacksburg | Drift | 1 | 0 | -47.313 |  | -38.2894 |  |  | -38.2665 |  | -38.2665 | -47.446 | -46.9366 | Drift Average | -47.7272 | -0.28119
32 | 4 | 2 | Blacksburg | Drift | 1 | 0 | -41.765 |  | -38.9729 |  |  | -38.9492 |  | -38.9492 | -47.446 | -47.5916 | Drift Vial 4 | -47.7279 | -0.2819
33 | 4 | 3 | Blacksburg | Drift | 1 | 0 | -40.47 |  | -39.2084 |  |  | -39.184 |  | -39.184 | -47.446 | -47.8169 | Drift Vial 9 | -47.3822 | 0.063789
34 | 4 | 4 | Blacksburg | Drift | 1 | 0 | -39.508 |  | -38.8162 |  |  | -38.791 |  | -38.791 | -47.446 | -47.4398 | Drift Vial 18 | -47.7189 | -0.27285
35 | 4 | 5 | Blacksburg | Drift | 1 | 0 | -39.504 |  | -39.0901 |  |  | -39.0642 |  | -39.0642 | -47.446 | -47.7019 | Drift Vial 27 | -47.8791 | -0.43309
36 | 4 | 6 | Blacksburg | Drift | 1 | 0 | -39.33 |  | -39.1168 |  |  | -39.0902 |  | -39.0902 | -47.446 | -47.7269
37 | 4 | 7 | Blacksburg | Drift | 1 | 0 | -39.365 |  | -39.1519 |  |  | -39.1245 |  | -39.1245 | -47.446 | -47.7599 | Reference Defined Values |  | d2H
38 | 4 | 8 | Blacksburg | Drift | 1 | 0 | -39.03 |  | -39.03 |  |  | -39.0019 |  | -39.0019 | -47.446 | -47.6422 | High Reference | North Myrtle Beach, SC | -10.453
39 | 4 | 9 | Blacksburg | Drift | 1 | 0 | -39.03 |  | -39.03 |  |  | -39.0012 |  | -39.0012 | -47.446 | -47.6415 | Low Reference | Homer, AK | -110.5169
40 | 4 | 10 | Blacksburg | Drift | 1 | 0 | -39.267 | 2.549 | -39.267 | 0.27902 |  | -39.2374 |  | -39.2374 | -47.446 | -47.8682 | Drift | Blacksburg | -47.446
41 | 5 | 1 | Hawaii | QC Check | 1 | 0 | -7.524 |  | -2.4998 |  |  | -2.4695 |  |  |  | -12.5899 | QC Check | Hawaii | -13.2294
42 | 5 | 2 | Hawaii | QC Check | 1 | 0 | -3.784 |  | -2.2003 |  |  | -2.1693 |  |  |  | -12.3019 | Calibration Date:
43 | 5 | 3 | Hawaii | QC Check | 1 | 0 | -3.107 |  | -2.3925 |  |  | -2.3607 |  |  |  | -12.4856
44 | 5 | 4 | Hawaii | QC Check | 1 | 0 | -2.853 | 2.1738 | -2.4643 | 0.13361 |  | -2.4318 |  |  |  | -12.5538
45 | 6 | 1 | CC4 30sep24 BOT |  | 1 | 0 | -19.206 |  | -21.7943 |  |  | -21.761 |  |  |  | -31.0999
46 | 6 | 2 | CC4 30sep24 BOT |  | 1 | 0 | -21.041 |  | -21.8528 |  |  | -21.8187 |  |  |  | -31.1552 | Denali National Park | -153.7172
47 | 6 | 3 | CC4 30sep24 BOT |  | 1 | 0 | -21.457 |  | -21.8246 |  |  | -21.7898 |  |  |  | -31.1275 | Edmonton, Alberta | -146.8245
48 | 6 | 4 | CC4 30sep24 BOT |  | 1 | 0 | -21.152 | 1.0206 | -21.3473 | 0.23947 |  | -21.3118 |  |  |  | -30.6689 | MT BHB | -140.5778
49 | 7 | 1 | CS1 28oct24 0.1m |  | 1 | 0 | -24.933 |  | -25.5314 |  |  | -25.4952 |  |  |  | -34.6827 | Anchorage, AK | -129.9028
50 | 7 | 2 | CS1 28oct24 0.1m |  | 1 | 0 | -25.204 |  | -25.3848 |  |  | -25.3479 |  |  |  | -34.5414 | Homer, AK | -110.5169
51 | 7 | 3 | CS1 28oct24 0.1m |  | 1 | 0 | -25.172 |  | -25.2514 |  |  | -25.2137 |  |  |  | -34.4126 | MT Creek | -97.7428
52 | 7 | 4 | CS1 28oct24 0.1m |  | 1 | 0 | -25.305 | 0.15755 | -25.3493 | 0.11609 |  | -25.3109 |  |  |  | -34.5059 | Nickel Mines, PA | -48.7951
53 | 8 | 1 | CC2 28oct24 BOT |  | 1 | 0 | -17.007 |  | -15.6936 |  |  | -15.6544 |  |  |  | -25.2407 | Blacksburg | -47.446
54 | 8 | 2 | CC2 28oct24 BOT |  | 1 | 0 | -16.323 |  | -15.9221 |  |  | -15.8822 |  |  |  | -25.4592 | Denton's Well | -41.771
55 | 8 | 3 | CC2 28oct24 BOT |  | 1 | 0 | -15.961 |  | -15.7764 |  |  | -15.7357 |  |  |  | -25.3187 | Hawaii | -13.2294
56 | 8 | 4 | CC2 28oct24 BOT |  | 1 | 0 | -15.666 | 0.57807 | -15.5631 | 0.15047 |  | -15.5217 |  |  |  | -25.1133 | North Myrtle Beach, SC | -10.453
57 | 9 | 1 | Blacksburg | Drift | 1 | 0 | -35.762 |  | -38.9427 |  | -38.9427 | -38.9006 | -38.9006 |  |  | -47.545 | WQL Heavy Oxygen | -15.7881
58 | 9 | 2 | Blacksburg | Drift | 1 | 0 | -37.791 |  | -38.7785 |  | -38.7785 | -38.7356 | -38.7356 |  |  | -47.3867 | WQL Heavy Hydrogen | 4.0332
59 | 9 | 3 | Blacksburg | Drift | 1 | 0 | -38.299 |  | -38.7462 |  | -38.7462 | -38.7026 | -38.7026 |  |  | -47.355 | Andres Heavy Water | 20.5382
60 | 9 | 4 | Blacksburg | Drift | 1 | 0 | -38.387 | 1.2269 | -38.6295 | 0.12928 | -38.6295 | -38.5851 | -38.5851 |  |  | -47.2423
61 | 10 | 1 | CC4 30sep24 1.5m |  | 1 | 0 | -17.567 |  | -14.2717 |  |  | -14.2265 |  |  |  | -23.8706
62 | 10 | 2 | CC4 30sep24 1.5m |  | 1 | 0 | -15.925 |  | -14.9225 |  |  | -14.8766 |  |  |  | -24.4944
63 | 10 | 3 | CC4 30sep24 1.5m |  | 1 | 0 | -15.606 |  | -15.1559 |  |  | -15.1093 |  |  |  | -24.7176
64 | 10 | 4 | CC4 30sep24 1.5m |  | 1 | 0 | -15.4 | 0.98562 | -15.1546 | 0.41768 |  | -15.1073 |  |  |  | -24.7157
65 | 11 | 1 | CC2 30sep24 BOT |  | 1 | 0 | -15.108 |  | -15.0618 |  |  | -15.0137 |  |  |  | -24.6259
66 | 11 | 2 | CC2 30sep24 BOT |  | 1 | 0 | -15.153 |  | -15.142 |  |  | -15.0932 |  |  |  | -24.7022
67 | 11 | 3 | CC2 30sep24 BOT |  | 1 | 0 | -15.136 |  | -15.1308 |  |  | -15.0812 |  |  |  | -24.6907
68 | 11 | 4 | CC2 30sep24 BOT |  | 1 | 0 | -15.312 | 0.091729 | -15.3111 | 0.10589 |  | -15.2608 |  |  |  | -24.863
69 | 12 | 1 | CCR rain 30sep24 |  | 1 | 0 | -10.36 |  | -9.5762 |  |  | -9.5252 |  |  |  | -19.3598
70 | 12 | 2 | CCR rain 30sep24 |  | 1 | 0 | -9.523 |  | -9.2646 |  |  | -9.2129 |  |  |  | -19.0601
71 | 12 | 3 | CCR rain 30sep24 |  | 1 | 0 | -9.593 |  | -9.48 |  |  | -9.4275 |  |  |  | -19.266
72 | 12 | 4 | CCR rain 30sep24 |  | 1 | 0 | -9.387 | 0.43793 | -9.3238 | 0.1427 |  | -9.2705 |  |  |  | -19.1154
73 | 13 | 1 | CC4 30sep24 0.1m |  | 1 | 0 | -14.401 |  | -15.1946 |  |  | -15.1406 |  |  |  | -24.7477
74 | 13 | 2 | CC4 30sep24 0.1m |  | 1 | 0 | -15.056 |  | -15.309 |  |  | -15.2543 |  |  |  | -24.8567
75 | 13 | 3 | CC4 30sep24 0.1m |  | 1 | 0 | -15.1 |  | -15.2129 |  |  | -15.1574 |  |  |  | -24.7638
76 | 13 | 4 | CC4 30sep24 0.1m |  | 1 | 0 | -15.131 | 0.34869 | -15.1923 | 0.055314 |  | -15.1361 |  |  |  | -24.7434
77 | 14 | 1 | CC3 30sep24 BOT |  | 1 | 0 | -15.265 |  | -15.2862 |  |  | -15.2293 |  |  |  | -24.8327
78 | 14 | 2 | CC3 30sep24 BOT |  | 1 | 0 | -15.379 |  | -15.3901 |  |  | -15.3324 |  |  |  | -24.9317
79 | 14 | 3 | CC3 30sep24 BOT |  | 1 | 0 | -15.311 |  | -15.3146 |  |  | -15.2561 |  |  |  | -24.8585
80 | 14 | 4 | CC3 30sep24 BOT |  | 1 | 0 | -15.297 | 0.048028 | -15.2988 | 0.046578 |  | -15.2396 |  |  |  | -24.8427
81 | 15 | 1 | CC4 16sep24 0.1m |  | 1 | 0 | -15.183 |  | -15.165 |  |  | -15.105 |  |  |  | -24.7136
82 | 15 | 2 | CC4 16sep24 0.1m |  | 1 | 0 | -15.692 |  | -15.7096 |  |  | -15.649 |  |  |  | -25.2355
83 | 15 | 3 | CC4 16sep24 0.1m |  | 1 | 0 | -15.485 |  | -15.4887 |  |  | -15.4273 |  |  |  | -25.0228
84 | 15 | 4 | CC4 16sep24 0.1m |  | 1 | 0 | -15.222 | 0.23887 | -15.2212 | 0.25225 |  | -15.1591 |  |  |  | -24.7654
85 | 16 | 1 | C50 17feb25 9m |  | 1 | 0 | -20.842 |  | -21.7315 |  |  | -21.6687 |  |  |  | -31.0112
86 | 16 | 2 | C50 17feb25 9m |  | 1 | 0 | -21.661 |  | -21.9484 |  |  | -21.8848 |  |  |  | -31.2186
87 | 16 | 3 | C50 17feb25 9m |  | 1 | 0 | -21.416 |  | -21.5384 |  |  | -21.474 |  |  |  | -30.8245
88 | 16 | 4 | C50 17feb25 9m |  | 1 | 0 | -21.397 | 0.34621 | -21.4629 | 0.21717 |  | -21.3978 |  |  |  | -30.7514
89 | 17 | 1 | CC3 16sep24 1.5m |  | 1 | 0 | -15.995 |  | -15.14 |  |  | -15.0742 |  |  |  | -24.6839
90 | 17 | 2 | CC3 16sep24 1.5m |  | 1 | 0 | -15.733 |  | -15.4802 |  |  | -15.4136 |  |  |  | -25.0096
91 | 17 | 3 | CC3 16sep24 1.5m |  | 1 | 0 | -15.437 |  | -15.3192 |  |  | -15.2519 |  |  |  | -24.8545
92 | 17 | 4 | CC3 16sep24 1.5m |  | 1 | 0 | -15.378 | 0.28546 | -15.3138 | 0.13896 |  | -15.2457 |  |  |  | -24.8485
93 | 18 | 1 | Blacksburg | Drift | 1 | 0 | -36.159 |  | -39.4482 |  | -39.4482 | -39.3794 | -39.3794 |  |  | -48.0044
94 | 18 | 2 | Blacksburg | Drift | 1 | 0 | -38.092 |  | -39.1058 |  | -39.1058 | -39.0362 | -39.0362 |  |  | -47.6751
95 | 18 | 3 | Blacksburg | Drift | 1 | 0 | -38.563 |  | -39.0211 |  | -39.0211 | -38.9508 | -38.9508 |  |  | -47.5932
96 | 18 | 4 | Blacksburg | Drift | 1 | 0 | -38.782 | 1.1952 | -39.0318 | 0.20119 | -39.0318 | -38.9608 | -38.9608 |  |  | -47.6028
97 | 19 | 1 | CC3 30sep24 0.1m |  | 1 | 0 | -18.172 |  | -14.9099 |  |  | -14.8382 |  |  |  | -24.4575
98 | 19 | 2 | CC3 30sep24 0.1m |  | 1 | 0 | -16.159 |  | -15.1493 |  |  | -15.0768 |  |  |  | -24.6865
99 | 19 | 3 | CC3 30sep24 0.1m |  | 1 | 0 | -15.891 |  | -15.4387 |  |  | -15.3655 |  |  |  | -24.9634
100 | 19 | 4 | CC3 30sep24 0.1m |  | 1 | 0 | -15.533 | 1.1836 | -15.2848 | 0.22423 |  | -15.2109 |  |  |  | -24.8151
101 | 20 | 1 | CC3 16sep24 BOT |  | 1 | 0 | -15.66 |  | -15.6801 |  |  | -15.6054 |  |  |  | -25.1936
102 | 20 | 2 | CC3 16sep24 BOT |  | 1 | 0 | -15.205 |  | -15.1904 |  |  | -15.1149 |  |  |  | -24.723
103 | 20 | 3 | CC3 16sep24 BOT |  | 1 | 0 | -15.697 |  | -15.7002 |  |  | -15.6241 |  |  |  | -25.2115
104 | 20 | 4 | CC3 16sep24 BOT |  | 1 | 0 | -15.559 | 0.22454 | -15.5593 | 0.23643 |  | -15.4824 |  |  |  | -25.0756
105 | 21 | 1 | CC3 17feb25 0.1m |  | 1 | 0 | -21.249 |  | -22.1496 |  |  | -22.0719 |  |  |  | -31.3982
106 | 21 | 2 | CC3 17feb25 0.1m |  | 1 | 0 | -22.143 |  | -22.4369 |  |  | -22.3585 |  |  |  | -31.6731
107 | 21 | 3 | CC3 17feb25 0.1m |  | 1 | 0 | -22.31 |  | -22.4434 |  |  | -22.3643 |  |  |  | -31.6786
108 | 21 | 4 | CC3 17feb25 0.1m |  | 1 | 0 | -22.11 | 0.47743 | -22.1799 | 0.15948 |  | -22.1 |  |  |  | -31.4251
109 | 22 | 1 | CC3 16sep24 0.1m |  | 1 | 0 | -16.009 |  | -15.0433 |  |  | -14.9627 |  |  |  | -24.577
110 | 22 | 2 | CC3 16sep24 0.1m |  | 1 | 0 | -15.64 |  | -15.3512 |  |  | -15.2699 |  |  |  | -24.8717
111 | 22 | 3 | CC3 16sep24 0.1m |  | 1 | 0 | -15.91 |  | -15.7875 |  |  | -15.7054 |  |  |  | -25.2896
112 | 22 | 4 | CC3 16sep24 0.1m |  | 1 | 0 | -15.415 | 0.26885 | -15.3435 | 0.30634 |  | -15.2607 |  |  |  | -24.8629
113 | 23 | 1 | CS2 17feb25 0.1m |  | 1 | 0 | -26.981 |  | -28.8116 |  |  | -28.7281 |  |  |  | -37.7846
114 | 23 | 2 | CS2 17feb25 0.1m |  | 1 | 0 | -28.336 |  | -28.9127 |  |  | -28.8284 |  |  |  | -37.8809
115 | 23 | 3 | CS2 17feb25 0.1m |  | 1 | 0 | -28.57 |  | -28.8299 |  |  | -28.7449 |  |  |  | -37.8007
116 | 23 | 4 | CS2 17feb25 0.1m |  | 1 | 0 | -28.85 | 0.82923 | -28.9934 | 0.083506 |  | -28.9076 |  |  |  | -37.9569
117 | 24 | 1 | C50 17feb25 0.1m |  | 1 | 0 | -22.807 |  | -21.8505 |  |  | -21.764 |  |  |  | -31.1027
118 | 24 | 2 | C50 17feb25 0.1m |  | 1 | 0 | -22.153 |  | -21.8541 |  |  | -21.7668 |  |  |  | -31.1054
119 | 24 | 3 | C50 17feb25 0.1m |  | 1 | 0 | -21.759 |  | -21.6189 |  |  | -21.5309 |  |  |  | -30.879
120 | 24 | 4 | C50 17feb25 0.1m |  | 1 | 0 | -21.935 | 0.45827 | -21.8612 | 0.11827 |  | -21.7724 |  |  |  | -31.1108
121 | 25 | 1 | C50 17feb25 9m - DUP |  | 1 | 0 | -21.884 |  | -21.8759 |  |  | -21.7864 |  |  |  | -31.1242
122 | 25 | 2 | C50 17feb25 9m - DUP |  | 1 | 0 | -21.563 |  | -21.5464 |  |  | -21.4562 |  |  |  | -30.8073
123 | 25 | 3 | C50 17feb25 9m - DUP |  | 1 | 0 | -21.798 |  | -21.7953 |  |  | -21.7043 |  |  |  | -31.0454
124 | 25 | 4 | C50 17feb25 9m - DUP |  | 1 | 0 | -22.079 | 0.21387 | -22.0805 | 0.2209 |  | -21.9888 |  |  |  | -31.3184
125 | 26 | 1 | Hawaii |  | 1 | 0 | -5.378 |  | -2.7346 |  |  | -2.6422 |  |  |  | -12.7556
126 | 26 | 2 | Hawaii |  | 1 | 0 | -3.309 |  | -2.4713 |  |  | -2.3781 |  |  |  | -12.5022
127 | 26 | 3 | Hawaii |  | 1 | 0 | -3.102 |  | -2.727 |  |  | -2.6331 |  |  |  | -12.7469
128 | 26 | 4 | Hawaii |  | 1 | 0 | -2.927 | 1.1434 | -2.7226 | 0.1285 |  | -2.6279 |  |  |  | -12.7419
129 | 27 | 1 | Blacksburg | Drift | 1 | 0 | -34.482 |  | -39.4764 |  | -39.4764 | -39.381 | -39.381 |  |  | -48.006
130 | 27 | 2 | Blacksburg | Drift | 1 | 0 | -38.191 |  | -39.7649 |  | -39.7649 | -39.6688 | -39.6688 |  |  | -48.282
131 | 27 | 3 | Blacksburg | Drift | 1 | 0 | -38.484 |  | -39.1865 |  | -39.1865 | -39.0897 | -39.0897 |  |  | -47.7264
132 | 27 | 4 | Blacksburg | Drift | 1 | 0 | -38.573 | 1.9738 | -38.9535 | 0.35213 | -38.9535 | -38.8559 | -38.8559 |  |  | -47.5021
Note on G2: Manually change first six injections to ignore "-1".
Note on G3: Manually change first six injections to ignore "-1".
Note on G4: Manually change first six injections to ignore "-1".
Note on G5: Manually change first six injections to ignore "-1".
Note on G6: Manually change first six injections to ignore "-1".
Note on G7: Manually change first six injections to ignore "-1".

## RESULTS
| Analyst: | DWH |  |  |  | Complete Sequence |  |  | Run Summary | 18O | 18O | 2H | 2H
| Analysis Date: | 2025-03-18 |  |  |  | Line | Inj Number | Identifier | Ignore | Ignore | d18O Final Value | d2H Final Value | Ignore
| Project: | CCR |  |  |  | 1 | 1 | Blacksburg | 0 | -1 |  |  | -1
[6 empty] | 2 | 2 | Blacksburg | 0 | -1 |  |  | -1
[6 empty] | 3 | 3 | Blacksburg | 0 | -1 |  |  | -1
Vial | Sample Name | d18O | 1 s.d. | d2H | 1 s.d. | 4 | 4 | Blacksburg | 0 | -1 |  |  | -1
|  | ‰ vs. VSMOW |  | ‰ vs. VSMOW |  | 5 | 5 | Blacksburg | 0 | -1 |  |  | -1
1 | Blacksburg | -7.4486 | 0.02238 | -47.9279 | 0.082062 | 6 | 6 | Blacksburg | 0 | -1 |  |  | -1
2 | Myrtle | -2.271 | 0.035416 | -10.3496 | 0.16835 | 7 | 7 | Blacksburg | 0 | 0 | -7.4524 | -48.0374 | 0
3 | Homer | -14.5237 | 0.036853 | -110.4537 | 0.16266 | 8 | 8 | Blacksburg | 0 | 0 | -7.4759 | -47.9139 | 0
4 | Blacksburg | -7.5976 | 0.069432 | -47.6125 | 0.2664 | 9 | 9 | Blacksburg | 0 | 0 | -7.4441 | -47.9218 | 0
5 | Hawaii | -3.8025 | 0.058153 | -12.4828 | 0.12813 | 10 | 10 | Blacksburg | 0 | 0 | -7.4218 | -47.8386 | 0
6 | CC4 30sep24 BOT | -5.3643 | 0.066449 | -31.0129 | 0.23045 | 11 | 1 | Myrtle | 0 | 0 | -2.2317 | -10.501 | 0
7 | CS1 28oct24 0.1m | -6.3942 | 0.047662 | -34.5357 | 0.11208 | 12 | 2 | Myrtle | 0 | 0 | -2.3302 | -10.4381 | 0
8 | CC2 28oct24 BOT | -4.5004 | 0.051209 | -25.283 | 0.1448 | 13 | 3 | Myrtle | 0 | 0 | -2.3223 | -10.3994 | 0
9 | Blacksburg | -7.8737 | 0.015802 | -47.3822 | 0.12493 | 14 | 4 | Myrtle | 0 | 0 | -2.2567 | -10.6604 | 0
10 | CC4 30sep24 1.5m | -4.1748 | 0.12656 | -24.4496 | 0.39994 | 15 | 5 | Myrtle | 0 | 0 | -2.2616 | -10.3626 | 0
11 | CC2 30sep24 BOT | -4.2263 | 0.052837 | -24.7204 | 0.10078 | 16 | 6 | Myrtle | 0 | 0 | -2.2904 | -10.2985 | 0
12 | CCR rain 30sep24 | -4.6069 | 0.059903 | -19.2003 | 0.13737 | 17 | 7 | Myrtle | 0 | 0 | -2.2468 | -10.0504 | 0
13 | CC4 30sep24 0.1m | -4.0713 | 0.03144 | -24.7779 | 0.053301 | 18 | 8 | Myrtle | 0 | 0 | -2.2493 | -10.3171 | 0
14 | CC3 30sep24 BOT | -4.0421 | 0.047306 | -24.8664 | 0.044796 | 19 | 9 | Myrtle | 0 | 0 | -2.2316 | -10.2848 | 0
15 | CC4 16sep24 0.1m | -3.8898 | 0.041661 | -24.9343 | 0.24206 | 20 | 10 | Myrtle | 0 | 0 | -2.2891 | -10.1833 | 0
16 | C50 17feb25 9m | -5.263 | 0.036546 | -30.9514 | 0.20904 | 21 | 1 | Homer | 0 | 0 | -14.4329 | -110.0816 | 0
17 | CC3 16sep24 1.5m | -3.7907 | 0.056572 | -24.8491 | 0.13303 | 22 | 2 | Homer | 0 | 0 | -14.5281 | -110.4737 | 0
18 | Blacksburg | -7.6002 | 0.061633 | -47.7189 | 0.19382 | 23 | 3 | Homer | 0 | 0 | -14.5289 | -110.5995 | 0
19 | CC3 30sep24 0.1m | -3.9044 | 0.035064 | -24.7306 | 0.2144 | 24 | 4 | Homer | 0 | 0 | -14.5143 | -110.4607 | 0
20 | CC3 16sep24 BOT | -3.8123 | 0.057589 | -25.0509 | 0.22678 | 25 | 5 | Homer | 0 | 0 | -14.5354 | -110.5144 | 0
21 | CC3 17feb25 0.1m | -5.2801 | 0.034413 | -31.5438 | 0.15295 | 26 | 6 | Homer | 0 | 0 | -14.5224 | -110.5285 | 0
22 | CC3 16sep24 0.1m | -3.7246 | 0.022753 | -24.9003 | 0.29342 | 27 | 7 | Homer | 0 | 0 | -14.5191 | -110.4103 | 0
23 | CS2 17feb25 0.1m | -6.702 | 0.041868 | -37.8558 | 0.07947 | 28 | 8 | Homer | 0 | 0 | -14.5763 | -110.6451 | 0
24 | C50 17feb25 0.1m | -5.0809 | 0.036265 | -31.0495 | 0.11369 | 29 | 9 | Homer | 0 | 0 | -14.5251 | -110.5273 | 0
25 | C50 17feb25 9m - DUP | -5.0781 | 0.047108 | -31.0739 | 0.21149 | 30 | 10 | Homer | 0 | 0 | -14.554 | -110.2963 | 0
26 | Hawaii | -3.5789 | 0.059122 | -12.6867 | 0.1231 | 31 | 1 | Blacksburg | 0 | 0 | -7.4707 | -46.9366 | 0
27 | Blacksburg | -7.5007 | 0.073609 | -47.8791 | 0.33858 | 32 | 2 | Blacksburg | 0 | 0 | -7.5622 | -47.5916 | 0
[6 empty] | 33 | 3 | Blacksburg | 0 | 0 | -7.5694 | -47.8169 | 0
[6 empty] | 34 | 4 | Blacksburg | 0 | 0 | -7.5919 | -47.4398 | 0
[6 empty] | 35 | 5 | Blacksburg | 0 | 0 | -7.6252 | -47.7019 | 0
| Quality Assurance Verification |  |  |  |  | 36 | 6 | Blacksburg | 0 | 0 | -7.581 | -47.7269 | 0
| Metric |  | Reviewer 1 | Reviewer 2 |  | 37 | 7 | Blacksburg | 0 | 0 | -7.5487 | -47.7599 | 0
| d18O Cal Curve R^2 | 0.999+ |  |  |  | 38 | 8 | Blacksburg | 0 | 0 | -7.6893 | -47.6422 | 0
| d18O QC checks | < 0.15 |  |  |  | 39 | 9 | Blacksburg | 0 | 0 | -7.7091 | -47.6415 | 0
| dD Cal Curve R^2 | 0.999+ |  |  |  | 40 | 10 | Blacksburg | 0 | 0 | -7.6283 | -47.8682 | 0
| dD QC checks | < 1.5 |  |  |  | 41 | 1 | Hawaii | 0 | 0 | -3.7378 | -12.5899 | 0
| Initials |  |  |  |  | 42 | 2 | Hawaii | 0 | 0 | -3.7702 | -12.3019 | 0
| Date |  |  |  |  | 43 | 3 | Hawaii | 0 | 0 | -3.8403 | -12.4856 | 0
| Comments: |  |  |  |  | 44 | 4 | Hawaii | 0 | 0 | -3.8615 | -12.5538 | 0
[6 empty] | 45 | 1 | CC4 30sep24 BOT | 0 | 0 | -5.3326 | -31.0999 | 0
[6 empty] | 46 | 2 | CC4 30sep24 BOT | 0 | 0 | -5.3489 | -31.1552 | 0
[6 empty] | 47 | 3 | CC4 30sep24 BOT | 0 | 0 | -5.3141 | -31.1275 | 0
[6 empty] | 48 | 4 | CC4 30sep24 BOT | 0 | 0 | -5.4617 | -30.6689 | 0
[6 empty] | 49 | 1 | CS1 28oct24 0.1m | 0 | 0 | -6.3943 | -34.6827 | 0
[6 empty] | 50 | 2 | CS1 28oct24 0.1m | 0 | 0 | -6.3576 | -34.5414 | 0
[6 empty] | 51 | 3 | CS1 28oct24 0.1m | 0 | 0 | -6.3635 | -34.4126 | 0
[6 empty] | 52 | 4 | CS1 28oct24 0.1m | 0 | 0 | -6.4615 | -34.5059 | 0
[6 empty] | 53 | 1 | CC2 28oct24 BOT | 0 | 0 | -4.4669 | -25.2407 | 0
[6 empty] | 54 | 2 | CC2 28oct24 BOT | 0 | 0 | -4.4865 | -25.4592 | 0
[6 empty] | 55 | 3 | CC2 28oct24 BOT | 0 | 0 | -4.4721 | -25.3187 | 0
[6 empty] | 56 | 4 | CC2 28oct24 BOT | 0 | 0 | -4.5763 | -25.1133 | 0
[6 empty] | 57 | 1 | Blacksburg | 0 | 0 | -7.8686 | -47.545 | 0
[6 empty] | 58 | 2 | Blacksburg | 0 | 0 | -7.8971 | -47.3867 | 0
[6 empty] | 59 | 3 | Blacksburg | 0 | 0 | -7.8632 | -47.355 | 0
[6 empty] | 60 | 4 | Blacksburg | 0 | 0 | -7.8658 | -47.2423 | 0
[6 empty] | 61 | 1 | CC4 30sep24 1.5m | 0 | 0 | -3.986 | -23.8706 | 0
[6 empty] | 62 | 2 | CC4 30sep24 1.5m | 0 | 0 | -4.25 | -24.4944 | 0
[6 empty] | 63 | 3 | CC4 30sep24 1.5m | 0 | 0 | -4.22 | -24.7176 | 0
[6 empty] | 64 | 4 | CC4 30sep24 1.5m | 0 | 0 | -4.2433 | -24.7157 | 0
[6 empty] | 65 | 1 | CC2 30sep24 BOT | 0 | 0 | -4.2397 | -24.6259 | 0
[6 empty] | 66 | 2 | CC2 30sep24 BOT | 0 | 0 | -4.1607 | -24.7022 | 0
[6 empty] | 67 | 3 | CC2 30sep24 BOT | 0 | 0 | -4.288 | -24.6907 | 0
[6 empty] | 68 | 4 | CC2 30sep24 BOT | 0 | 0 | -4.217 | -24.863 | 0
[6 empty] | 69 | 1 | CCR rain 30sep24 | 0 | 0 | -4.5722 | -19.3598 | 0
[6 empty] | 70 | 2 | CCR rain 30sep24 | 0 | 0 | -4.6966 | -19.0601 | 0
[6 empty] | 71 | 3 | CCR rain 30sep24 | 0 | 0 | -4.5766 | -19.266 | 0
[6 empty] | 72 | 4 | CCR rain 30sep24 | 0 | 0 | -4.5824 | -19.1154 | 0
[6 empty] | 73 | 1 | CC4 30sep24 0.1m | 0 | 0 | -4.0785 | -24.7477 | 0
[6 empty] | 74 | 2 | CC4 30sep24 0.1m | 0 | 0 | -4.0551 | -24.8567 | 0
[6 empty] | 75 | 3 | CC4 30sep24 0.1m | 0 | 0 | -4.112 | -24.7638 | 0
[6 empty] | 76 | 4 | CC4 30sep24 0.1m | 0 | 0 | -4.0397 | -24.7434 | 0
[6 empty] | 77 | 1 | CC3 30sep24 BOT | 0 | 0 | -3.9886 | -24.8327 | 0
[6 empty] | 78 | 2 | CC3 30sep24 BOT | 0 | 0 | -4.0209 | -24.9317 | 0
[6 empty] | 79 | 3 | CC3 30sep24 BOT | 0 | 0 | -4.0617 | -24.8585 | 0
[6 empty] | 80 | 4 | CC3 30sep24 BOT | 0 | 0 | -4.097 | -24.8427 | 0
[6 empty] | 81 | 1 | CC4 16sep24 0.1m | 0 | 0 | -3.862 | -24.7136 | 0
[6 empty] | 82 | 2 | CC4 16sep24 0.1m | 0 | 0 | -3.8677 | -25.2355 | 0
[6 empty] | 83 | 3 | CC4 16sep24 0.1m | 0 | 0 | -3.9515 | -25.0228 | 0
[6 empty] | 84 | 4 | CC4 16sep24 0.1m | 0 | 0 | -3.8781 | -24.7654 | 0
[6 empty] | 85 | 1 | C50 17feb25 9m | 0 | 0 | -5.2363 | -31.0112 | 0
[6 empty] | 86 | 2 | C50 17feb25 9m | 0 | 0 | -5.3082 | -31.2186 | 0
[6 empty] | 87 | 3 | C50 17feb25 9m | 0 | 0 | -5.2305 | -30.8245 | 0
[6 empty] | 88 | 4 | C50 17feb25 9m | 0 | 0 | -5.277 | -30.7514 | 0
[6 empty] | 89 | 1 | CC3 16sep24 1.5m | 0 | 0 | -3.7079 | -24.6839 | 0
[6 empty] | 90 | 2 | CC3 16sep24 1.5m | 0 | 0 | -3.8093 | -25.0096 | 0
[6 empty] | 91 | 3 | CC3 16sep24 1.5m | 0 | 0 | -3.8355 | -24.8545 | 0
[6 empty] | 92 | 4 | CC3 16sep24 1.5m | 0 | 0 | -3.8103 | -24.8485 | 0
[6 empty] | 93 | 1 | Blacksburg | 0 | 0 | -7.6555 | -48.0044 | 0
[6 empty] | 94 | 2 | Blacksburg | 0 | 0 | -7.5947 | -47.6751 | 0
[6 empty] | 95 | 3 | Blacksburg | 0 | 0 | -7.6347 | -47.5932 | 0
[6 empty] | 96 | 4 | Blacksburg | 0 | 0 | -7.5159 | -47.6028 | 0
[6 empty] | 97 | 1 | CC3 30sep24 0.1m | 0 | 0 | -3.8746 | -24.4575 | 0
[6 empty] | 98 | 2 | CC3 30sep24 0.1m | 0 | 0 | -3.9247 | -24.6865 | 0
[6 empty] | 99 | 3 | CC3 30sep24 0.1m | 0 | 0 | -3.9434 | -24.9634 | 0
[6 empty] | 100 | 4 | CC3 30sep24 0.1m | 0 | 0 | -3.8749 | -24.8151 | 0
[6 empty] | 101 | 1 | CC3 16sep24 BOT | 0 | 0 | -3.8704 | -25.1936 | 0
[6 empty] | 102 | 2 | CC3 16sep24 BOT | 0 | 0 | -3.821 | -24.723 | 0
[6 empty] | 103 | 3 | CC3 16sep24 BOT | 0 | 0 | -3.825 | -25.2115 | 0
[6 empty] | 104 | 4 | CC3 16sep24 BOT | 0 | 0 | -3.7327 | -25.0756 | 0
[6 empty] | 105 | 1 | CC3 17feb25 0.1m | 0 | 0 | -5.3138 | -31.3982 | 0
[6 empty] | 106 | 2 | CC3 17feb25 0.1m | 0 | 0 | -5.2646 | -31.6731 | 0
[6 empty] | 107 | 3 | CC3 17feb25 0.1m | 0 | 0 | -5.2392 | -31.6786 | 0
[6 empty] | 108 | 4 | CC3 17feb25 0.1m | 0 | 0 | -5.3025 | -31.4251 | 0
[6 empty] | 109 | 1 | CC3 16sep24 0.1m | 0 | 0 | -3.6919 | -24.577 | 0
[6 empty] | 110 | 2 | CC3 16sep24 0.1m | 0 | 0 | -3.7441 | -24.8717 | 0
[6 empty] | 111 | 3 | CC3 16sep24 0.1m | 0 | 0 | -3.7287 | -25.2896 | 0
[6 empty] | 112 | 4 | CC3 16sep24 0.1m | 0 | 0 | -3.7339 | -24.8629 | 0
[6 empty] | 113 | 1 | CS2 17feb25 0.1m | 0 | 0 | -6.7603 | -37.7846 | 0
[6 empty] | 114 | 2 | CS2 17feb25 0.1m | 0 | 0 | -6.6751 | -37.8809 | 0
[6 empty] | 115 | 3 | CS2 17feb25 0.1m | 0 | 0 | -6.6684 | -37.8007 | 0
[6 empty] | 116 | 4 | CS2 17feb25 0.1m | 0 | 0 | -6.7044 | -37.9569 | 0
[6 empty] | 117 | 1 | C50 17feb25 0.1m | 0 | 0 | -5.0935 | -31.1027 | 0
[6 empty] | 118 | 2 | C50 17feb25 0.1m | 0 | 0 | -5.0275 | -31.1054 | 0
[6 empty] | 119 | 3 | C50 17feb25 0.1m | 0 | 0 | -5.0937 | -30.879 | 0
[6 empty] | 120 | 4 | C50 17feb25 0.1m | 0 | 0 | -5.1087 | -31.1108 | 0
[6 empty] | 121 | 1 | C50 17feb25 9m - DUP | 0 | 0 | -5.0325 | -31.1242 | 0
[6 empty] | 122 | 2 | C50 17feb25 9m - DUP | 0 | 0 | -5.1306 | -30.8073 | 0
[6 empty] | 123 | 3 | C50 17feb25 9m - DUP | 0 | 0 | -5.1048 | -31.0454 | 0
[6 empty] | 124 | 4 | C50 17feb25 9m - DUP | 0 | 0 | -5.0447 | -31.3184 | 0
[6 empty] | 125 | 1 | Hawaii | 0 | 0 | -3.5385 | -12.7556 | 0
[6 empty] | 126 | 2 | Hawaii | 0 | 0 | -3.5534 | -12.5022 | 0
[6 empty] | 127 | 3 | Hawaii | 0 | 0 | -3.6668 | -12.7469 | 0
[6 empty] | 128 | 4 | Hawaii | 0 | 0 | -3.557 | -12.7419 | 0
[6 empty] | 129 | 1 | Blacksburg | 0 | 0 | -7.5785 | -48.006 | 0
[6 empty] | 130 | 2 | Blacksburg | 0 | 0 | -7.4137 | -48.282 | 0
[6 empty] | 131 | 3 | Blacksburg | 0 | 0 | -7.5413 | -47.7264 | 0
[6 empty] | 132 | 4 | Blacksburg | 0 | 0 | -7.4693 | -47.5021 | 0