## Instructions
Picarro Isotope Analyzer
Data Processor 3.0 (KAP)
How to Use this Spreadsheet:
Step 1: Retrieve raw result file from Piccaro CPU (C:Piccaro\IsotopeData). Add initials and project name to end
of file name. Copy file into your folder on Water Server.
Step 2: On syncing computer, move raw file into Raw Data folder on Google Drive:
Instruments and Analyses --> Picarro Water Isotope Analyzer --> Raw Data
Step 3: Save this spreadsheet in the Processed Data folder on Google Drive:
Instruments and Analyses --> Picarro Water Isotope Analyzer --> Processed Data
Use file name: YYYYMMDD_Initials_(anything else to identify ex. project, run number, etc.)
Step 4: Enter run information and sample positions in the "Tray Configuration" tab. (orange cells)
(It is important to use the same names for the standards throughout this spreadsheet)
Step 5: Copy text from raw result file and paste into "Picarro Output" tab.
Step 6: In Run Summary tab, check that H2O Mean is within "good" bounds. (between 17,000 - 23,000)
Step 7: In 18O tab, verify that the reference defined values match the references used in run (cells U39 - 42). Run Solver and
verify stdev memory corrected are reduced (column K values less than corresponding column I values). Verify stdev of
drift corrected reference is lower than memory corrected. Verify normalization R2 is 0.99+. Verify QC checks are within
0.15 tolerance.
Step 8: In dD tab, verify that the reference defined values match the references used in run (cells U39 - 42. Run Solver and
verify stdev memory corrected are reduced (column K values less than corresponding column I values). Verify stdev of
drift corrected reference is lower than memory corrected. Verify normalization R2 is 0.99+. Verify QC checks are within
1.5 tolerance.
Step 9: Fill in Quality Assurance Verification chart on RESULTS tab. Dr. Scott or Kelly Peeler will fill in as reviewer 2.
Step 10: Add reference results to Water Isotope Analyzer run log.

## Tray Configuration
| Analyst |  | DWH
| PI |  | Durelle Scott
| Project |  | CCR
| Sampling Date |  | 2024 - 2025
| Sampling Location |  | Carvins Cove Reservoir
| Analysis Date |  | 2025-05-28
| Notes |  | raw data file shows file date as 2006 (I forgot to change the clock on the Picarro), but data in file shows correct date
Rear Rack
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
1 | Blacksburg | Myrtle | Homer | Blacksburg | Hawaii | CP1 22may24 BOT | C50 16sep24 0.1m | CC2 19jun24 0.1m | Blacksburg | 9
10 | CP1 28oct24 0.1m | CC4 11jul24 9m | CC4 15aug24 BOT | C50 16apr25 0.1m | C50 17feb25 1.5m | C50 11jul24 9m | CC4 22may24 0.1m | C50 22may24 0.1m | Blacksburg | 18
19 | CS2 22may24 0.1m | CCR raindeploy - 11jul24 R1 | C50 16apr25 6m | CC4 16sep24 1.5m | C50 19jun24 1.5m | CP2 22may24 0.1m | CC4 22may24 0.1m - DUP | Hawaii | Blacksburg | 27
28 | [9 empty] | 36
37 | [9 empty] | 45
46 | [9 empty] | 54
Rear Rack
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
1 | Drift | High reference | Low Reference | Drift | QC Check | Sample | Sample | Sample | Drift | 9
10 | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Drift | 18
19 | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Drift | 27
28 | [9 empty] | 36
37 | [9 empty] | 45
46 | [9 empty] | 54
Standards
|  | Standard ID | Actual d(18_16) | Actual d(D_H)
| #1 | Denali | -19.9287 | -153.7172 | [6 empty] | -3.37
| #2 | Alberta | -18.8392 | -146.8245 | [6 empty] | -20.92
| #3 | MT BHB | -18.431 | -140.5778 | [6 empty] | -17.35
| #4 | Anchorage | -16.4204 | -129.9028 | [6 empty] | -7.39
| #5 | Homer | -14.5399 | -110.5169
| #6 | MT Creek | -13.3562 | -97.7428
| #7 | Nickel Mines | -7.876 | -48.7951
| #8 | Blacksburg | -7.5615 | -47.446
| #9 | Denton's Well | -6.9583 | -42.6011
| #10 | Hawaii | -3.6108 | -13.2294
| #11 | Myrtle Beach | -2.2908 | -10.453
| #12 | Heavy Oxygen | 3.1753 | -15.7881
| #13 | Heavy Hydrogen | 11.1372 | 4.0332
| #14 | Andres Heavy Water | 17.5664 | 20.5382
| #15 | Picarro ZERO | 0.3 | 1.8
| #16 | Picarro MID | -20.6 | -159
| #17 | Picarro DEPL | -29.6 | -235
#18
#19
#20
|  | DI | -6 | -36
Note on C34: Nate: The Standard ID must match the Standard ID used above in the trey protocol. For example, if you are using the Boulder Colorado Standard, and you labeled it "B" above, it must be labled "B" in this collum.

## Timing
| Enter start time and date --> | 2019-01-10 09:30
Tray Position | Sample ID | Start Time | # Injections
1 | Blacksburg | 2019-01-10 10:58 | 10
2 | Myrtle | 2019-01-10 12:27 | 10
3 | Homer | 2019-01-10 13:55 | 10
4 | Blacksburg | 2019-01-10 15:23 | 10
5 | Hawaii | 2019-01-10 15:59 | 4
6 | CP1 22may24 BOT | 2019-01-10 16:34 | 4
7 | C50 16sep24 0.1m | 2019-01-10 17:09 | 4
8 | CC2 19jun24 0.1m | 2019-01-10 17:45 | 4
9 | Blacksburg | 2019-01-10 18:20 | 4
10 | CP1 28oct24 0.1m | 2019-01-10 18:55 | 4
11 | CC4 11jul24 9m | 2019-01-10 19:31 | 4
12 | CC4 15aug24 BOT | 2019-01-10 20:06 | 4
13 | C50 16apr25 0.1m | 2019-01-10 20:41 | 4
14 | C50 17feb25 1.5m | 2019-01-10 21:17 | 4
15 | C50 11jul24 9m | 2019-01-10 21:52 | 4
16 | CC4 22may24 0.1m | 2019-01-10 22:27 | 4
17 | C50 22may24 0.1m | 2019-01-10 23:03 | 4
18 | Blacksburg | 2019-01-10 23:38 | 4
19 | CS2 22may24 0.1m | 2019-01-11 00:13 | 4
20 | CCR raindeploy - 11jul24 R1 | 2019-01-11 00:49 | 4
21 | C50 16apr25 6m | 2019-01-11 01:24 | 4
22 | CC4 16sep24 1.5m | 2019-01-11 01:59 | 4
23 | C50 19jun24 1.5m | 2019-01-11 02:35 | 4
24 | CP2 22may24 0.1m | 2019-01-11 03:10 | 4
25 | CC4 22may24 0.1m - DUP | 2019-01-11 03:45 | 4
26 | Hawaii | 2019-01-11 04:21 | 4
27 | Blacksburg | 2019-01-11 04:56 | 4
Formulas:
Start Time = [Start Time, row above]+"0:08:50"*10

## Picarro Output
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | days_Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | BASELINE_Mean | BASELINE_SD | DAS Temp | Tray | Sample | Job | Method
1 | P-11728 | 2025/05/28 10:22:30 | MT1-Rear-01 | 1 | -8.596 | -46.973 | 24311 | -1 | 0 |  |  | H2O | 0.93281 | 0.142 | 0.371 | 113 | 0.034 | -0.072 | 34.09 | 845 | 0 | 30.563 | MT1-Rear | 1 | 66 | BSE
2 | P-11728 | 2025/05/28 10:31:16 | MT1-Rear-01 | 2 | -9.022 | -45.287 | 19602 | -1 | 1 |  |  | H2O | 0.93894 | 0.133 | 0.294 | 105 | -0.015 | -0.037 | 31.63 | 844 | 0 | 30.688 | MT1-Rear | 1 | 66 | BSE
3 | P-11728 | 2025/05/28 10:40:03 | MT1-Rear-01 | 3 | -8.849 | -44.776 | 19434 | -1 | 1 |  |  | H2O | 0.94491 | 0.11 | 0.427 | 103 | -0.005 | -0.039 | 30.91 | 844 | 0 | 30.75 | MT1-Rear | 1 | 66 | BSE
4 | P-11728 | 2025/05/28 10:48:49 | MT1-Rear-01 | 4 | -8.826 | -44.877 | 19255 | 0 | 1 |  |  | H2O | 0.95103 | 0.201 | 0.366 | 105 | -0.036 | -0.04 | 31.58 | 844 | 0 | 30.75 | MT1-Rear | 1 | 66 | BSE
5 | P-11728 | 2025/05/28 10:57:36 | MT1-Rear-01 | 5 | -8.767 | -44.464 | 19549 | 0 | 1 |  |  | H2O | 0.95715 | 0.151 | 0.325 | 103 | -0.038 | -0.017 | 31.11 | 844 | 0 | 30.625 | MT1-Rear | 1 | 66 | BSE
6 | P-11728 | 2025/05/28 11:06:22 | MT1-Rear-01 | 6 | -8.797 | -44.194 | 19108 | 0 | 1 |  |  | H2O | 0.96327 | 0.121 | 0.263 | 105 | -0.007 | -0.037 | 31.53 | 844 | 0 | 30.625 | MT1-Rear | 1 | 66 | BSE
7 | P-11728 | 2025/05/28 11:15:09 | MT1-Rear-01 | 7 | -8.857 | -43.957 | 19370 | 0 | 1 |  |  | H2O | 0.9694 | 0.129 | 0.286 | 106 | -0.023 | -0.025 | 31.95 | 844 | 0 | 30.5 | MT1-Rear | 1 | 66 | BSE
8 | P-11728 | 2025/05/28 11:23:55 | MT1-Rear-01 | 8 | -8.848 | -44.191 | 19353 | 0 | 1 |  |  | H2O | 0.97545 | 0.098 | 0.417 | 96 | -0.018 | -0.07 | 31.86 | 844 | 0 | 30.625 | MT1-Rear | 1 | 66 | BSE
9 | P-11728 | 2025/05/28 11:32:42 | MT1-Rear-01 | 9 | -8.762 | -43.925 | 19494 | 0 | 1 |  |  | H2O | 0.9815 | 0.13 | 0.242 | 103 | -0.012 | -0.02 | 31.14 | 844 | 0 | 30.688 | MT1-Rear | 1 | 66 | BSE
10 | P-11728 | 2025/05/28 11:41:29 | MT1-Rear-01 | 10 | -8.889 | -43.99 | 19311 | 0 | 1 |  |  | H2O | 0.98762 | 0.115 | 0.284 | 105 | -0.013 | -0.006 | 31.5 | 844 | 0 | 30.813 | MT1-Rear | 1 | 66 | BSE
11 | P-11729 | 2025/05/28 11:50:16 | MT1-Rear-02 | 1 | -3.8 | -9.933 | 19440 | -1 | 1 |  |  | H2O | 0.99375 | 0.137 | 0.258 | 107 | -0.031 | 0.04 | 32.22 | 844 | 0 | 30.563 | MT1-Rear | 2 | 66 | BSE
12 | P-11729 | 2025/05/28 11:59:02 | MT1-Rear-02 | 2 | -3.403 | -6.116 | 19531 | -1 | 1 |  |  | H2O | 0.99988 | 0.144 | 0.652 | 104 | -0.032 | -0.154 | 31.18 | 844 | 0 | 30.5 | MT1-Rear | 2 | 66 | BSE
13 | P-11729 | 2025/05/28 12:07:49 | MT1-Rear-02 | 3 | -3.42 | -5.319 | 21298 | -1 | 1 |  |  | H2O | 1.006 | 0.135 | 0.303 | 102 | -0.005 | 0.005 | 33.59 | 844 | 0 | 30.75 | MT1-Rear | 2 | 66 | BSE
14 | P-11729 | 2025/05/28 12:16:35 | MT1-Rear-02 | 4 | -3.315 | -4.857 | 19506 | 0 | 1 |  |  | H2O | 1.012 | 0.116 | 0.308 | 103 | -0.014 | -0.052 | 30.89 | 844 | 0 | 30.625 | MT1-Rear | 2 | 66 | BSE
15 | P-11729 | 2025/05/28 12:25:22 | MT1-Rear-02 | 5 | -3.299 | -4.853 | 19560 | 0 | 1 |  |  | H2O | 1.0182 | 0.124 | 0.402 | 105 | -0.027 | -0.102 | 31.69 | 844 | 0 | 30.625 | MT1-Rear | 2 | 66 | BSE
16 | P-11729 | 2025/05/28 12:34:09 | MT1-Rear-02 | 6 | -3.228 | -4.782 | 19611 | 0 | 1 |  |  | H2O | 1.0243 | 0.117 | 0.39 | 103 | -0.002 | -0.009 | 31.18 | 844 | 0 | 30.75 | MT1-Rear | 2 | 66 | BSE
17 | P-11729 | 2025/05/28 12:42:55 | MT1-Rear-02 | 7 | -3.338 | -4.586 | 19464 | 0 | 1 |  |  | H2O | 1.0303 | 0.088 | 0.182 | 104 | 0.004 | -0.026 | 31.41 | 844 | 0 | 30.625 | MT1-Rear | 2 | 66 | BSE
18 | P-11729 | 2025/05/28 12:51:42 | MT1-Rear-02 | 8 | -3.336 | -4.557 | 19840 | 0 | 1 |  |  | H2O | 1.0364 | 0.124 | 0.408 | 106 | -0.023 | 0.005 | 31.94 | 844 | 0 | 30.625 | MT1-Rear | 2 | 66 | BSE
19 | P-11729 | 2025/05/28 13:00:29 | MT1-Rear-02 | 9 | -3.36 | -4.473 | 19606 | 0 | 1 |  |  | H2O | 1.0426 | 0.104 | 0.275 | 105 | -0.001 | 0.005 | 31.77 | 844 | 0 | 30.75 | MT1-Rear | 2 | 66 | BSE
20 | P-11729 | 2025/05/28 13:09:16 | MT1-Rear-02 | 10 | -3.251 | -4.442 | 19782 | 0 | 1 |  |  | H2O | 1.0486 | 0.125 | 0.285 | 99 | -0.026 | 0.046 | 32.72 | 844 | 0 | 30.75 | MT1-Rear | 2 | 66 | BSE
21 | P-11730 | 2025/05/28 13:18:03 | MT1-Rear-03 | 1 | -14.813 | -93.78 | 19804 | -1 | 1 |  |  | H2O | 1.0547 | 0.1 | 0.558 | 104 | 0.005 | 0.091 | 31.33 | 844 | 0 | 30.688 | MT1-Rear | 3 | 66 | BSE
22 | P-11730 | 2025/05/28 13:26:49 | MT1-Rear-03 | 2 | -15.687 | -104.711 | 19503 | -1 | 1 |  |  | H2O | 1.0608 | 0.105 | 0.721 | 106 | 0.013 | 0.19 | 32.08 | 844 | 0 | 30.75 | MT1-Rear | 3 | 66 | BSE
23 | P-11730 | 2025/05/28 13:35:36 | MT1-Rear-03 | 3 | -15.812 | -106.635 | 19616 | -1 | 1 |  |  | H2O | 1.0669 | 0.109 | 0.387 | 94 | -0.015 | 0.084 | 30.9 | 844 | 0 | 30.875 | MT1-Rear | 3 | 66 | BSE
24 | P-11730 | 2025/05/28 13:44:23 | MT1-Rear-03 | 4 | -15.915 | -107.535 | 19830 | 0 | 1 |  |  | H2O | 1.073 | 0.076 | 0.596 | 108 | 0.002 | 0.173 | 32.42 | 844 | 0 | 30.563 | MT1-Rear | 3 | 66 | BSE
25 | P-11730 | 2025/05/28 13:53:11 | MT1-Rear-03 | 5 | -15.9 | -108.061 | 19529 | 0 | 1 |  |  | H2O | 1.0791 | 0.132 | 0.689 | 107 | 0.005 | 0.141 | 32.34 | 844 | 0 | 30.75 | MT1-Rear | 3 | 66 | BSE
26 | P-11730 | 2025/05/28 14:01:58 | MT1-Rear-03 | 6 | -15.887 | -108.108 | 19691 | 0 | 1 |  |  | H2O | 1.0853 | 0.109 | 0.382 | 104 | 0.024 | 0.069 | 31.46 | 844 | 0 | 31 | MT1-Rear | 3 | 66 | BSE
27 | P-11730 | 2025/05/28 14:10:45 | MT1-Rear-03 | 7 | -15.656 | -106.868 | 21834 | 0 | 1 |  |  | H2O | 1.0913 | 0.077 | 0.424 | 109 | -0.005 | 0.093 | 32.9 | 844 | 0 | 30.813 | MT1-Rear | 3 | 66 | BSE
28 | P-11730 | 2025/05/28 14:19:33 | MT1-Rear-03 | 8 | -15.943 | -108.093 | 20743 | 0 | 1 |  |  | H2O | 1.0975 | 0.093 | 0.377 | 110 | 0.002 | 0.067 | 33.1 | 844 | 0 | 30.75 | MT1-Rear | 3 | 66 | BSE
29 | P-11730 | 2025/05/28 14:28:21 | MT1-Rear-03 | 9 | -16.09 | -108.993 | 19279 | 0 | 1 |  |  | H2O | 1.1035 | 0.093 | 0.583 | 96 | 0 | 0.093 | 31.58 | 844 | 0 | 30.688 | MT1-Rear | 3 | 66 | BSE
30 | P-11730 | 2025/05/28 14:37:07 | MT1-Rear-03 | 10 | -15.946 | -108.845 | 19558 | 0 | 1 |  |  | H2O | 1.1096 | 0.12 | 0.273 | 105 | -0.013 | -0.014 | 31.7 | 844 | 0 | 30.75 | MT1-Rear | 3 | 66 | BSE
31 | P-11731 | 2025/05/28 14:45:55 | MT1-Rear-04 | 1 | -9.359 | -52.697 | 20179 | -1 | 1 |  |  | H2O | 1.1158 | 0.197 | 0.444 | 107 | -0.053 | 0.02 | 32.25 | 844 | 0 | 30.75 | MT1-Rear | 4 | 66 | BSE
32 | P-11731 | 2025/05/28 14:54:43 | MT1-Rear-04 | 2 | -9.056 | -46.746 | 19418 | -1 | 1 |  |  | H2O | 1.1218 | 0.126 | 0.516 | 102 | 0.018 | 0.137 | 33.6 | 844 | 0 | 30.75 | MT1-Rear | 4 | 66 | BSE
33 | P-11731 | 2025/05/28 15:03:29 | MT1-Rear-04 | 3 | -8.884 | -44.807 | 19651 | -1 | 1 |  |  | H2O | 1.1279 | 0.095 | 0.406 | 109 | -0.001 | 0.025 | 32.89 | 844 | 0 | 30.688 | MT1-Rear | 4 | 66 | BSE
34 | P-11731 | 2025/05/28 15:12:16 | MT1-Rear-04 | 4 | -8.87 | -44.788 | 19462 | 0 | 1 |  |  | H2O | 1.1341 | 0.13 | 0.262 | 105 | -0.021 | 0.049 | 31.63 | 844 | 0 | 30.75 | MT1-Rear | 4 | 66 | BSE
35 | P-11731 | 2025/05/28 15:21:04 | MT1-Rear-04 | 5 | -8.884 | -44.922 | 19795 | 0 | 1 |  |  | H2O | 1.1401 | 0.154 | 0.691 | 106 | 0.04 | 0.176 | 34.91 | 844 | 0 | 30.688 | MT1-Rear | 4 | 66 | BSE
36 | P-11731 | 2025/05/28 15:29:51 | MT1-Rear-04 | 6 | -8.742 | -44.182 | 19686 | 0 | 1 |  |  | H2O | 1.1462 | 0.088 | 0.357 | 107 | 0.002 | 0.078 | 32.3 | 844 | 0 | 30.438 | MT1-Rear | 4 | 66 | BSE
37 | P-11731 | 2025/05/28 15:38:39 | MT1-Rear-04 | 7 | -8.849 | -44.32 | 19361 | 0 | 1 |  |  | H2O | 1.1524 | 0.052 | 0.204 | 105 | 0.002 | 0.023 | 31.62 | 844 | 0 | 30.813 | MT1-Rear | 4 | 66 | BSE
38 | P-11731 | 2025/05/28 15:47:27 | MT1-Rear-04 | 8 | -8.824 | -44.19 | 19337 | 0 | 1 |  |  | H2O | 1.1585 | 0.106 | 0.301 | 102 | -0.012 | 0.01 | 30.83 | 844 | 0 | 30.688 | MT1-Rear | 4 | 66 | BSE
39 | P-11731 | 2025/05/28 15:56:15 | MT1-Rear-04 | 9 | -8.865 | -44.306 | 19655 | 0 | 1 |  |  | H2O | 1.1646 | 0.074 | 0.383 | 92 | -0.012 | -0.084 | 30.3 | 844 | 0 | 30.75 | MT1-Rear | 4 | 66 | BSE
40 | P-11731 | 2025/05/28 16:05:02 | MT1-Rear-04 | 10 | -8.855 | -44.238 | 19411 | 0 | 1 |  |  | H2O | 1.1707 | 0.081 | 0.372 | 102 | 0.006 | 0.078 | 30.65 | 844 | 0 | 30.75 | MT1-Rear | 4 | 66 | BSE
41 | P-11732 | 2025/05/28 16:13:50 | MT1-Rear-05 | 1 | -4.265 | -10.89 | 19802 | -1 | 1 |  |  | H2O | 1.1768 | 0.162 | 0.278 | 107 | -0.028 | -0.03 | 32.12 | 844 | 0 | 30.875 | MT1-Rear | 5 | 67 | BSE
42 | P-11732 | 2025/05/28 16:22:39 | MT1-Rear-05 | 2 | -3.984 | -6.882 | 19686 | -1 | 1 |  |  | H2O | 1.1829 | 0.142 | 0.475 | 94 | -0.021 | -0.139 | 30.94 | 844 | 0 | 30.813 | MT1-Rear | 5 | 67 | BSE
43 | P-11732 | 2025/05/28 16:31:27 | MT1-Rear-05 | 3 | -3.982 | -6.409 | 19227 | -1 | 1 |  |  | H2O | 1.189 | 0.105 | 0.257 | 105 | -0.021 | -0.041 | 31.58 | 844 | 0 | 30.875 | MT1-Rear | 5 | 67 | BSE
44 | P-11732 | 2025/05/28 16:40:15 | MT1-Rear-05 | 4 | -3.929 | -5.993 | 19644 | 0 | 1 |  |  | H2O | 1.1951 | 0.115 | 0.27 | 103 | -0.023 | -0.019 | 31.06 | 844 | 0 | 30.813 | MT1-Rear | 5 | 67 | BSE
45 | P-11733 | 2025/05/28 16:49:03 | MT1-Rear-06 | 1 | -6.922 | -27.362 | 19774 | -1 | 1 |  |  | H2O | 1.2013 | 0.1 | 0.365 | 104 | -0.009 | 0.08 | 31.49 | 844 | 0 | 30.688 | MT1-Rear | 6 | 67 | BSE
46 | P-11733 | 2025/05/28 16:57:51 | MT1-Rear-06 | 2 | -7.218 | -29.471 | 19344 | -1 | 1 |  |  | H2O | 1.2074 | 0.122 | 0.43 | 105 | -0.027 | -0.018 | 31.75 | 844 | 0 | 30.75 | MT1-Rear | 6 | 67 | BSE
47 | P-11733 | 2025/05/28 17:06:38 | MT1-Rear-06 | 3 | -7.142 | -30.48 | 19270 | -1 | 1 |  |  | H2O | 1.2134 | 0.131 | 0.325 | 103 | -0.033 | -0.035 | 30.89 | 844 | 0 | 30.875 | MT1-Rear | 6 | 67 | BSE
48 | P-11733 | 2025/05/28 17:15:27 | MT1-Rear-06 | 4 | -7.171 | -30.276 | 19505 | 0 | 1 |  |  | H2O | 1.2196 | 0.154 | 0.391 | 106 | -0.032 | -0.005 | 31.79 | 844 | 0 | 30.563 | MT1-Rear | 6 | 67 | BSE
49 | P-11734 | 2025/05/28 17:24:14 | MT1-Rear-07 | 1 | -4.995 | -21.571 | 19347 | -1 | 1 |  |  | H2O | 1.2257 | 0.121 | 0.284 | 104 | -0.023 | 0.035 | 31.3 | 844 | 0 | 30.688 | MT1-Rear | 7 | 67 | BSE
50 | P-11734 | 2025/05/28 17:33:02 | MT1-Rear-07 | 2 | -4.827 | -20.29 | 19785 | -1 | 1 |  |  | H2O | 1.2318 | 0.159 | 0.301 | 95 | -0.024 | 0.067 | 31.24 | 844 | 0 | 30.938 | MT1-Rear | 7 | 67 | BSE
51 | P-11734 | 2025/05/28 17:41:51 | MT1-Rear-07 | 3 | -4.781 | -20.114 | 19774 | -1 | 1 |  |  | H2O | 1.2379 | 0.098 | 0.354 | 102 | -0.019 | -0.017 | 30.58 | 844 | 0 | 30.875 | MT1-Rear | 7 | 67 | BSE
52 | P-11734 | 2025/05/28 17:50:39 | MT1-Rear-07 | 4 | -4.785 | -20.188 | 19799 | 0 | 1 |  |  | H2O | 1.2441 | 0.12 | 0.293 | 108 | 0.013 | 0.002 | 32.53 | 844 | 0 | 30.813 | MT1-Rear | 7 | 67 | BSE
53 | P-11735 | 2025/05/28 17:59:28 | MT1-Rear-08 | 1 | -5.712 | -23.364 | 19759 | -1 | 1 |  |  | H2O | 1.2502 | 0.078 | 0.438 | 96 | 0 | 0.083 | 31.71 | 844 | 0 | 31.063 | MT1-Rear | 8 | 67 | BSE
54 | P-11735 | 2025/05/28 18:08:16 | MT1-Rear-08 | 2 | -5.742 | -23.808 | 19757 | -1 | 1 |  |  | H2O | 1.2562 | 0.102 | 0.362 | 105 | -0.008 | 0.089 | 31.65 | 844 | 0 | 30.875 | MT1-Rear | 8 | 67 | BSE
55 | P-11735 | 2025/05/28 18:17:04 | MT1-Rear-08 | 3 | -5.734 | -23.751 | 19837 | -1 | 1 |  |  | H2O | 1.2624 | 0.119 | 0.342 | 101 | -0.019 | -0.077 | 30.51 | 844 | 0 | 30.938 | MT1-Rear | 8 | 67 | BSE
56 | P-11735 | 2025/05/28 18:25:53 | MT1-Rear-08 | 4 | -5.813 | -24.158 | 19517 | 0 | 1 |  |  | H2O | 1.2686 | 0.148 | 0.235 | 105 | -0.023 | 0.025 | 31.5 | 844 | 0 | 31.125 | MT1-Rear | 8 | 67 | BSE
57 | P-11736 | 2025/05/28 18:34:41 | MT1-Rear-09 | 1 | -8.516 | -40.729 | 19810 | -1 | 1 |  |  | H2O | 1.2746 | 0.139 | 0.401 | 96 | -0.024 | 0.087 | 31.66 | 844 | 0 | 31 | MT1-Rear | 9 | 67 | BSE
58 | P-11736 | 2025/05/28 18:43:30 | MT1-Rear-09 | 2 | -8.655 | -43.047 | 19849 | -1 | 1 |  |  | H2O | 1.2807 | 0.111 | 0.552 | 108 | 0.005 | 0.132 | 32.41 | 844 | 0 | 31 | MT1-Rear | 9 | 67 | BSE
59 | P-11736 | 2025/05/28 18:52:19 | MT1-Rear-09 | 3 | -8.712 | -43.539 | 19653 | -1 | 1 |  |  | H2O | 1.2869 | 0.122 | 0.337 | 104 | -0.006 | 0.028 | 31.31 | 844 | 0 | 31 | MT1-Rear | 9 | 67 | BSE
60 | P-11736 | 2025/05/28 19:01:08 | MT1-Rear-09 | 4 | -8.782 | -43.612 | 19694 | 0 | 1 |  |  | H2O | 1.293 | 0.178 | 0.269 | 104 | -0.024 | 0.037 | 31.25 | 844 | 0 | 30.875 | MT1-Rear | 9 | 67 | BSE
61 | P-11737 | 2025/05/28 19:09:58 | MT1-Rear-10 | 1 | -7.303 | -31.612 | 19448 | -1 | 1 |  |  | H2O | 1.2991 | 0.159 | 0.276 | 97 | -0.034 | 0.029 | 32.07 | 844 | 0 | 31.188 | MT1-Rear | 10 | 67 | BSE
62 | P-11737 | 2025/05/28 19:18:47 | MT1-Rear-10 | 2 | -7.235 | -30.47 | 19675 | -1 | 1 |  |  | H2O | 1.3053 | 0.116 | 0.329 | 98 | -0.019 | 0.089 | 32.22 | 844 | 0 | 31.063 | MT1-Rear | 10 | 67 | BSE
63 | P-11737 | 2025/05/28 19:27:37 | MT1-Rear-10 | 3 | -7.231 | -30.183 | 19631 | -1 | 1 |  |  | H2O | 1.3114 | 0.127 | 0.378 | 107 | 0.017 | 0.055 | 32.21 | 844 | 0 | 31 | MT1-Rear | 10 | 67 | BSE
64 | P-11737 | 2025/05/28 19:36:26 | MT1-Rear-10 | 4 | -7.123 | -29.866 | 20147 | 0 | 1 |  |  | H2O | 1.3175 | 0.125 | 0.509 | 109 | 0.016 | 0.114 | 32.91 | 844 | 0 | 30.875 | MT1-Rear | 10 | 67 | BSE
65 | P-11738 | 2025/05/28 19:45:16 | MT1-Rear-11 | 1 | -6.451 | -27.991 | 19728 | -1 | 1 |  |  | H2O | 1.3236 | 0.121 | 0.211 | 97 | -0.018 | 0.041 | 32.02 | 844 | 0 | 30.875 | MT1-Rear | 11 | 67 | BSE
66 | P-11738 | 2025/05/28 19:54:04 | MT1-Rear-11 | 2 | -6.405 | -27.818 | 19771 | -1 | 1 |  |  | H2O | 1.3298 | 0.097 | 0.201 | 94 | -0.02 | -0.014 | 31.12 | 844 | 0 | 31 | MT1-Rear | 11 | 67 | BSE
67 | P-11738 | 2025/05/28 20:02:53 | MT1-Rear-11 | 3 | -6.435 | -27.54 | 19926 | -1 | 1 |  |  | H2O | 1.3359 | 0.091 | 0.266 | 104 | -0.01 | 0.033 | 31.39 | 844 | 0 | 30.813 | MT1-Rear | 11 | 67 | BSE
68 | P-11738 | 2025/05/28 20:11:42 | MT1-Rear-11 | 4 | -6.342 | -27.398 | 19823 | 0 | 1 |  |  | H2O | 1.342 | 0.1 | 0.427 | 102 | -0.017 | -0.052 | 30.77 | 844 | 0 | 30.75 | MT1-Rear | 11 | 67 | BSE
69 | P-11739 | 2025/05/28 20:20:32 | MT1-Rear-12 | 1 | -6.204 | -27.428 | 19658 | -1 | 1 |  |  | H2O | 1.3481 | 0.167 | 0.59 | 96 | -0.029 | 0.125 | 31.75 | 844 | 0 | 30.75 | MT1-Rear | 12 | 67 | BSE
70 | P-11739 | 2025/05/28 20:29:21 | MT1-Rear-12 | 2 | -6.244 | -27.059 | 20242 | -1 | 1 |  |  | H2O | 1.3543 | 0.064 | 0.312 | 95 | -0.001 | -0.073 | 31.27 | 844 | 0 | 30.563 | MT1-Rear | 12 | 67 | BSE
71 | P-11739 | 2025/05/28 20:38:10 | MT1-Rear-12 | 3 | -6.293 | -27.31 | 19610 | -1 | 1 |  |  | H2O | 1.3604 | 0.107 | 0.385 | 109 | 0.01 | 0.069 | 32.72 | 844 | 0 | 30.563 | MT1-Rear | 12 | 67 | BSE
72 | P-11739 | 2025/05/28 20:46:59 | MT1-Rear-12 | 4 | -6.195 | -26.915 | 20119 | 0 | 1 |  |  | H2O | 1.3665 | 0.1 | 0.371 | 98 | 0 | 0.064 | 32.46 | 844 | 0 | 30.563 | MT1-Rear | 12 | 67 | BSE
73 | P-11740 | 2025/05/28 20:55:49 | MT1-Rear-13 | 1 | -6.095 | -25.988 | 19804 | -1 | 1 |  |  | H2O | 1.3726 | 0.126 | 0.278 | 96 | -0.003 | 0.027 | 31.56 | 844 | 0 | 30.688 | MT1-Rear | 13 | 67 | BSE
74 | P-11740 | 2025/05/28 21:04:38 | MT1-Rear-13 | 2 | -6.04 | -26.054 | 20021 | -1 | 1 |  |  | H2O | 1.3787 | 0.125 | 0.393 | 100 | -0.004 | -0.022 | 30.17 | 844 | 0 | 30.5 | MT1-Rear | 13 | 67 | BSE
75 | P-11740 | 2025/05/28 21:13:28 | MT1-Rear-13 | 3 | -6.167 | -26.148 | 19580 | -1 | 1 |  |  | H2O | 1.3849 | 0.099 | 0.418 | 106 | -0.021 | 0.063 | 31.9 | 844 | 0 | 30.375 | MT1-Rear | 13 | 67 | BSE
76 | P-11740 | 2025/05/28 21:22:17 | MT1-Rear-13 | 4 | -6.116 | -25.887 | 19804 | 0 | 1 |  |  | H2O | 1.391 | 0.084 | 0.493 | 106 | -0.017 | 0.098 | 31.98 | 844 | 0 | 30.375 | MT1-Rear | 13 | 67 | BSE
77 | P-11741 | 2025/05/28 21:31:08 | MT1-Rear-14 | 1 | -6.08 | -26.309 | 19666 | -1 | 1 |  |  | H2O | 1.3972 | 0.099 | 0.358 | 102 | -0.023 | -0.063 | 30.82 | 844 | 0 | 30.438 | MT1-Rear | 14 | 67 | BSE
78 | P-11741 | 2025/05/28 21:39:57 | MT1-Rear-14 | 2 | -6.158 | -25.973 | 19472 | -1 | 1 |  |  | H2O | 1.4033 | 0.093 | 0.346 | 94 | 0.015 | 0.016 | 30.94 | 844 | 0 | 30.5 | MT1-Rear | 14 | 67 | BSE
79 | P-11741 | 2025/05/28 21:48:47 | MT1-Rear-14 | 3 | -6.142 | -25.61 | 19751 | -1 | 1 |  |  | H2O | 1.4094 | 0.108 | 0.243 | 94 | -0.014 | 0.014 | 30.87 | 844 | 0 | 30.313 | MT1-Rear | 14 | 67 | BSE
80 | P-11741 | 2025/05/28 21:57:37 | MT1-Rear-14 | 4 | -6.116 | -26.297 | 19606 | 0 | 1 |  |  | H2O | 1.4155 | 0.084 | 0.409 | 100 | -0.016 | -0.027 | 29.98 | 844 | 0 | 30.313 | MT1-Rear | 14 | 67 | BSE
81 | P-11742 | 2025/05/28 22:06:26 | MT1-Rear-15 | 1 | -6.171 | -26.825 | 19687 | -1 | 1 |  |  | H2O | 1.4217 | 0.162 | 0.466 | 105 | -0.015 | 0.109 | 31.68 | 844 | 0 | 30.375 | MT1-Rear | 15 | 67 | BSE
82 | P-11742 | 2025/05/28 22:15:16 | MT1-Rear-15 | 2 | -6.225 | -26.585 | 20167 | -1 | 1 |  |  | H2O | 1.4279 | 0.137 | 0.405 | 103 | -0.015 | -0.097 | 31.11 | 844 | 0 | 30.438 | MT1-Rear | 15 | 67 | BSE
83 | P-11742 | 2025/05/28 22:24:06 | MT1-Rear-15 | 3 | -6.235 | -26.818 | 19789 | -1 | 1 |  |  | H2O | 1.4339 | 0.135 | 0.316 | 96 | 0.001 | 0.055 | 31.75 | 844 | 0 | 30.063 | MT1-Rear | 15 | 67 | BSE
84 | P-11742 | 2025/05/28 22:32:55 | MT1-Rear-15 | 4 | -6.128 | -26.683 | 20013 | 0 | 1 |  |  | H2O | 1.4401 | 0.099 | 0.362 | 94 | -0.019 | -0.01 | 31.16 | 844 | 0 | 29.688 | MT1-Rear | 15 | 67 | BSE
85 | P-11743 | 2025/05/28 22:41:45 | MT1-Rear-16 | 1 | -5.938 | -25.702 | 20183 | -1 | 1 |  |  | H2O | 1.4462 | 0.085 | 0.304 | 103 | -0.012 | -0.044 | 30.99 | 844 | 0 | 29.938 | MT1-Rear | 16 | 67 | BSE
86 | P-11743 | 2025/05/28 22:50:35 | MT1-Rear-16 | 2 | -5.954 | -25.824 | 19683 | -1 | 1 |  |  | H2O | 1.4524 | 0.111 | 0.416 | 101 | -0.018 | -0.015 | 30.29 | 844 | 0 | 30.313 | MT1-Rear | 16 | 67 | BSE
87 | P-11743 | 2025/05/28 22:59:26 | MT1-Rear-16 | 3 | -5.968 | -25.601 | 19622 | -1 | 1 |  |  | H2O | 1.4585 | 0.162 | 0.239 | 100 | -0.036 | -0.009 | 30.15 | 844 | 0 | 30.5 | MT1-Rear | 16 | 67 | BSE
88 | P-11743 | 2025/05/28 23:08:15 | MT1-Rear-16 | 4 | -6.009 | -25.596 | 19832 | 0 | 1 |  |  | H2O | 1.4646 | 0.143 | 0.336 | 92 | -0.035 | -0.063 | 30.47 | 844 | 0 | 30.75 | MT1-Rear | 16 | 67 | BSE
89 | P-11744 | 2025/05/28 23:17:05 | MT1-Rear-17 | 1 | -6.043 | -25.534 | 19796 | -1 | 1 |  |  | H2O | 1.4708 | 0.085 | 0.408 | 100 | -0.012 | 0.103 | 33.04 | 844 | 0 | 30.938 | MT1-Rear | 17 | 67 | BSE
90 | P-11744 | 2025/05/28 23:25:55 | MT1-Rear-17 | 2 | -5.985 | -25.518 | 19768 | -1 | 1 |  |  | H2O | 1.4769 | 0.109 | 0.294 | 102 | -0.022 | -0.046 | 30.75 | 844 | 0 | 31.063 | MT1-Rear | 17 | 67 | BSE
91 | P-11744 | 2025/05/28 23:34:46 | MT1-Rear-17 | 3 | -6.047 | -25.52 | 19730 | -1 | 1 |  |  | H2O | 1.483 | 0.129 | 0.48 | 104 | -0.023 | 0.063 | 31.38 | 844 | 0 | 31.125 | MT1-Rear | 17 | 67 | BSE
92 | P-11744 | 2025/05/28 23:43:36 | MT1-Rear-17 | 4 | -5.989 | -25.618 | 19568 | 0 | 1 |  |  | H2O | 1.4892 | 0.126 | 0.249 | 103 | -0.011 | -0.03 | 31 | 844 | 0 | 30.375 | MT1-Rear | 17 | 67 | BSE
93 | P-11745 | 2025/05/28 23:52:25 | MT1-Rear-18 | 1 | -8.502 | -41.551 | 19717 | -1 | 1 |  |  | H2O | 1.4953 | 0.088 | 0.506 | 98 | 0.009 | 0.144 | 32.42 | 844 | 0 | 30.375 | MT1-Rear | 18 | 67 | BSE
94 | P-11745 | 2025/05/29 00:01:16 | MT1-Rear-18 | 2 | -8.614 | -43.09 | 19515 | -1 | 1 |  |  | H2O | 1.5014 | 0.093 | 0.405 | 102 | -0.017 | 0.076 | 30.76 | 844 | 0 | 30.563 | MT1-Rear | 18 | 67 | BSE
95 | P-11745 | 2025/05/29 00:10:05 | MT1-Rear-18 | 3 | -8.618 | -43.706 | 19449 | -1 | 1 |  |  | H2O | 1.5076 | 0.123 | 0.389 | 104 | -0.014 | 0.049 | 31.36 | 844 | 0 | 30.688 | MT1-Rear | 18 | 67 | BSE
96 | P-11745 | 2025/05/29 00:18:56 | MT1-Rear-18 | 4 | -8.652 | -43.71 | 20325 | 0 | 1 |  |  | H2O | 1.5137 | 0.134 | 0.428 | 109 | 0.027 | 0.099 | 32.92 | 844 | 0 | 30.688 | MT1-Rear | 18 | 67 | BSE
97 | P-11746 | 2025/05/29 00:27:47 | MT1-Rear-19 | 1 | -7.862 | -35.696 | 19661 | -1 | 1 |  |  | H2O | 1.5198 | 0.143 | 0.206 | 93 | -0.034 | -0.018 | 30.56 | 844 | 0 | 30.625 | MT1-Rear | 19 | 67 | BSE
98 | P-11746 | 2025/05/29 00:36:37 | MT1-Rear-19 | 2 | -7.868 | -34.811 | 19606 | -1 | 1 |  |  | H2O | 1.526 | 0.134 | 0.55 | 94 | 0.001 | 0.157 | 31.12 | 844 | 0 | 30.438 | MT1-Rear | 19 | 67 | BSE
99 | P-11746 | 2025/05/29 00:45:28 | MT1-Rear-19 | 3 | -7.875 | -34.657 | 19694 | -1 | 1 |  |  | H2O | 1.5321 | 0.088 | 0.388 | 96 | 0 | 0.085 | 31.77 | 844 | 0 | 30.5 | MT1-Rear | 19 | 67 | BSE
100 | P-11746 | 2025/05/29 00:54:18 | MT1-Rear-19 | 4 | -7.845 | -34.418 | 19685 | 0 | 1 |  |  | H2O | 1.5382 | 0.143 | 0.242 | 102 | -0.032 | 0.02 | 30.57 | 844 | 0 | 30.5 | MT1-Rear | 19 | 67 | BSE
101 | P-11747 | 2025/05/29 01:03:09 | MT1-Rear-20 | 1 | -3.507 | -5.04 | 19813 | -1 | 1 |  |  | H2O | 1.5444 | 0.155 | 0.251 | 103 | -0.037 | 0.015 | 31.02 | 844 | 0 | 30.25 | MT1-Rear | 20 | 67 | BSE
102 | P-11747 | 2025/05/29 01:12:00 | MT1-Rear-20 | 2 | -3.21 | -2.149 | 19725 | -1 | 1 |  |  | H2O | 1.5506 | 0.141 | 0.307 | 104 | -0.021 | -0.05 | 31.25 | 844 | 0 | 30 | MT1-Rear | 20 | 67 | BSE
103 | P-11747 | 2025/05/29 01:20:50 | MT1-Rear-20 | 3 | -3.123 | -1.468 | 19471 | -1 | 1 |  |  | H2O | 1.5567 | 0.112 | 0.193 | 105 | -0.024 | 0.022 | 31.56 | 844 | 0 | 29.563 | MT1-Rear | 20 | 67 | BSE
104 | P-11747 | 2025/05/29 01:29:41 | MT1-Rear-20 | 4 | -3.085 | -1.26 | 19514 | 0 | 1 |  |  | H2O | 1.5628 | 0.091 | 0.387 | 98 | -0.002 | 0.074 | 32.19 | 844 | 0 | 30 | MT1-Rear | 20 | 67 | BSE
105 | P-11748 | 2025/05/29 01:38:32 | MT1-Rear-21 | 1 | -5.875 | -22.671 | 19978 | -1 | 1 |  |  | H2O | 1.569 | 0.063 | 0.397 | 95 | -0.001 | 0.096 | 31.51 | 844 | 0 | 30.438 | MT1-Rear | 21 | 67 | BSE
106 | P-11748 | 2025/05/29 01:47:22 | MT1-Rear-21 | 2 | -6.012 | -25.111 | 20209 | -1 | 1 |  |  | H2O | 1.5751 | 0.127 | 0.324 | 102 | -0.014 | 0.041 | 30.69 | 844 | 0 | 30.5 | MT1-Rear | 21 | 67 | BSE
107 | P-11748 | 2025/05/29 01:56:12 | MT1-Rear-21 | 3 | -6.06 | -25.624 | 20117 | -1 | 1 |  |  | H2O | 1.5813 | 0.085 | 0.341 | 101 | -0.014 | -0.056 | 30.37 | 844 | 0 | 30.875 | MT1-Rear | 21 | 67 | BSE
108 | P-11748 | 2025/05/29 02:05:04 | MT1-Rear-21 | 4 | -6.098 | -25.716 | 19771 | 0 | 1 |  |  | H2O | 1.5874 | 0.108 | 0.354 | 101 | -0.011 | -0.003 | 30.51 | 844 | 0 | 31 | MT1-Rear | 21 | 67 | BSE
109 | P-11749 | 2025/05/29 02:13:56 | MT1-Rear-22 | 1 | -4.874 | -20.285 | 19627 | -1 | 1 |  |  | H2O | 1.5936 | 0.133 | 0.273 | 101 | -0.034 | -0.034 | 30.33 | 844 | 0 | 30.938 | MT1-Rear | 22 | 67 | BSE
110 | P-11749 | 2025/05/29 02:22:47 | MT1-Rear-22 | 2 | -4.738 | -19.591 | 19755 | -1 | 1 |  |  | H2O | 1.5998 | 0.121 | 0.373 | 102 | -0.006 | 0.022 | 30.64 | 844 | 0 | 30.563 | MT1-Rear | 22 | 67 | BSE
111 | P-11749 | 2025/05/29 02:31:39 | MT1-Rear-22 | 3 | -4.672 | -19.608 | 19832 | -1 | 1 |  |  | H2O | 1.6058 | 0.086 | 0.264 | 97 | -0.018 | 0.024 | 31.95 | 844 | 0 | 30.438 | MT1-Rear | 22 | 67 | BSE
112 | P-11749 | 2025/05/29 02:40:30 | MT1-Rear-22 | 4 | -4.573 | -19.822 | 19540 | 0 | 1 |  |  | H2O | 1.612 | 0.143 | 0.305 | 93 | -0.028 | 0.028 | 30.67 | 844 | 0 | 30.438 | MT1-Rear | 22 | 67 | BSE
113 | P-11750 | 2025/05/29 02:49:20 | MT1-Rear-23 | 1 | -5.511 | -23.047 | 19793 | -1 | 1 |  |  | H2O | 1.6182 | 0.14 | 0.402 | 94 | -0.004 | 0.054 | 30.94 | 844 | 0 | 30.563 | MT1-Rear | 23 | 67 | BSE
114 | P-11750 | 2025/05/29 02:58:11 | MT1-Rear-23 | 2 | -5.507 | -23.407 | 19980 | -1 | 1 |  |  | H2O | 1.6243 | 0.108 | 0.349 | 93 | -0.021 | -0.023 | 30.78 | 844 | 0 | 30.5 | MT1-Rear | 23 | 67 | BSE
115 | P-11750 | 2025/05/29 03:07:02 | MT1-Rear-23 | 3 | -5.496 | -23.668 | 19956 | -1 | 1 |  |  | H2O | 1.6304 | 0.134 | 0.377 | 103 | -0.025 | -0.08 | 31.14 | 844 | 0 | 30.688 | MT1-Rear | 23 | 67 | BSE
116 | P-11750 | 2025/05/29 03:15:54 | MT1-Rear-23 | 4 | -5.532 | -23.725 | 20070 | 0 | 1 |  |  | H2O | 1.6366 | 0.121 | 0.336 | 103 | -0.016 | -0.04 | 31.04 | 844 | 0 | 30.688 | MT1-Rear | 23 | 67 | BSE
117 | P-11751 | 2025/05/29 03:24:46 | MT1-Rear-24 | 1 | -5.827 | -25.546 | 20191 | -1 | 1 |  |  | H2O | 1.6428 | 0.111 | 0.393 | 108 | -0.004 | 0.08 | 32.48 | 844 | 0 | 30.563 | MT1-Rear | 24 | 67 | BSE
118 | P-11751 | 2025/05/29 03:33:38 | MT1-Rear-24 | 2 | -5.897 | -25.68 | 19954 | -1 | 1 |  |  | H2O | 1.6489 | 0.141 | 0.643 | 108 | 0.035 | 0.148 | 32.52 | 844 | 0 | 30.563 | MT1-Rear | 24 | 67 | BSE
119 | P-11751 | 2025/05/29 03:42:30 | MT1-Rear-24 | 3 | -5.83 | -25.249 | 20105 | -1 | 1 |  |  | H2O | 1.6551 | 0.077 | 0.314 | 106 | -0.019 | 0.058 | 31.87 | 844 | 0 | 30.625 | MT1-Rear | 24 | 67 | BSE
120 | P-11751 | 2025/05/29 03:51:22 | MT1-Rear-24 | 4 | -5.896 | -25.378 | 19725 | 0 | 1 |  |  | H2O | 1.6613 | 0.081 | 0.304 | 102 | -0.012 | -0.018 | 30.66 | 844 | 0 | 29.875 | MT1-Rear | 24 | 67 | BSE
121 | P-11752 | 2025/05/29 04:00:14 | MT1-Rear-25 | 1 | -5.861 | -25.251 | 20011 | -1 | 1 |  |  | H2O | 1.6674 | 0.076 | 0.424 | 105 | 0.008 | 0.045 | 31.53 | 844 | 0 | 29.5 | MT1-Rear | 25 | 67 | BSE
122 | P-11752 | 2025/05/29 04:09:05 | MT1-Rear-25 | 2 | -5.773 | -25.645 | 19928 | -1 | 1 |  |  | H2O | 1.6736 | 0.111 | 0.324 | 103 | -0.021 | -0.005 | 31.12 | 844 | 0 | 29.688 | MT1-Rear | 25 | 67 | BSE
123 | P-11752 | 2025/05/29 04:17:57 | MT1-Rear-25 | 3 | -5.854 | -25.54 | 19764 | -1 | 1 |  |  | H2O | 1.6797 | 0.131 | 0.346 | 104 | -0.026 | 0.064 | 31.44 | 844 | 0 | 30.313 | MT1-Rear | 25 | 67 | BSE
124 | P-11752 | 2025/05/29 04:26:48 | MT1-Rear-25 | 4 | -5.867 | -25.078 | 19626 | 0 | 1 |  |  | H2O | 1.6859 | 0.132 | 0.246 | 106 | -0.01 | 0.021 | 31.91 | 844 | 0 | 30.813 | MT1-Rear | 25 | 67 | BSE
125 | P-11753 | 2025/05/29 04:35:39 | MT1-Rear-26 | 1 | -2.935 | -5.593 | 20222 | -1 | 1 |  |  | H2O | 1.692 | 0.13 | 0.304 | 92 | -0.038 | -0.074 | 30.25 | 844 | 0 | 30.813 | MT1-Rear | 26 | 67 | BSE
126 | P-11753 | 2025/05/29 04:44:32 | MT1-Rear-26 | 2 | -2.729 | -3.328 | 19778 | -1 | 1 |  |  | H2O | 1.6982 | 0.109 | 0.347 | 104 | -0.019 | -0.031 | 31.2 | 844 | 0 | 30.938 | MT1-Rear | 26 | 67 | BSE
127 | P-11753 | 2025/05/29 04:53:24 | MT1-Rear-26 | 3 | -2.706 | -3.356 | 19612 | -1 | 1 |  |  | H2O | 1.7043 | 0.123 | 0.434 | 93 | 0.023 | 0.111 | 30.58 | 844 | 0 | 31 | MT1-Rear | 26 | 67 | BSE
128 | P-11753 | 2025/05/29 05:02:16 | MT1-Rear-26 | 4 | -2.732 | -2.39 | 20198 | 0 | 1 |  |  | H2O | 1.7105 | 0.153 | 0.347 | 92 | -0.045 | -0.021 | 30.49 | 844 | 0 | 31 | MT1-Rear | 26 | 67 | BSE
129 | P-11754 | 2025/05/29 05:11:08 | MT1-Rear-27 | 1 | -8.166 | -38.3 | 19696 | -1 | 1 |  |  | H2O | 1.7167 | 0.136 | 0.407 | 91 | -0.031 | 0.105 | 30.15 | 844 | 0 | 30.875 | MT1-Rear | 27 | 67 | BSE
130 | P-11754 | 2025/05/29 05:20:00 | MT1-Rear-27 | 2 | -8.554 | -42.089 | 19594 | -1 | 1 |  |  | H2O | 1.7228 | 0.152 | 0.35 | 95 | 0.002 | 0.079 | 31.27 | 844 | 0 | 30.813 | MT1-Rear | 27 | 67 | BSE
131 | P-11754 | 2025/05/29 05:28:52 | MT1-Rear-27 | 3 | -8.552 | -42.67 | 19527 | -1 | 1 |  |  | H2O | 1.7289 | 0.094 | 0.245 | 90 | -0.018 | 0.003 | 29.64 | 844 | 0 | 30.625 | MT1-Rear | 27 | 67 | BSE
132 | P-11754 | 2025/05/29 05:37:43 | MT1-Rear-27 | 4 | -8.68 | -43.174 | 19763 | 0 | 1 |  |  | H2O | 1.7351 | 0.143 | 0.401 | 93 | -0.028 | 0.057 | 30.65 | 844 | 0 | 30.563 | MT1-Rear | 27 | 67 | BSE

## Run Summary
Line | Sample | Inj Nr | Sample Name | H2O_Mean | Good | Ignore | Time Code | d(18_16)Mean | d(D_H)Mean
1 | 1 | 1 | Blacksburg | 24311 | 0 | -1 | 2025/05/28 10:22:30 | -8.596 | -46.973 |  | Min | Max
2 | 1 | 2 | Blacksburg | 19602 | 1 | 0 | 2025/05/28 10:31:16 | -9.022 | -45.287 | 0 | 17000 | 23000
3 | 1 | 3 | Blacksburg | 19434 | 1 | 0 | 2025/05/28 10:40:03 | -8.849 | -44.776 | 500 | 17000 | 23000
4 | 1 | 4 | Blacksburg | 19255 | 1 | 0 | 2025/05/28 10:48:49 | -8.826 | -44.877
5 | 1 | 5 | Blacksburg | 19549 | 1 | 0 | 2025/05/28 10:57:36 | -8.767 | -44.464
6 | 1 | 6 | Blacksburg | 19108 | 1 | 0 | 2025/05/28 11:06:22 | -8.797 | -44.194
7 | 1 | 7 | Blacksburg | 19370 | 1 | 0 | 2025/05/28 11:15:09 | -8.857 | -43.957
8 | 1 | 8 | Blacksburg | 19353 | 1 | 0 | 2025/05/28 11:23:55 | -8.848 | -44.191
9 | 1 | 9 | Blacksburg | 19494 | 1 | 0 | 2025/05/28 11:32:42 | -8.762 | -43.925
10 | 1 | 10 | Blacksburg | 19311 | 1 | 0 | 2025/05/28 11:41:29 | -8.889 | -43.99
11 | 2 | 1 | Myrtle | 19440 | 1 | 0 | 2025/05/28 11:50:16 | -3.8 | -9.933
12 | 2 | 2 | Myrtle | 19531 | 1 | 0 | 2025/05/28 11:59:02 | -3.403 | -6.116
13 | 2 | 3 | Myrtle | 21298 | 1 | 0 | 2025/05/28 12:07:49 | -3.42 | -5.319
14 | 2 | 4 | Myrtle | 19506 | 1 | 0 | 2025/05/28 12:16:35 | -3.315 | -4.857
15 | 2 | 5 | Myrtle | 19560 | 1 | 0 | 2025/05/28 12:25:22 | -3.299 | -4.853
16 | 2 | 6 | Myrtle | 19611 | 1 | 0 | 2025/05/28 12:34:09 | -3.228 | -4.782
17 | 2 | 7 | Myrtle | 19464 | 1 | 0 | 2025/05/28 12:42:55 | -3.338 | -4.586 | Concentration of water vapor should be about 20,000 ppmV (roughly +/- 1000; but not outside 17,000 and 23,000 ppmV).
18 | 2 | 8 | Myrtle | 19840 | 1 | 0 | 2025/05/28 12:51:42 | -3.336 | -4.557 | Large scatter may indicate a problem with the syringe or the septa.
19 | 2 | 9 | Myrtle | 19606 | 1 | 0 | 2025/05/28 13:00:29 | -3.36 | -4.473
20 | 2 | 10 | Myrtle | 19782 | 1 | 0 | 2025/05/28 13:09:16 | -3.251 | -4.442
21 | 3 | 1 | Homer | 19804 | 1 | 0 | 2025/05/28 13:18:03 | -14.813 | -93.78
22 | 3 | 2 | Homer | 19503 | 1 | 0 | 2025/05/28 13:26:49 | -15.687 | -104.711
23 | 3 | 3 | Homer | 19616 | 1 | 0 | 2025/05/28 13:35:36 | -15.812 | -106.635
24 | 3 | 4 | Homer | 19830 | 1 | 0 | 2025/05/28 13:44:23 | -15.915 | -107.535
25 | 3 | 5 | Homer | 19529 | 1 | 0 | 2025/05/28 13:53:11 | -15.9 | -108.061
26 | 3 | 6 | Homer | 19691 | 1 | 0 | 2025/05/28 14:01:58 | -15.887 | -108.108
27 | 3 | 7 | Homer | 21834 | 1 | 0 | 2025/05/28 14:10:45 | -15.656 | -106.868
28 | 3 | 8 | Homer | 20743 | 1 | 0 | 2025/05/28 14:19:33 | -15.943 | -108.093
29 | 3 | 9 | Homer | 19279 | 1 | 0 | 2025/05/28 14:28:21 | -16.09 | -108.993
30 | 3 | 10 | Homer | 19558 | 1 | 0 | 2025/05/28 14:37:07 | -15.946 | -108.845
31 | 4 | 1 | Blacksburg | 20179 | 1 | 0 | 2025/05/28 14:45:55 | -9.359 | -52.697
32 | 4 | 2 | Blacksburg | 19418 | 1 | 0 | 2025/05/28 14:54:43 | -9.056 | -46.746
33 | 4 | 3 | Blacksburg | 19651 | 1 | 0 | 2025/05/28 15:03:29 | -8.884 | -44.807
34 | 4 | 4 | Blacksburg | 19462 | 1 | 0 | 2025/05/28 15:12:16 | -8.87 | -44.788
35 | 4 | 5 | Blacksburg | 19795 | 1 | 0 | 2025/05/28 15:21:04 | -8.884 | -44.922
36 | 4 | 6 | Blacksburg | 19686 | 1 | 0 | 2025/05/28 15:29:51 | -8.742 | -44.182
37 | 4 | 7 | Blacksburg | 19361 | 1 | 0 | 2025/05/28 15:38:39 | -8.849 | -44.32
38 | 4 | 8 | Blacksburg | 19337 | 1 | 0 | 2025/05/28 15:47:27 | -8.824 | -44.19
39 | 4 | 9 | Blacksburg | 19655 | 1 | 0 | 2025/05/28 15:56:15 | -8.865 | -44.306
40 | 4 | 10 | Blacksburg | 19411 | 1 | 0 | 2025/05/28 16:05:02 | -8.855 | -44.238
41 | 5 | 1 | Hawaii | 19802 | 1 | 0 | 2025/05/28 16:13:50 | -4.265 | -10.89
42 | 5 | 2 | Hawaii | 19686 | 1 | 0 | 2025/05/28 16:22:39 | -3.984 | -6.882
43 | 5 | 3 | Hawaii | 19227 | 1 | 0 | 2025/05/28 16:31:27 | -3.982 | -6.409
44 | 5 | 4 | Hawaii | 19644 | 1 | 0 | 2025/05/28 16:40:15 | -3.929 | -5.993
45 | 6 | 1 | CP1 22may24 BOT | 19774 | 1 | 0 | 2025/05/28 16:49:03 | -6.922 | -27.362
46 | 6 | 2 | CP1 22may24 BOT | 19344 | 1 | 0 | 2025/05/28 16:57:51 | -7.218 | -29.471
47 | 6 | 3 | CP1 22may24 BOT | 19270 | 1 | 0 | 2025/05/28 17:06:38 | -7.142 | -30.48
48 | 6 | 4 | CP1 22may24 BOT | 19505 | 1 | 0 | 2025/05/28 17:15:27 | -7.171 | -30.276
49 | 7 | 1 | C50 16sep24 0.1m | 19347 | 1 | 0 | 2025/05/28 17:24:14 | -4.995 | -21.571
50 | 7 | 2 | C50 16sep24 0.1m | 19785 | 1 | 0 | 2025/05/28 17:33:02 | -4.827 | -20.29
51 | 7 | 3 | C50 16sep24 0.1m | 19774 | 1 | 0 | 2025/05/28 17:41:51 | -4.781 | -20.114
52 | 7 | 4 | C50 16sep24 0.1m | 19799 | 1 | 0 | 2025/05/28 17:50:39 | -4.785 | -20.188
53 | 8 | 1 | CC2 19jun24 0.1m | 19759 | 1 | 0 | 2025/05/28 17:59:28 | -5.712 | -23.364
54 | 8 | 2 | CC2 19jun24 0.1m | 19757 | 1 | 0 | 2025/05/28 18:08:16 | -5.742 | -23.808
55 | 8 | 3 | CC2 19jun24 0.1m | 19837 | 1 | 0 | 2025/05/28 18:17:04 | -5.734 | -23.751
56 | 8 | 4 | CC2 19jun24 0.1m | 19517 | 1 | 0 | 2025/05/28 18:25:53 | -5.813 | -24.158
57 | 9 | 1 | Blacksburg | 19810 | 1 | 0 | 2025/05/28 18:34:41 | -8.516 | -40.729
58 | 9 | 2 | Blacksburg | 19849 | 1 | 0 | 2025/05/28 18:43:30 | -8.655 | -43.047
59 | 9 | 3 | Blacksburg | 19653 | 1 | 0 | 2025/05/28 18:52:19 | -8.712 | -43.539
60 | 9 | 4 | Blacksburg | 19694 | 1 | 0 | 2025/05/28 19:01:08 | -8.782 | -43.612
61 | 10 | 1 | CP1 28oct24 0.1m | 19448 | 1 | 0 | 2025/05/28 19:09:58 | -7.303 | -31.612
62 | 10 | 2 | CP1 28oct24 0.1m | 19675 | 1 | 0 | 2025/05/28 19:18:47 | -7.235 | -30.47
63 | 10 | 3 | CP1 28oct24 0.1m | 19631 | 1 | 0 | 2025/05/28 19:27:37 | -7.231 | -30.183
64 | 10 | 4 | CP1 28oct24 0.1m | 20147 | 1 | 0 | 2025/05/28 19:36:26 | -7.123 | -29.866
65 | 11 | 1 | CC4 11jul24 9m | 19728 | 1 | 0 | 2025/05/28 19:45:16 | -6.451 | -27.991
66 | 11 | 2 | CC4 11jul24 9m | 19771 | 1 | 0 | 2025/05/28 19:54:04 | -6.405 | -27.818
67 | 11 | 3 | CC4 11jul24 9m | 19926 | 1 | 0 | 2025/05/28 20:02:53 | -6.435 | -27.54
68 | 11 | 4 | CC4 11jul24 9m | 19823 | 1 | 0 | 2025/05/28 20:11:42 | -6.342 | -27.398
69 | 12 | 1 | CC4 15aug24 BOT | 19658 | 1 | 0 | 2025/05/28 20:20:32 | -6.204 | -27.428
70 | 12 | 2 | CC4 15aug24 BOT | 20242 | 1 | 0 | 2025/05/28 20:29:21 | -6.244 | -27.059
71 | 12 | 3 | CC4 15aug24 BOT | 19610 | 1 | 0 | 2025/05/28 20:38:10 | -6.293 | -27.31
72 | 12 | 4 | CC4 15aug24 BOT | 20119 | 1 | 0 | 2025/05/28 20:46:59 | -6.195 | -26.915
73 | 13 | 1 | C50 16apr25 0.1m | 19804 | 1 | 0 | 2025/05/28 20:55:49 | -6.095 | -25.988
74 | 13 | 2 | C50 16apr25 0.1m | 20021 | 1 | 0 | 2025/05/28 21:04:38 | -6.04 | -26.054
75 | 13 | 3 | C50 16apr25 0.1m | 19580 | 1 | 0 | 2025/05/28 21:13:28 | -6.167 | -26.148
76 | 13 | 4 | C50 16apr25 0.1m | 19804 | 1 | 0 | 2025/05/28 21:22:17 | -6.116 | -25.887
77 | 14 | 1 | C50 17feb25 1.5m | 19666 | 1 | 0 | 2025/05/28 21:31:08 | -6.08 | -26.309
78 | 14 | 2 | C50 17feb25 1.5m | 19472 | 1 | 0 | 2025/05/28 21:39:57 | -6.158 | -25.973
79 | 14 | 3 | C50 17feb25 1.5m | 19751 | 1 | 0 | 2025/05/28 21:48:47 | -6.142 | -25.61
80 | 14 | 4 | C50 17feb25 1.5m | 19606 | 1 | 0 | 2025/05/28 21:57:37 | -6.116 | -26.297
81 | 15 | 1 | C50 11jul24 9m | 19687 | 1 | 0 | 2025/05/28 22:06:26 | -6.171 | -26.825
82 | 15 | 2 | C50 11jul24 9m | 20167 | 1 | 0 | 2025/05/28 22:15:16 | -6.225 | -26.585
83 | 15 | 3 | C50 11jul24 9m | 19789 | 1 | 0 | 2025/05/28 22:24:06 | -6.235 | -26.818
84 | 15 | 4 | C50 11jul24 9m | 20013 | 1 | 0 | 2025/05/28 22:32:55 | -6.128 | -26.683
85 | 16 | 1 | CC4 22may24 0.1m | 20183 | 1 | 0 | 2025/05/28 22:41:45 | -5.938 | -25.702
86 | 16 | 2 | CC4 22may24 0.1m | 19683 | 1 | 0 | 2025/05/28 22:50:35 | -5.954 | -25.824
87 | 16 | 3 | CC4 22may24 0.1m | 19622 | 1 | 0 | 2025/05/28 22:59:26 | -5.968 | -25.601
88 | 16 | 4 | CC4 22may24 0.1m | 19832 | 1 | 0 | 2025/05/28 23:08:15 | -6.009 | -25.596
89 | 17 | 1 | C50 22may24 0.1m | 19796 | 1 | 0 | 2025/05/28 23:17:05 | -6.043 | -25.534
90 | 17 | 2 | C50 22may24 0.1m | 19768 | 1 | 0 | 2025/05/28 23:25:55 | -5.985 | -25.518
91 | 17 | 3 | C50 22may24 0.1m | 19730 | 1 | 0 | 2025/05/28 23:34:46 | -6.047 | -25.52
92 | 17 | 4 | C50 22may24 0.1m | 19568 | 1 | 0 | 2025/05/28 23:43:36 | -5.989 | -25.618
93 | 18 | 1 | Blacksburg | 19717 | 1 | 0 | 2025/05/28 23:52:25 | -8.502 | -41.551
94 | 18 | 2 | Blacksburg | 19515 | 1 | 0 | 2025/05/29 00:01:16 | -8.614 | -43.09
95 | 18 | 3 | Blacksburg | 19449 | 1 | 0 | 2025/05/29 00:10:05 | -8.618 | -43.706
96 | 18 | 4 | Blacksburg | 20325 | 1 | 0 | 2025/05/29 00:18:56 | -8.652 | -43.71
97 | 19 | 1 | CS2 22may24 0.1m | 19661 | 1 | 0 | 2025/05/29 00:27:47 | -7.862 | -35.696
98 | 19 | 2 | CS2 22may24 0.1m | 19606 | 1 | 0 | 2025/05/29 00:36:37 | -7.868 | -34.811
99 | 19 | 3 | CS2 22may24 0.1m | 19694 | 1 | 0 | 2025/05/29 00:45:28 | -7.875 | -34.657
100 | 19 | 4 | CS2 22may24 0.1m | 19685 | 1 | 0 | 2025/05/29 00:54:18 | -7.845 | -34.418
101 | 20 | 1 | CCR raindeploy - 11jul24 R1 | 19813 | 1 | 0 | 2025/05/29 01:03:09 | -3.507 | -5.04
102 | 20 | 2 | CCR raindeploy - 11jul24 R1 | 19725 | 1 | 0 | 2025/05/29 01:12:00 | -3.21 | -2.149
103 | 20 | 3 | CCR raindeploy - 11jul24 R1 | 19471 | 1 | 0 | 2025/05/29 01:20:50 | -3.123 | -1.468
104 | 20 | 4 | CCR raindeploy - 11jul24 R1 | 19514 | 1 | 0 | 2025/05/29 01:29:41 | -3.085 | -1.26
105 | 21 | 1 | C50 16apr25 6m | 19978 | 1 | 0 | 2025/05/29 01:38:32 | -5.875 | -22.671
106 | 21 | 2 | C50 16apr25 6m | 20209 | 1 | 0 | 2025/05/29 01:47:22 | -6.012 | -25.111
107 | 21 | 3 | C50 16apr25 6m | 20117 | 1 | 0 | 2025/05/29 01:56:12 | -6.06 | -25.624
108 | 21 | 4 | C50 16apr25 6m | 19771 | 1 | 0 | 2025/05/29 02:05:04 | -6.098 | -25.716
109 | 22 | 1 | CC4 16sep24 1.5m | 19627 | 1 | 0 | 2025/05/29 02:13:56 | -4.874 | -20.285
110 | 22 | 2 | CC4 16sep24 1.5m | 19755 | 1 | 0 | 2025/05/29 02:22:47 | -4.738 | -19.591
111 | 22 | 3 | CC4 16sep24 1.5m | 19832 | 1 | 0 | 2025/05/29 02:31:39 | -4.672 | -19.608
112 | 22 | 4 | CC4 16sep24 1.5m | 19540 | 1 | 0 | 2025/05/29 02:40:30 | -4.573 | -19.822
113 | 23 | 1 | C50 19jun24 1.5m | 19793 | 1 | 0 | 2025/05/29 02:49:20 | -5.511 | -23.047
114 | 23 | 2 | C50 19jun24 1.5m | 19980 | 1 | 0 | 2025/05/29 02:58:11 | -5.507 | -23.407
115 | 23 | 3 | C50 19jun24 1.5m | 19956 | 1 | 0 | 2025/05/29 03:07:02 | -5.496 | -23.668
116 | 23 | 4 | C50 19jun24 1.5m | 20070 | 1 | 0 | 2025/05/29 03:15:54 | -5.532 | -23.725
117 | 24 | 1 | CP2 22may24 0.1m | 20191 | 1 | 0 | 2025/05/29 03:24:46 | -5.827 | -25.546
118 | 24 | 2 | CP2 22may24 0.1m | 19954 | 1 | 0 | 2025/05/29 03:33:38 | -5.897 | -25.68
119 | 24 | 3 | CP2 22may24 0.1m | 20105 | 1 | 0 | 2025/05/29 03:42:30 | -5.83 | -25.249
120 | 24 | 4 | CP2 22may24 0.1m | 19725 | 1 | 0 | 2025/05/29 03:51:22 | -5.896 | -25.378
121 | 25 | 1 | CC4 22may24 0.1m - DUP | 20011 | 1 | 0 | 2025/05/29 04:00:14 | -5.861 | -25.251
122 | 25 | 2 | CC4 22may24 0.1m - DUP | 19928 | 1 | 0 | 2025/05/29 04:09:05 | -5.773 | -25.645
123 | 25 | 3 | CC4 22may24 0.1m - DUP | 19764 | 1 | 0 | 2025/05/29 04:17:57 | -5.854 | -25.54
124 | 25 | 4 | CC4 22may24 0.1m - DUP | 19626 | 1 | 0 | 2025/05/29 04:26:48 | -5.867 | -25.078
125 | 26 | 1 | Hawaii | 20222 | 1 | 0 | 2025/05/29 04:35:39 | -2.935 | -5.593
126 | 26 | 2 | Hawaii | 19778 | 1 | 0 | 2025/05/29 04:44:32 | -2.729 | -3.328
127 | 26 | 3 | Hawaii | 19612 | 1 | 0 | 2025/05/29 04:53:24 | -2.706 | -3.356
128 | 26 | 4 | Hawaii | 20198 | 1 | 0 | 2025/05/29 05:02:16 | -2.732 | -2.39
129 | 27 | 1 | Blacksburg | 19696 | 1 | 0 | 2025/05/29 05:11:08 | -8.166 | -38.3
130 | 27 | 2 | Blacksburg | 19594 | 1 | 0 | 2025/05/29 05:20:00 | -8.554 | -42.089
131 | 27 | 3 | Blacksburg | 19527 | 1 | 0 | 2025/05/29 05:28:52 | -8.552 | -42.67
132 | 27 | 4 | Blacksburg | 19763 | 1 | 0 | 2025/05/29 05:37:43 | -8.68 | -43.174

## 18O
Line | Sample | Injection | Sample Name | Correction Type | Good | Ignore | d(18_16) Picarro Raw | Stdev Raw | d(18_16) Memory Corrected | Stdev Memory Corrected | d(18_16) Memory Corrected Drifts | d(18_16) Drift Corrected | Drift Corrected Drifts | d(18_16) Drift Corrected High, Low, and Drift | d(18_16) Reference Defined | d18O Final Corrected Value | Post run corrections:
1 | 1 | 1 | Blacksburg | Drift | 0 | -1 | -8.596 | [9 empty] | (1) Memory correction |  |  | (1) Graphs for memory correction
2 | 1 | 2 | Blacksburg | Drift | 1 | -1 | -9.022 | [9 empty] | Combined Stdev | 0.11832
3 | 1 | 3 | Blacksburg | Drift | 1 | -1 | -8.849 | [9 empty] | Date:
4 | 1 | 4 | Blacksburg | Drift | 1 | -1 | -8.826 | [9 empty] | injection | 18O memory coef
5 | 1 | 5 | Blacksburg | Drift | 1 | -1 | -8.767 | [9 empty] | 1 | 0.90296
6 | 1 | 6 | Blacksburg | Drift | 1 | -1 | -8.797 | [9 empty] | 2 | 0.97438
7 | 1 | 7 | Blacksburg | Drift | 1 | 0 | -8.857 |  | -8.857 |  | -8.857 | -8.8663 | -8.8663 |  |  | -7.6291 | 3 | 0.98597
8 | 1 | 8 | Blacksburg | Drift | 1 | 0 | -8.848 |  | -8.848 |  | -8.848 | -8.8586 | -8.8586 |  |  | -7.6217 | 4 | 0.99185
9 | 1 | 9 | Blacksburg | Drift | 1 | 0 | -8.762 |  | -8.762 |  | -8.762 | -8.7739 | -8.7739 |  |  | -7.54 | 5 | 0.99185
10 | 1 | 10 | Blacksburg | Drift | 1 | 0 | -8.889 | 0.054265 | -8.889 | 0.054265 | -8.889 | -8.9022 | -8.9022 |  |  | -7.6638 | 6 | 0.99185
11 | 2 | 1 | Myrtle | High Reference | 1 | 0 | -3.8 |  | -3.3061 |  |  | -3.3207 |  | -3.3207 | -2.2908 | -2.2791 | 7 | 0.99185
12 | 2 | 2 | Myrtle | High Reference | 1 | 0 | -3.403 |  | -3.2625 |  |  | -3.2783 |  | -3.2783 | -2.2908 | -2.2382 | 8 | 0.99664
13 | 2 | 3 | Myrtle | High Reference | 1 | 0 | -3.42 |  | -3.3432 |  |  | -3.3604 |  | -3.3604 | -2.2908 | -2.3174 | 9 | 1
14 | 2 | 4 | Myrtle | High Reference | 1 | 0 | -3.315 |  | -3.2696 |  |  | -3.2881 |  | -3.2881 | -2.2908 | -2.2476 | 10 | 1
15 | 2 | 5 | Myrtle | High Reference | 1 | 0 | -3.299 |  | -3.2535 |  |  | -3.2733 |  | -3.2733 | -2.2908 | -2.2333
16 | 2 | 6 | Myrtle | High Reference | 1 | 0 | -3.228 |  | -3.1819 |  |  | -3.203 |  | -3.203 | -2.2908 | -2.1655 | (2) Drift correction
17 | 2 | 7 | Myrtle | High Reference | 1 | 0 | -3.338 |  | -3.2928 |  |  | -3.3153 |  | -3.3153 | -2.2908 | -2.2738 | Slope m: | 0.0013225
18 | 2 | 8 | Myrtle | High Reference | 1 | 0 | -3.336 |  | -3.3173 |  |  | -3.3411 |  | -3.3411 | -2.2908 | -2.2988 | Stdev Mem Corr. | 0.074712
19 | 2 | 9 | Myrtle | High Reference | 1 | 0 | -3.36 |  | -3.36 |  |  | -3.3851 |  | -3.3851 | -2.2908 | -2.3412 | Stdev Drift Corr. | 0.042234
20 | 2 | 10 | Myrtle | High Reference | 1 | 0 | -3.251 | 0.16087 | -3.251 | 0.051733 |  | -3.2774 |  | -3.2774 | -2.2908 | -2.2373
21 | 3 | 1 | Homer | Low Reference | 1 | 0 | -14.813 |  | -15.935 |  |  | -15.9628 |  | -15.9628 | -14.5399 | -14.4755 | (3) Normalization to VSMOW scale
22 | 3 | 2 | Homer | Low Reference | 1 | 0 | -15.687 |  | -16.0056 |  |  | -16.0346 |  | -16.0346 | -14.5399 | -14.5448 | slope m | 0.96475
23 | 3 | 3 | Homer | Low Reference | 1 | 0 | -15.812 |  | -15.9883 |  |  | -16.0187 |  | -16.0187 | -14.5399 | -14.5294 | interception t | 0.92457
24 | 3 | 4 | Homer | Low Reference | 1 | 0 | -15.915 |  | -16.0182 |  |  | -16.0499 |  | -16.0499 | -14.5399 | -14.5595
25 | 3 | 5 | Homer | Low Reference | 1 | 0 | -15.9 |  | -16.0031 |  |  | -16.0361 |  | -16.0361 | -14.5399 | -14.5462 | (4) Quality control |  | offset
26 | 3 | 6 | Homer | Low Reference | 1 | 0 | -15.887 |  | -15.99 |  |  | -16.0243 |  | -16.0243 | -14.5399 | -14.5349 | QC Check Defined | -3.6108 | accepted ± 0.15
27 | 3 | 7 | Homer | Low Reference | 1 | 0 | -15.656 |  | -15.7571 |  |  | -15.7928 |  | -15.7928 | -14.5399 | -14.3115 | QC Check Vial 5 | -2.8635 | 0.74729
28 | 3 | 8 | Homer | Low Reference | 1 | 0 | -15.943 |  | -15.9856 |  |  | -16.0227 |  | -16.0227 | -14.5399 | -14.5333 | QC Check Vial 26 | -1.8096 | 1.8012
29 | 3 | 9 | Homer | Low Reference | 1 | 0 | -16.09 |  | -16.09 |  |  | -16.1284 |  | -16.1284 | -14.5399 | -14.6352
30 | 3 | 10 | Homer | Low Reference | 1 | 0 | -15.946 | 0.3578 | -15.946 | 0.086383 |  | -15.9857 |  | -15.9857 | -14.5399 | -14.4976 | Drift Defined | -7.5615 | accepted ± 0.15
31 | 4 | 1 | Blacksburg | Drift | 1 | 0 | -9.359 |  | -8.7198 |  |  | -8.7608 |  | -8.7608 | -7.5615 | -7.5274 | Drift Average | -7.6132 | -0.051707
32 | 4 | 2 | Blacksburg | Drift | 1 | 0 | -9.056 |  | -8.8795 |  |  | -8.9218 |  | -8.9218 | -7.5615 | -7.6827 | Drift Vial 4 | -7.6412 | -0.07967
33 | 4 | 3 | Blacksburg | Drift | 1 | 0 | -8.884 |  | -8.7849 |  |  | -8.8285 |  | -8.8285 | -7.5615 | -7.5927 | Drift Vial 9 | -7.6073 | -0.045791
34 | 4 | 4 | Blacksburg | Drift | 1 | 0 | -8.87 |  | -8.8123 |  |  | -8.8573 |  | -8.8573 | -7.5615 | -7.6205 | Drift Vial 18 | -7.5786 | -0.017127
35 | 4 | 5 | Blacksburg | Drift | 1 | 0 | -8.884 |  | -8.8265 |  |  | -8.8727 |  | -8.8727 | -7.5615 | -7.6354 | Drift Vial 27 | -7.6252 | -0.063759
36 | 4 | 6 | Blacksburg | Drift | 1 | 0 | -8.742 |  | -8.6833 |  |  | -8.7309 |  | -8.7309 | -7.5615 | -7.4986
37 | 4 | 7 | Blacksburg | Drift | 1 | 0 | -8.849 |  | -8.7912 |  |  | -8.8401 |  | -8.8401 | -7.5615 | -7.6039 | Reference Defined Values |  | d18O
38 | 4 | 8 | Blacksburg | Drift | 1 | 0 | -8.824 |  | -8.8001 |  |  | -8.8503 |  | -8.8503 | -7.5615 | -7.6138 | High Reference | Myrtle Beach | -2.2908
39 | 4 | 9 | Blacksburg | Drift | 1 | 0 | -8.865 |  | -8.865 |  |  | -8.9166 |  | -8.9166 | -7.5615 | -7.6777 | Low Reference | Homer | -14.5399
40 | 4 | 10 | Blacksburg | Drift | 1 | 0 | -8.855 | 0.17304 | -8.855 | 0.062142 |  | -8.9079 |  | -8.9079 | -7.5615 | -7.6693 | Drift | Blacksburg | -7.5615
41 | 5 | 1 | Hawaii | QC Check | 1 | 0 | -4.265 |  | -3.8196 |  |  | -3.8738 |  |  |  | -2.8127 | QC Check | Hawaii | -3.6108
42 | 5 | 2 | Hawaii | QC Check | 1 | 0 | -3.984 |  | -3.8592 |  |  | -3.9148 |  |  |  | -2.8522 | Calibration Date:
43 | 5 | 3 | Hawaii | QC Check | 1 | 0 | -3.982 |  | -3.9136 |  |  | -3.9705 |  |  |  | -2.9059
44 | 5 | 4 | Hawaii | QC Check | 1 | 0 | -3.929 | 0.15215 | -3.8889 | 0.040484 |  | -3.9471 |  |  |  | -2.8833
45 | 6 | 1 | CP1 22may24 BOT |  | 1 | 0 | -6.922 |  | -7.2125 |  |  | -7.272 |  |  |  | -6.091
46 | 6 | 2 | CP1 22may24 BOT |  | 1 | 0 | -7.218 |  | -7.3022 |  |  | -7.3631 |  |  |  | -6.1789 | Denali | -19.9287
47 | 6 | 3 | CP1 22may24 BOT |  | 1 | 0 | -7.142 |  | -7.1871 |  |  | -7.2492 |  |  |  | -6.0691 | Alberta | -18.8392
48 | 6 | 4 | CP1 22may24 BOT |  | 1 | 0 | -7.171 | 0.13129 | -7.1974 | 0.052671 |  | -7.2609 |  |  |  | -6.0804 | MT BHB | -18.431
49 | 7 | 1 | C50 16sep24 0.1m |  | 1 | 0 | -4.995 |  | -4.7838 |  |  | -4.8486 |  |  |  | -3.7531 | Anchorage | -16.4204
50 | 7 | 2 | C50 16sep24 0.1m |  | 1 | 0 | -4.827 |  | -4.767 |  |  | -4.8331 |  |  |  | -3.7381 | Homer | -14.5399
51 | 7 | 3 | C50 16sep24 0.1m |  | 1 | 0 | -4.781 |  | -4.7475 |  |  | -4.8149 |  |  |  | -3.7206 | MT Creek | -13.3562
52 | 7 | 4 | C50 16sep24 0.1m |  | 1 | 0 | -4.785 | 0.10084 | -4.7656 | 0.014865 |  | -4.8343 |  |  |  | -3.7393 | Nickel Mines | -7.876
53 | 8 | 1 | CC2 19jun24 0.1m |  | 1 | 0 | -5.712 |  | -5.802 |  |  | -5.8721 |  |  |  | -4.7405 | Blacksburg | -7.5615
54 | 8 | 2 | CC2 19jun24 0.1m |  | 1 | 0 | -5.742 |  | -5.7665 |  |  | -5.8379 |  |  |  | -4.7076 | Denton's Well | -6.87
55 | 8 | 3 | CC2 19jun24 0.1m |  | 1 | 0 | -5.734 |  | -5.7473 |  |  | -5.8201 |  |  |  | -4.6903 | Hawaii | -3.6108
56 | 8 | 4 | CC2 19jun24 0.1m |  | 1 | 0 | -5.813 | 0.043714 | -5.8214 | 0.033518 |  | -5.8954 |  |  |  | -4.763 | Myrtle Beach | -2.2908
57 | 9 | 1 | Blacksburg | Drift | 1 | 0 | -8.516 |  | -8.7783 |  | -8.7783 | -8.8537 | -8.8537 |  |  | -7.617 | Heavy Oxygen | 3.1753
58 | 9 | 2 | Blacksburg | Drift | 1 | 0 | -8.655 |  | -8.7278 |  | -8.7278 | -8.8045 | -8.8045 |  |  | -7.5695 | Heavy Hydrogen | 11.1372
59 | 9 | 3 | Blacksburg | Drift | 1 | 0 | -8.712 |  | -8.7527 |  | -8.7527 | -8.8307 | -8.8307 |  |  | -7.5948 | Andres Heavy Water | 17.5664
60 | 9 | 4 | Blacksburg | Drift | 1 | 0 | -8.782 | 0.11283 | -8.8062 | 0.03368 | -8.8062 | -8.8855 | -8.8855 |  |  | -7.6477 | Picarro ZERO | 0.3
61 | 10 | 1 | CP1 28oct24 0.1m |  | 1 | 0 | -7.303 |  | -7.1595 |  |  | -7.2401 |  |  |  | -6.0603 | Picarro MID | -20.6
62 | 10 | 2 | CP1 28oct24 0.1m |  | 1 | 0 | -7.235 |  | -7.1954 |  |  | -7.2774 |  |  |  | -6.0962 | Picarro DEPL | -29.6
63 | 10 | 3 | CP1 28oct24 0.1m |  | 1 | 0 | -7.231 |  | -7.2092 |  |  | -7.2925 |  |  |  | -6.1109
64 | 10 | 4 | CP1 28oct24 0.1m |  | 1 | 0 | -7.123 | 0.074404 | -7.1095 | 0.044519 |  | -7.1941 |  |  |  | -6.0159
65 | 11 | 1 | CC4 11jul24 9m |  | 1 | 0 | -6.451 |  | -6.3858 |  |  | -6.4717 |  |  |  | -5.319
66 | 11 | 2 | CC4 11jul24 9m |  | 1 | 0 | -6.405 |  | -6.3866 |  |  | -6.4739 |  |  |  | -5.3211
67 | 11 | 3 | CC4 11jul24 9m |  | 1 | 0 | -6.435 |  | -6.4253 |  |  | -6.5139 |  |  |  | -5.3597
68 | 11 | 4 | CC4 11jul24 9m |  | 1 | 0 | -6.342 | 0.048107 | -6.3356 | 0.036772 |  | -6.4256 |  |  |  | -5.2745
69 | 12 | 1 | CC4 15aug24 BOT |  | 1 | 0 | -6.204 |  | -6.1906 |  |  | -6.2819 |  |  |  | -5.1358
70 | 12 | 2 | CC4 15aug24 BOT |  | 1 | 0 | -6.244 |  | -6.2415 |  |  | -6.3341 |  |  |  | -5.1862
71 | 12 | 3 | CC4 15aug24 BOT |  | 1 | 0 | -6.293 |  | -6.2923 |  |  | -6.3862 |  |  |  | -5.2365
72 | 12 | 4 | CC4 15aug24 BOT |  | 1 | 0 | -6.195 | 0.044729 | -6.1938 | 0.047875 |  | -6.289 |  |  |  | -5.1427
73 | 13 | 1 | C50 16apr25 0.1m |  | 1 | 0 | -6.095 |  | -6.0853 |  |  | -6.1818 |  |  |  | -5.0393
74 | 13 | 2 | C50 16apr25 0.1m |  | 1 | 0 | -6.04 |  | -6.036 |  |  | -6.1339 |  |  |  | -4.9931
75 | 13 | 3 | C50 16apr25 0.1m |  | 1 | 0 | -6.167 |  | -6.1666 |  |  | -6.2658 |  |  |  | -5.1203
76 | 13 | 4 | C50 16apr25 0.1m |  | 1 | 0 | -6.116 | 0.052564 | -6.1154 | 0.054705 |  | -6.2159 |  |  |  | -5.0722
77 | 14 | 1 | C50 17feb25 1.5m |  | 1 | 0 | -6.08 |  | -6.0765 |  |  | -6.1783 |  |  |  | -5.036
78 | 14 | 2 | C50 17feb25 1.5m |  | 1 | 0 | -6.158 |  | -6.1591 |  |  | -6.2622 |  |  |  | -5.1169
79 | 14 | 3 | C50 17feb25 1.5m |  | 1 | 0 | -6.142 |  | -6.1424 |  |  | -6.2468 |  |  |  | -5.1021
80 | 14 | 4 | C50 17feb25 1.5m |  | 1 | 0 | -6.116 | 0.034059 | -6.116 | 0.035992 |  | -6.2218 |  |  |  | -5.0779
81 | 15 | 1 | C50 11jul24 9m |  | 1 | 0 | -6.171 |  | -6.1763 |  |  | -6.2835 |  |  |  | -5.1374
82 | 15 | 2 | C50 11jul24 9m |  | 1 | 0 | -6.225 |  | -6.2278 |  |  | -6.3362 |  |  |  | -5.1883
83 | 15 | 3 | C50 11jul24 9m |  | 1 | 0 | -6.235 |  | -6.2367 |  |  | -6.3464 |  |  |  | -5.1981
84 | 15 | 4 | C50 11jul24 9m |  | 1 | 0 | -6.128 | 0.049849 | -6.1281 | 0.050349 |  | -6.2392 |  |  |  | -5.0947
85 | 16 | 1 | CC4 22may24 0.1m |  | 1 | 0 | -5.938 |  | -5.9196 |  |  | -6.032 |  |  |  | -4.8948
86 | 16 | 2 | CC4 22may24 0.1m |  | 1 | 0 | -5.954 |  | -5.9495 |  |  | -6.0633 |  |  |  | -4.925
87 | 16 | 3 | CC4 22may24 0.1m |  | 1 | 0 | -5.968 |  | -5.9658 |  |  | -6.0808 |  |  |  | -4.9419
88 | 16 | 4 | CC4 22may24 0.1m |  | 1 | 0 | -6.009 | 0.030412 | -6.008 | 0.03689 |  | -6.1244 |  |  |  | -4.9839
89 | 17 | 1 | C50 22may24 0.1m |  | 1 | 0 | -6.043 |  | -6.0463 |  |  | -6.164 |  |  |  | -5.0221
90 | 17 | 2 | C50 22may24 0.1m |  | 1 | 0 | -5.985 |  | -5.9844 |  |  | -6.1034 |  |  |  | -4.9637
91 | 17 | 3 | C50 22may24 0.1m |  | 1 | 0 | -6.047 |  | -6.0475 |  |  | -6.1679 |  |  |  | -5.0259
92 | 17 | 4 | C50 22may24 0.1m |  | 1 | 0 | -5.989 | 0.033566 | -5.9888 | 0.034868 |  | -6.1105 |  |  |  | -4.9705
93 | 18 | 1 | Blacksburg | Drift | 1 | 0 | -8.502 |  | -8.7459 |  | -8.7459 | -8.8689 | -8.8689 |  |  | -7.6316
94 | 18 | 2 | Blacksburg | Drift | 1 | 0 | -8.614 |  | -8.6812 |  | -8.6812 | -8.8056 | -8.8056 |  |  | -7.5706
95 | 18 | 3 | Blacksburg | Drift | 1 | 0 | -8.618 |  | -8.6549 |  | -8.6549 | -8.7805 | -8.7805 |  |  | -7.5464
96 | 18 | 4 | Blacksburg | Drift | 1 | 0 | -8.652 | 0.065266 | -8.6737 | 0.039547 | -8.6737 | -8.8007 | -8.8007 |  |  | -7.5658
97 | 19 | 1 | CS2 22may24 0.1m |  | 1 | 0 | -7.862 |  | -7.7853 |  |  | -7.9136 |  |  |  | -6.7101
98 | 19 | 2 | CS2 22may24 0.1m |  | 1 | 0 | -7.868 |  | -7.8479 |  |  | -7.9775 |  |  |  | -6.7717
99 | 19 | 3 | CS2 22may24 0.1m |  | 1 | 0 | -7.875 |  | -7.8641 |  |  | -7.995 |  |  |  | -6.7886
100 | 19 | 4 | CS2 22may24 0.1m |  | 1 | 0 | -7.845 | 0.012819 | -7.8384 | 0.034094 |  | -7.9707 |  |  |  | -6.7651
101 | 20 | 1 | CCR raindeploy - 11jul24 R1 |  | 1 | 0 | -3.507 |  | -3.086 |  |  | -3.2196 |  |  |  | -2.1815
102 | 20 | 2 | CCR raindeploy - 11jul24 R1 |  | 1 | 0 | -3.21 |  | -3.0913 |  |  | -3.2262 |  |  |  | -2.1879
103 | 20 | 3 | CCR raindeploy - 11jul24 R1 |  | 1 | 0 | -3.123 |  | -3.0567 |  |  | -3.1929 |  |  |  | -2.1558
104 | 20 | 4 | CCR raindeploy - 11jul24 R1 |  | 1 | 0 | -3.085 | 0.19113 | -3.0462 | 0.021992 |  | -3.1838 |  |  |  | -2.1469
105 | 21 | 1 | C50 16apr25 6m |  | 1 | 0 | -5.875 |  | -6.1458 |  |  | -6.2846 |  |  |  | -5.1385
106 | 21 | 2 | C50 16apr25 6m |  | 1 | 0 | -6.012 |  | -6.087 |  |  | -6.2272 |  |  |  | -5.0831
107 | 21 | 3 | C50 16apr25 6m |  | 1 | 0 | -6.06 |  | -6.1018 |  |  | -6.2433 |  |  |  | -5.0986
108 | 21 | 4 | C50 16apr25 6m |  | 1 | 0 | -6.098 | 0.097411 | -6.1225 | 0.025569 |  | -6.2654 |  |  |  | -5.1199
109 | 22 | 1 | CC4 16sep24 1.5m |  | 1 | 0 | -4.874 |  | -4.7552 |  |  | -4.8994 |  |  |  | -3.8021
110 | 22 | 2 | CC4 16sep24 1.5m |  | 1 | 0 | -4.738 |  | -4.7032 |  |  | -4.8486 |  |  |  | -3.7531
111 | 22 | 3 | CC4 16sep24 1.5m |  | 1 | 0 | -4.672 |  | -4.652 |  |  | -4.7988 |  |  |  | -3.705
112 | 22 | 4 | CC4 16sep24 1.5m |  | 1 | 0 | -4.573 | 0.12625 | -4.5606 | 0.082945 |  | -4.7087 |  |  |  | -3.6181
113 | 23 | 1 | C50 19jun24 1.5m |  | 1 | 0 | -5.511 |  | -5.602 |  |  | -5.7515 |  |  |  | -4.6241
114 | 23 | 2 | C50 19jun24 1.5m |  | 1 | 0 | -5.507 |  | -5.5309 |  |  | -5.6817 |  |  |  | -4.5568
115 | 23 | 3 | C50 19jun24 1.5m |  | 1 | 0 | -5.496 |  | -5.509 |  |  | -5.661 |  |  |  | -4.5369
116 | 23 | 4 | C50 19jun24 1.5m |  | 1 | 0 | -5.532 | 0.015067 | -5.5398 | 0.039899 |  | -5.6932 |  |  |  | -4.5679
117 | 24 | 1 | CP2 22may24 0.1m |  | 1 | 0 | -5.827 |  | -5.8556 |  |  | -6.0104 |  |  |  | -4.8739
118 | 24 | 2 | CP2 22may24 0.1m |  | 1 | 0 | -5.897 |  | -5.9063 |  |  | -6.0624 |  |  |  | -4.9241
119 | 24 | 3 | CP2 22may24 0.1m |  | 1 | 0 | -5.83 |  | -5.8342 |  |  | -5.9916 |  |  |  | -4.8558
120 | 24 | 4 | CP2 22may24 0.1m |  | 1 | 0 | -5.896 | 0.039281 | -5.899 | 0.034605 |  | -6.0577 |  |  |  | -4.9195
121 | 25 | 1 | CC4 22may24 0.1m - DUP |  | 1 | 0 | -5.861 |  | -5.8576 |  |  | -6.0176 |  |  |  | -4.8809
122 | 25 | 2 | CC4 22may24 0.1m - DUP |  | 1 | 0 | -5.773 |  | -5.7698 |  |  | -5.9312 |  |  |  | -4.7975
123 | 25 | 3 | CC4 22may24 0.1m - DUP |  | 1 | 0 | -5.854 |  | -5.8534 |  |  | -6.0161 |  |  |  | -4.8794
124 | 25 | 4 | CC4 22may24 0.1m - DUP |  | 1 | 0 | -5.867 | 0.044154 | -5.8668 | 0.045051 |  | -6.0307 |  |  |  | -4.8936
125 | 26 | 1 | Hawaii |  | 1 | 0 | -2.935 |  | -2.6505 |  |  | -2.8158 |  |  |  | -1.7919
126 | 26 | 2 | Hawaii |  | 1 | 0 | -2.729 |  | -2.6486 |  |  | -2.8152 |  |  |  | -1.7914
127 | 26 | 3 | Hawaii |  | 1 | 0 | -2.706 |  | -2.6616 |  |  | -2.8296 |  |  |  | -1.8053
128 | 26 | 4 | Hawaii |  | 1 | 0 | -2.732 | 0.10697 | -2.7065 | 0.02706 |  | -2.8757 |  |  |  | -1.8498
129 | 27 | 1 | Blacksburg | Drift | 1 | 0 | -8.166 |  | -8.6933 |  | -8.6933 | -8.8639 | -8.8639 |  |  | -7.6269
130 | 27 | 2 | Blacksburg | Drift | 1 | 0 | -8.554 |  | -8.7031 |  | -8.7031 | -8.8751 | -8.8751 |  |  | -7.6376
131 | 27 | 3 | Blacksburg | Drift | 1 | 0 | -8.552 |  | -8.6337 |  | -8.6337 | -8.8069 | -8.8069 |  |  | -7.5719
132 | 27 | 4 | Blacksburg | Drift | 1 | 0 | -8.68 | 0.22286 | -8.7285 | 0.040142 | -8.7285 | -8.903 | -8.903 |  |  | -7.6646
Note on G2: Manually change first six injections to ignore "-1".
Note on G3: Manually change first six injections to ignore "-1".
Note on G4: Manually change first six injections to ignore "-1".
Note on G5: Manually change first six injections to ignore "-1".
Note on G6: Manually change first six injections to ignore "-1".
Note on G7: Manually change first six injections to ignore "-1".

## 2H
Line | Sample | Injection | Sample Name | Correction Type | Good | Ignore | dD_H) Picarro Raw | Stdev Raw | d(D_H) Memory Corrected | Stdev Memory Corrected | d(D_H) Memory Corrected Drift | d(D_H) Drift Corrected | Drift Corrected Drift | d(D_H) Drift Corrected High, Low, Drift | d(D_H) Reference Defined | dD Final Corrected Value | Post run corrections:
1 | 1 | 1 | Blacksburg | Drift | 0 | -1 | -46.973 | [9 empty] | (1) Memory correction |  |  | (1) Graphs for memory correction
2 | 1 | 2 | Blacksburg | Drift | 1 | -1 | -45.287 | [9 empty] | Combined Stdev | 0.50451
3 | 1 | 3 | Blacksburg | Drift | 1 | -1 | -44.776 | [9 empty] | Date:
4 | 1 | 4 | Blacksburg | Drift | 1 | -1 | -44.877 | [9 empty] | injection | 18O memory coef
5 | 1 | 5 | Blacksburg | Drift | 1 | -1 | -44.464 | [9 empty] | 1 | 0.83639
6 | 1 | 6 | Blacksburg | Drift | 1 | -1 | -44.194 | [9 empty] | 2 | 0.95847
7 | 1 | 7 | Blacksburg | Drift | 1 | 0 | -43.957 |  | -43.957 |  | -43.957 | -43.9677 | -43.9677 |  |  | -48.0428 | 3 | 0.98129
8 | 1 | 8 | Blacksburg | Drift | 1 | 0 | -44.191 |  | -44.191 |  | -44.191 | -44.2032 | -44.2032 |  |  | -48.2692 | 4 | 0.98829
9 | 1 | 9 | Blacksburg | Drift | 1 | 0 | -43.925 |  | -43.925 |  | -43.925 | -43.9388 | -43.9388 |  |  | -48.015 | 5 | 0.99077
10 | 1 | 10 | Blacksburg | Drift | 1 | 0 | -43.99 | 0.11981 | -43.99 | 0.11981 | -43.99 | -44.0053 | -44.0053 |  |  | -48.0789 | 6 | 0.99077
11 | 2 | 1 | Myrtle | High Reference | 1 | 0 | -9.933 |  | -4.361 |  |  | -4.3778 |  | -4.3778 | -10.453 | -9.9968 | 7 | 0.99077
12 | 2 | 2 | Myrtle | High Reference | 1 | 0 | -6.116 |  | -4.5429 |  |  | -4.5613 |  | -4.5613 | -10.453 | -10.1731 | 8 | 0.99494
13 | 2 | 3 | Myrtle | High Reference | 1 | 0 | -5.319 |  | -4.5953 |  |  | -4.6152 |  | -4.6152 | -10.453 | -10.2249 | 9 | 1
14 | 2 | 4 | Myrtle | High Reference | 1 | 0 | -4.857 |  | -4.3989 |  |  | -4.4204 |  | -4.4204 | -10.453 | -10.0377 | 10 | 1
15 | 2 | 5 | Myrtle | High Reference | 1 | 0 | -4.853 |  | -4.4919 |  |  | -4.5149 |  | -4.5149 | -10.453 | -10.1285
16 | 2 | 6 | Myrtle | High Reference | 1 | 0 | -4.782 |  | -4.4202 |  |  | -4.4447 |  | -4.4447 | -10.453 | -10.0611 | (2) Drift correction
17 | 2 | 7 | Myrtle | High Reference | 1 | 0 | -4.586 |  | -4.2224 |  |  | -4.2485 |  | -4.2485 | -10.453 | -9.8725 | Slope m: | 0.0015305
18 | 2 | 8 | Myrtle | High Reference | 1 | 0 | -4.557 |  | -4.3576 |  |  | -4.3852 |  | -4.3852 | -10.453 | -10.0039 | Stdev Mem Corr. | 0.2298
19 | 2 | 9 | Myrtle | High Reference | 1 | 0 | -4.473 |  | -4.473 |  |  | -4.5021 |  | -4.5021 | -10.453 | -10.1162 | Stdev Drift Corr. | 0.21845
20 | 2 | 10 | Myrtle | High Reference | 1 | 0 | -4.442 | 1.6731 | -4.442 | 0.10543 |  | -4.4726 |  | -4.4726 | -10.453 | -10.0879
21 | 3 | 1 | Homer | Low Reference | 1 | 0 | -93.78 |  | -108.3964 |  |  | -108.4286 |  | -108.4286 | -110.5169 | -109.9899 | (3) Normalization to VSMOW scale
22 | 3 | 2 | Homer | Low Reference | 1 | 0 | -104.711 |  | -108.8756 |  |  | -108.9093 |  | -108.9093 | -110.5169 | -110.4519 | slope m | 0.961
23 | 3 | 3 | Homer | Low Reference | 1 | 0 | -106.635 |  | -108.5474 |  |  | -108.5826 |  | -108.5826 | -110.5169 | -110.138 | interception t | -5.7897
24 | 3 | 4 | Homer | Low Reference | 1 | 0 | -107.535 |  | -108.7417 |  |  | -108.7785 |  | -108.7785 | -110.5169 | -110.3262
25 | 3 | 5 | Homer | Low Reference | 1 | 0 | -108.061 |  | -109.017 |  |  | -109.0553 |  | -109.0553 | -110.5169 | -110.5922 | (4) Quality control |  | offset
26 | 3 | 6 | Homer | Low Reference | 1 | 0 | -108.108 |  | -109.0645 |  |  | -109.1043 |  | -109.1043 | -110.5169 | -110.6393 | QC Check Defined | -13.2294 | accepted ±1.5
27 | 3 | 7 | Homer | Low Reference | 1 | 0 | -106.868 |  | -107.813 |  |  | -107.8544 |  | -107.8544 | -110.5169 | -109.4381 | QC Check Vial 5 | -11.1403 | 2.0891
28 | 3 | 8 | Homer | Low Reference | 1 | 0 | -108.093 |  | -108.617 |  |  | -108.6599 |  | -108.6599 | -110.5169 | -110.2122 | QC Check Vial 26 | -8.3551 | 4.8743
29 | 3 | 9 | Homer | Low Reference | 1 | 0 | -108.993 |  | -108.993 |  |  | -109.0374 |  | -109.0374 | -110.5169 | -110.575
30 | 3 | 10 | Homer | Low Reference | 1 | 0 | -108.845 | 4.526 | -108.845 | 0.37728 |  | -108.8909 |  | -108.8909 | -110.5169 | -110.4343 | Drift Defined | -47.446 | accepted ±1.5
31 | 4 | 1 | Blacksburg | Drift | 1 | 0 | -52.697 |  | -43.5107 |  |  | -43.5582 |  | -43.5582 | -47.446 | -47.6492 | Drift Average | -48.0815 | -0.63548
32 | 4 | 2 | Blacksburg | Drift | 1 | 0 | -46.746 |  | -44.1668 |  |  | -44.2157 |  | -44.2157 | -47.446 | -48.2812 | Drift Vial 4 | -48.1621 | -0.71612
33 | 4 | 3 | Blacksburg | Drift | 1 | 0 | -44.807 |  | -43.6086 |  |  | -43.6591 |  | -43.6591 | -47.446 | -47.7462 | Drift Vial 9 | -47.9228 | -0.47677
34 | 4 | 4 | Blacksburg | Drift | 1 | 0 | -44.788 |  | -44.0382 |  |  | -44.0902 |  | -44.0902 | -47.446 | -48.1606 | Drift Vial 18 | -48.1984 | -0.75234
35 | 4 | 5 | Blacksburg | Drift | 1 | 0 | -44.922 |  | -44.3322 |  |  | -44.3858 |  | -44.3858 | -47.446 | -48.4446 | Drift Vial 27 | -48.0227 | -0.57669
36 | 4 | 6 | Blacksburg | Drift | 1 | 0 | -44.182 |  | -43.5854 |  |  | -43.6405 |  | -43.6405 | -47.446 | -47.7284
37 | 4 | 7 | Blacksburg | Drift | 1 | 0 | -44.32 |  | -43.7247 |  |  | -43.7813 |  | -43.7813 | -47.446 | -47.8637 | Reference Defined Values |  | d2H
38 | 4 | 8 | Blacksburg | Drift | 1 | 0 | -44.19 |  | -43.8631 |  |  | -43.9213 |  | -43.9213 | -47.446 | -47.9982 | High Reference | North Myrtle Beach, SC | -10.453
39 | 4 | 9 | Blacksburg | Drift | 1 | 0 | -44.306 |  | -44.306 |  |  | -44.3657 |  | -44.3657 | -47.446 | -48.4253 | Low Reference | Homer, AK | -110.5169
40 | 4 | 10 | Blacksburg | Drift | 1 | 0 | -44.238 | 2.6359 | -44.238 | 0.31793 |  | -44.2992 |  | -44.2992 | -47.446 | -48.3614 | Drift | Blacksburg | -47.446
41 | 5 | 1 | Hawaii | QC Check | 1 | 0 | -10.89 |  | -5.434 |  |  | -5.4967 |  |  |  | -11.0721 | QC Check | Hawaii | -13.2294
42 | 5 | 2 | Hawaii | QC Check | 1 | 0 | -6.882 |  | -5.3304 |  |  | -5.3947 |  |  |  | -10.974 | Calibration Date:
43 | 5 | 3 | Hawaii | QC Check | 1 | 0 | -6.409 |  | -5.7011 |  |  | -5.7669 |  |  |  | -11.3317
44 | 5 | 4 | Hawaii | QC Check | 1 | 0 | -5.993 | 2.2604 | -5.5453 | 0.15871 |  | -5.6127 |  |  |  | -11.1835
45 | 6 | 1 | CP1 22may24 BOT |  | 1 | 0 | -27.362 |  | -30.8581 |  |  | -30.927 |  |  |  | -35.5107
46 | 6 | 2 | CP1 22may24 BOT |  | 1 | 0 | -29.471 |  | -30.4461 |  |  | -30.5165 |  |  |  | -35.1162 | Denali National Park | -153.7172
47 | 6 | 3 | CP1 22may24 BOT |  | 1 | 0 | -30.48 |  | -30.9382 |  |  | -31.0102 |  |  |  | -35.5906 | Edmonton, Alberta | -146.8245
48 | 6 | 4 | CP1 22may24 BOT |  | 1 | 0 | -30.276 | 1.425 | -30.5602 | 0.23505 |  | -30.6337 |  |  |  | -35.2288 | MT BHB | -140.5778
49 | 7 | 1 | C50 16sep24 0.1m |  | 1 | 0 | -21.571 |  | -20.1468 |  |  | -20.2218 |  |  |  | -25.2229 | Anchorage, AK | -129.9028
50 | 7 | 2 | C50 16sep24 0.1m |  | 1 | 0 | -20.29 |  | -19.8752 |  |  | -19.9518 |  |  |  | -24.9634 | Homer, AK | -110.5169
51 | 7 | 3 | C50 16sep24 0.1m |  | 1 | 0 | -20.114 |  | -19.9238 |  |  | -20.0019 |  |  |  | -25.0116 | MT Creek | -97.7428
52 | 7 | 4 | C50 16sep24 0.1m |  | 1 | 0 | -20.188 | 0.69061 | -20.0699 | 0.12615 |  | -20.1495 |  |  |  | -25.1534 | Nickel Mines, PA | -48.7951
53 | 8 | 1 | CC2 19jun24 0.1m |  | 1 | 0 | -23.364 |  | -23.8836 |  |  | -23.9647 |  |  |  | -28.8199 | Blacksburg | -47.446
54 | 8 | 2 | CC2 19jun24 0.1m |  | 1 | 0 | -23.808 |  | -23.9584 |  |  | -24.041 |  |  |  | -28.8932 | Denton's Well | -41.771
55 | 8 | 3 | CC2 19jun24 0.1m |  | 1 | 0 | -23.751 |  | -23.8177 |  |  | -23.9019 |  |  |  | -28.7595 | Hawaii | -13.2294
56 | 8 | 4 | CC2 19jun24 0.1m |  | 1 | 0 | -24.158 | 0.32516 | -24.2045 | 0.16903 |  | -24.2902 |  |  |  | -29.1326 | North Myrtle Beach, SC | -10.453
57 | 9 | 1 | Blacksburg | Drift | 1 | 0 | -40.729 |  | -43.4402 |  | -43.4402 | -43.5274 | -43.5274 |  |  | -47.6197 | WQL Heavy Oxygen | -15.7881
58 | 9 | 2 | Blacksburg | Drift | 1 | 0 | -43.047 |  | -43.8315 |  | -43.8315 | -43.9203 | -43.9203 |  |  | -47.9973 | WQL Heavy Hydrogen | 4.0332
59 | 9 | 3 | Blacksburg | Drift | 1 | 0 | -43.539 |  | -43.9017 |  | -43.9017 | -43.992 | -43.992 |  |  | -48.0662 | Andres Heavy Water | 20.5382
60 | 9 | 4 | Blacksburg | Drift | 1 | 0 | -43.612 | 1.3585 | -43.8397 | 0.21109 | -43.8397 | -43.9315 | -43.9315 |  |  | -48.0081
61 | 10 | 1 | CP1 28oct24 0.1m |  | 1 | 0 | -31.612 |  | -29.6487 |  |  | -29.7421 |  |  |  | -34.3719
62 | 10 | 2 | CP1 28oct24 0.1m |  | 1 | 0 | -30.47 |  | -29.9242 |  |  | -30.019 |  |  |  | -34.6381
63 | 10 | 3 | CP1 28oct24 0.1m |  | 1 | 0 | -30.183 |  | -29.9317 |  |  | -30.0281 |  |  |  | -34.6468
64 | 10 | 4 | CP1 28oct24 0.1m |  | 1 | 0 | -29.866 | 0.76061 | -29.7051 | 0.14678 |  | -29.8031 |  |  |  | -34.4305
65 | 11 | 1 | CC4 11jul24 9m |  | 1 | 0 | -27.991 |  | -27.6842 |  |  | -27.7837 |  |  |  | -32.4899
66 | 11 | 2 | CC4 11jul24 9m |  | 1 | 0 | -27.818 |  | -27.7329 |  |  | -27.834 |  |  |  | -32.5382
67 | 11 | 3 | CC4 11jul24 9m |  | 1 | 0 | -27.54 |  | -27.4965 |  |  | -27.599 |  |  |  | -32.3124
68 | 11 | 4 | CC4 11jul24 9m |  | 1 | 0 | -27.398 | 0.26752 | -27.3691 | 0.16868 |  | -27.4732 |  |  |  | -32.1915
69 | 12 | 1 | CC4 15aug24 BOT |  | 1 | 0 | -27.428 |  | -27.4329 |  |  | -27.5385 |  |  |  | -32.2543
70 | 12 | 2 | CC4 15aug24 BOT |  | 1 | 0 | -27.059 |  | -27.0449 |  |  | -27.1521 |  |  |  | -31.8829
71 | 12 | 3 | CC4 15aug24 BOT |  | 1 | 0 | -27.31 |  | -27.3084 |  |  | -27.417 |  |  |  | -32.1375
72 | 12 | 4 | CC4 15aug24 BOT |  | 1 | 0 | -26.915 | 0.23328 | -26.9093 | 0.2393 |  | -27.0195 |  |  |  | -31.7556
73 | 13 | 1 | C50 16apr25 0.1m |  | 1 | 0 | -25.988 |  | -25.8363 |  |  | -25.9481 |  |  |  | -30.7259
74 | 13 | 2 | C50 16apr25 0.1m |  | 1 | 0 | -26.054 |  | -26.0182 |  |  | -26.1315 |  |  |  | -30.9021
75 | 13 | 3 | C50 16apr25 0.1m |  | 1 | 0 | -26.148 |  | -26.1336 |  |  | -26.2484 |  |  |  | -31.0145
76 | 13 | 4 | C50 16apr25 0.1m |  | 1 | 0 | -25.887 | 0.10993 | -25.875 | 0.13655 |  | -25.9913 |  |  |  | -30.7674
77 | 14 | 1 | C50 17feb25 1.5m |  | 1 | 0 | -26.309 |  | -26.378 |  |  | -26.4959 |  |  |  | -31.2523
78 | 14 | 2 | C50 17feb25 1.5m |  | 1 | 0 | -25.973 |  | -25.9766 |  |  | -26.096 |  |  |  | -30.868
79 | 14 | 3 | C50 17feb25 1.5m |  | 1 | 0 | -25.61 |  | -25.6048 |  |  | -25.7257 |  |  |  | -30.5122
80 | 14 | 4 | C50 17feb25 1.5m |  | 1 | 0 | -26.297 | 0.33045 | -26.3018 | 0.35292 |  | -26.4242 |  |  |  | -31.1835
81 | 15 | 1 | C50 11jul24 9m |  | 1 | 0 | -26.825 |  | -26.9114 |  |  | -27.0354 |  |  |  | -31.7708
82 | 15 | 2 | C50 11jul24 9m |  | 1 | 0 | -26.585 |  | -26.597 |  |  | -26.7225 |  |  |  | -31.4701
83 | 15 | 3 | C50 11jul24 9m |  | 1 | 0 | -26.818 |  | -26.8277 |  |  | -26.9548 |  |  |  | -31.6933
84 | 15 | 4 | C50 11jul24 9m |  | 1 | 0 | -26.683 | 0.11545 | -26.6875 | 0.14056 |  | -26.8161 |  |  |  | -31.56
85 | 16 | 1 | CC4 22may24 0.1m |  | 1 | 0 | -25.702 |  | -25.5415 |  |  | -25.6716 |  |  |  | -30.4602
86 | 16 | 2 | CC4 22may24 0.1m |  | 1 | 0 | -25.824 |  | -25.7883 |  |  | -25.9199 |  |  |  | -30.6988
87 | 16 | 3 | CC4 22may24 0.1m |  | 1 | 0 | -25.601 |  | -25.5808 |  |  | -25.7139 |  |  |  | -30.5008
88 | 16 | 4 | CC4 22may24 0.1m |  | 1 | 0 | -25.596 | 0.10726 | -25.5833 | 0.11156 |  | -25.718 |  |  |  | -30.5047
89 | 17 | 1 | C50 22may24 0.1m |  | 1 | 0 | -25.534 |  | -25.5239 |  |  | -25.6601 |  |  |  | -30.4491
90 | 17 | 2 | C50 22may24 0.1m |  | 1 | 0 | -25.518 |  | -25.5148 |  |  | -25.6525 |  |  |  | -30.4418
91 | 17 | 3 | C50 22may24 0.1m |  | 1 | 0 | -25.52 |  | -25.5186 |  |  | -25.6579 |  |  |  | -30.447
92 | 17 | 4 | C50 22may24 0.1m |  | 1 | 0 | -25.618 | 0.047536 | -25.6183 | 0.049736 |  | -25.7591 |  |  |  | -30.5442
93 | 18 | 1 | Blacksburg | Drift | 1 | 0 | -41.551 |  | -44.1578 |  | -44.1578 | -44.3001 | -44.3001 |  |  | -48.3622
94 | 18 | 2 | Blacksburg | Drift | 1 | 0 | -43.09 |  | -43.8157 |  | -43.8157 | -43.9596 | -43.9596 |  |  | -48.035
95 | 18 | 3 | Blacksburg | Drift | 1 | 0 | -43.706 |  | -44.0445 |  | -44.0445 | -44.1899 | -44.1899 |  |  | -48.2563
96 | 18 | 4 | Blacksburg | Drift | 1 | 0 | -43.71 | 1.0181 | -43.9218 | 0.14838 | -43.9218 | -44.0687 | -44.0687 |  |  | -48.1399
97 | 19 | 1 | CS2 22may24 0.1m |  | 1 | 0 | -35.696 |  | -34.3848 |  |  | -34.5333 |  |  |  | -38.9763
98 | 19 | 2 | CS2 22may24 0.1m |  | 1 | 0 | -34.811 |  | -34.4414 |  |  | -34.5914 |  |  |  | -39.0321
99 | 19 | 3 | CS2 22may24 0.1m |  | 1 | 0 | -34.657 |  | -34.4876 |  |  | -34.6391 |  |  |  | -39.078
100 | 19 | 4 | CS2 22may24 0.1m |  | 1 | 0 | -34.418 | 0.55762 | -34.3092 | 0.076852 |  | -34.4623 |  |  |  | -38.9081
101 | 20 | 1 | CCR raindeploy - 11jul24 R1 |  | 1 | 0 | -5.04 |  | -0.23352 |  |  | -0.3881 |  |  |  | -6.1626
102 | 20 | 2 | CCR raindeploy - 11jul24 R1 |  | 1 | 0 | -2.149 |  | -0.80873 |  |  | -0.96484 |  |  |  | -6.7169
103 | 20 | 3 | CCR raindeploy - 11jul24 R1 |  | 1 | 0 | -1.468 |  | -0.85138 |  |  | -1.009 |  |  |  | -6.7594
104 | 20 | 4 | CCR raindeploy - 11jul24 R1 |  | 1 | 0 | -1.26 | 1.7489 | -0.87188 | 0.30637 |  | -1.0311 |  |  |  | -6.7805
105 | 21 | 1 | C50 16apr25 6m |  | 1 | 0 | -22.671 |  | -26.174 |  |  | -26.3347 |  |  |  | -31.0974
106 | 21 | 2 | C50 16apr25 6m |  | 1 | 0 | -25.111 |  | -26.1016 |  |  | -26.2639 |  |  |  | -31.0294
107 | 21 | 3 | C50 16apr25 6m |  | 1 | 0 | -25.624 |  | -26.0799 |  |  | -26.2437 |  |  |  | -31.01
108 | 21 | 4 | C50 16apr25 6m |  | 1 | 0 | -25.716 | 1.4313 | -26.0023 | 0.070691 |  | -26.1676 |  |  |  | -30.9368
109 | 22 | 1 | CC4 16sep24 1.5m |  | 1 | 0 | -20.285 |  | -19.3964 |  |  | -19.5633 |  |  |  | -24.5901
110 | 22 | 2 | CC4 16sep24 1.5m |  | 1 | 0 | -19.591 |  | -19.3366 |  |  | -19.505 |  |  |  | -24.534
111 | 22 | 3 | CC4 16sep24 1.5m |  | 1 | 0 | -19.608 |  | -19.4937 |  |  | -19.6636 |  |  |  | -24.6865
112 | 22 | 4 | CC4 16sep24 1.5m |  | 1 | 0 | -19.822 | 0.32324 | -19.753 | 0.18382 |  | -19.9244 |  |  |  | -24.9371
113 | 23 | 1 | C50 19jun24 1.5m |  | 1 | 0 | -23.047 |  | -23.5746 |  |  | -23.7476 |  |  |  | -28.6112
114 | 23 | 2 | C50 19jun24 1.5m |  | 1 | 0 | -23.407 |  | -23.5559 |  |  | -23.7304 |  |  |  | -28.5947
115 | 23 | 3 | C50 19jun24 1.5m |  | 1 | 0 | -23.668 |  | -23.74 |  |  | -23.916 |  |  |  | -28.773
116 | 23 | 4 | C50 19jun24 1.5m |  | 1 | 0 | -23.725 | 0.30922 | -23.7707 | 0.11071 |  | -23.9482 |  |  |  | -28.804
117 | 24 | 1 | CP2 22may24 0.1m |  | 1 | 0 | -25.546 |  | -25.8439 |  |  | -26.023 |  |  |  | -30.7979
118 | 24 | 2 | CP2 22may24 0.1m |  | 1 | 0 | -25.68 |  | -25.7612 |  |  | -25.9418 |  |  |  | -30.7198
119 | 24 | 3 | CP2 22may24 0.1m |  | 1 | 0 | -25.249 |  | -25.2775 |  |  | -25.4597 |  |  |  | -30.2565
120 | 24 | 4 | CP2 22may24 0.1m |  | 1 | 0 | -25.378 | 0.18886 | -25.3973 | 0.27504 |  | -25.581 |  |  |  | -30.3731
121 | 25 | 1 | CC4 22may24 0.1m - DUP |  | 1 | 0 | -25.251 |  | -25.2302 |  |  | -25.4154 |  |  |  | -30.214
122 | 25 | 2 | CC4 22may24 0.1m - DUP |  | 1 | 0 | -25.645 |  | -25.6561 |  |  | -25.8428 |  |  |  | -30.6247
123 | 25 | 3 | CC4 22may24 0.1m - DUP |  | 1 | 0 | -25.54 |  | -25.543 |  |  | -25.7313 |  |  |  | -30.5175
124 | 25 | 4 | CC4 22may24 0.1m - DUP |  | 1 | 0 | -25.078 | 0.26055 | -25.0745 | 0.26988 |  | -25.2643 |  |  |  | -30.0687
125 | 26 | 1 | Hawaii |  | 1 | 0 | -5.593 |  | -2.4051 |  |  | -2.5964 |  |  |  | -8.2848
126 | 26 | 2 | Hawaii |  | 1 | 0 | -3.328 |  | -2.4246 |  |  | -2.6175 |  |  |  | -8.3051
127 | 26 | 3 | Hawaii |  | 1 | 0 | -3.356 |  | -2.9495 |  |  | -3.1439 |  |  |  | -8.811
128 | 26 | 4 | Hawaii |  | 1 | 0 | -2.39 | 1.3604 | -2.1244 | 0.34422 |  | -2.3203 |  |  |  | -8.0195
129 | 27 | 1 | Blacksburg | Drift | 1 | 0 | -38.3 |  | -44.1752 |  | -44.1752 | -44.3726 | -44.3726 |  |  | -48.4319
130 | 27 | 2 | Blacksburg | Drift | 1 | 0 | -42.089 |  | -43.7379 |  | -43.7379 | -43.9368 | -43.9368 |  |  | -48.0131
131 | 27 | 3 | Blacksburg | Drift | 1 | 0 | -42.67 |  | -43.4238 |  | -43.4238 | -43.6243 | -43.6243 |  |  | -47.7128
132 | 27 | 4 | Blacksburg | Drift | 1 | 0 | -43.174 | 2.2169 | -43.6514 | 0.31465 | -43.6514 | -43.8534 | -43.8534 |  |  | -47.933
Note on G2: Manually change first six injections to ignore "-1".
Note on G3: Manually change first six injections to ignore "-1".
Note on G4: Manually change first six injections to ignore "-1".
Note on G5: Manually change first six injections to ignore "-1".
Note on G6: Manually change first six injections to ignore "-1".
Note on G7: Manually change first six injections to ignore "-1".

## RESULTS
| Analyst: | DWH |  |  |  | Complete Sequence |  |  | Run Summary | 18O | 18O | 2H | 2H
| Analysis Date: | 2025-05-28 |  |  |  | Line | Inj Number | Identifier | Ignore | Ignore | d18O Final Value | d2H Final Value | Ignore
| Project: | CCR |  |  |  | 1 | 1 | Blacksburg | -1 | -1 |  |  | -1
[6 empty] | 2 | 2 | Blacksburg | 0 | -1 |  |  | -1
[6 empty] | 3 | 3 | Blacksburg | 0 | -1 |  |  | -1
Vial | Sample Name | d18O | 1 s.d. | d2H | 1 s.d. | 4 | 4 | Blacksburg | 0 | -1 |  |  | -1
|  | ‰ vs. VSMOW |  | ‰ vs. VSMOW |  | 5 | 5 | Blacksburg | 0 | -1 |  |  | -1
1 | Blacksburg | -7.6137 | 0.052417 | -48.1015 | 0.11481 | 6 | 6 | Blacksburg | 0 | -1 |  |  | -1
2 | Myrtle | -2.2632 | 0.050134 | -10.0702 | 0.10021 | 7 | 7 | Blacksburg | 0 | 0 | -7.6291 | -48.0428 | 0
3 | Homer | -14.5168 | 0.083329 | -110.2797 | 0.36313 | 8 | 8 | Blacksburg | 0 | 0 | -7.6217 | -48.2692 | 0
4 | Blacksburg | -7.6122 | 0.061186 | -48.0659 | 0.30739 | 9 | 9 | Blacksburg | 0 | 0 | -7.54 | -48.015 | 0
5 | Hawaii | -2.8635 | 0.040444 | -11.1403 | 0.15361 | 10 | 10 | Blacksburg | 0 | 0 | -7.6638 | -48.0789 | 0
6 | CP1 22may24 BOT | -6.1049 | 0.05019 | -35.3616 | 0.22547 | 11 | 1 | Myrtle | 0 | 0 | -2.2791 | -9.9968 | 0
7 | C50 16sep24 0.1m | -3.7378 | 0.013337 | -25.0878 | 0.12089 | 12 | 2 | Myrtle | 0 | 0 | -2.2382 | -10.1731 | 0
8 | CC2 19jun24 0.1m | -4.7253 | 0.032625 | -28.9013 | 0.16364 | 13 | 3 | Myrtle | 0 | 0 | -2.3174 | -10.2249 | 0
9 | Blacksburg | -7.6073 | 0.033212 | -47.9228 | 0.20433 | 14 | 4 | Myrtle | 0 | 0 | -2.2476 | -10.0377 | 0
10 | CP1 28oct24 0.1m | -6.0709 | 0.042326 | -34.5218 | 0.14136 | 15 | 5 | Myrtle | 0 | 0 | -2.2333 | -10.1285 | 0
11 | CC4 11jul24 9m | -5.3186 | 0.034863 | -32.383 | 0.16039 | 16 | 6 | Myrtle | 0 | 0 | -2.1655 | -10.0611 | 0
12 | CC4 15aug24 BOT | -5.1753 | 0.046484 | -32.0076 | 0.22863 | 17 | 7 | Myrtle | 0 | 0 | -2.2738 | -9.8725 | 0
13 | C50 16apr25 0.1m | -5.0562 | 0.053653 | -30.8525 | 0.13165 | 18 | 8 | Myrtle | 0 | 0 | -2.2988 | -10.0039 | 0
14 | C50 17feb25 1.5m | -5.0832 | 0.035357 | -30.954 | 0.33875 | 19 | 9 | Myrtle | 0 | 0 | -2.3412 | -10.1162 | 0
15 | C50 11jul24 9m | -5.1546 | 0.048025 | -31.6235 | 0.13433 | 20 | 10 | Myrtle | 0 | 0 | -2.2373 | -10.0879 | 0
16 | CC4 22may24 0.1m | -4.9364 | 0.037214 | -30.5412 | 0.10704 | 21 | 1 | Homer | 0 | 0 | -14.4755 | -109.9899 | 0
17 | C50 22may24 0.1m | -4.9955 | 0.033007 | -30.4705 | 0.049228 | 22 | 2 | Homer | 0 | 0 | -14.5448 | -110.4519 | 0
18 | Blacksburg | -7.5786 | 0.03686 | -48.1984 | 0.14182 | 23 | 3 | Homer | 0 | 0 | -14.5294 | -110.138 | 0
19 | CS2 22may24 0.1m | -6.7589 | 0.034009 | -38.9986 | 0.073301 | 24 | 4 | Homer | 0 | 0 | -14.5595 | -110.3262 | 0
20 | CCR raindeploy - 11jul24 R1 | -2.168 | 0.019741 | -6.6049 | 0.29599 | 25 | 5 | Homer | 0 | 0 | -14.5462 | -110.5922 | 0
21 | C50 16apr25 6m | -5.11 | 0.024264 | -31.0184 | 0.066074 | 26 | 6 | Homer | 0 | 0 | -14.5349 | -110.6393 | 0
22 | CC4 16sep24 1.5m | -3.7196 | 0.078393 | -24.6869 | 0.17829 | 27 | 7 | Homer | 0 | 0 | -14.3115 | -109.4381 | 0
23 | C50 19jun24 1.5m | -4.5715 | 0.0374 | -28.6957 | 0.10811 | 28 | 8 | Homer | 0 | 0 | -14.5333 | -110.2122 | 0
24 | CP2 22may24 0.1m | -4.8933 | 0.033779 | -30.5368 | 0.2627 | 29 | 9 | Homer | 0 | 0 | -14.6352 | -110.575 | 0
25 | CC4 22may24 0.1m - DUP | -4.8629 | 0.044015 | -30.3562 | 0.25884 | 30 | 10 | Homer | 0 | 0 | -14.4976 | -110.4343 | 0
26 | Hawaii | -1.8096 | 0.027541 | -8.3551 | 0.33058 | 31 | 1 | Blacksburg | 0 | 0 | -7.5274 | -47.6492 | 0
27 | Blacksburg | -7.6252 | 0.038952 | -48.0227 | 0.30092 | 32 | 2 | Blacksburg | 0 | 0 | -7.6827 | -48.2812 | 0
[6 empty] | 33 | 3 | Blacksburg | 0 | 0 | -7.5927 | -47.7462 | 0
[6 empty] | 34 | 4 | Blacksburg | 0 | 0 | -7.6205 | -48.1606 | 0
[6 empty] | 35 | 5 | Blacksburg | 0 | 0 | -7.6354 | -48.4446 | 0
| Quality Assurance Verification |  |  |  |  | 36 | 6 | Blacksburg | 0 | 0 | -7.4986 | -47.7284 | 0
| Metric |  | Reviewer 1 | Reviewer 2 |  | 37 | 7 | Blacksburg | 0 | 0 | -7.6039 | -47.8637 | 0
| d18O Cal Curve R^2 | 0.999+ | pass |  |  | 38 | 8 | Blacksburg | 0 | 0 | -7.6138 | -47.9982 | 0
| d18O QC checks | < 0.15 | DWH; see comment |  |  | 39 | 9 | Blacksburg | 0 | 0 | -7.6777 | -48.4253 | 0
| dD Cal Curve R^2 | 0.999+ | pass |  |  | 40 | 10 | Blacksburg | 0 | 0 | -7.6693 | -48.3614 | 0
| dD QC checks | < 1.5 | DWH; see comment |  |  | 41 | 1 | Hawaii | 0 | 0 | -2.8127 | -11.0721 | 0
| Initials |  | DWH |  |  | 42 | 2 | Hawaii | 0 | 0 | -2.8522 | -10.974 | 0
| Date |  | 2025-05-29 |  |  | 43 | 3 | Hawaii | 0 | 0 | -2.9059 | -11.3317 | 0
| Comments: QC checks look bad, but could be do to using older preprepped vials. Going to use same references from this run for 29may run but with new Hawaii vials |  |  |  |  | 44 | 4 | Hawaii | 0 | 0 | -2.8833 | -11.1835 | 0
[6 empty] | 45 | 1 | CP1 22may24 BOT | 0 | 0 | -6.091 | -35.5107 | 0
[6 empty] | 46 | 2 | CP1 22may24 BOT | 0 | 0 | -6.1789 | -35.1162 | 0
[6 empty] | 47 | 3 | CP1 22may24 BOT | 0 | 0 | -6.0691 | -35.5906 | 0
[6 empty] | 48 | 4 | CP1 22may24 BOT | 0 | 0 | -6.0804 | -35.2288 | 0
[6 empty] | 49 | 1 | C50 16sep24 0.1m | 0 | 0 | -3.7531 | -25.2229 | 0
[6 empty] | 50 | 2 | C50 16sep24 0.1m | 0 | 0 | -3.7381 | -24.9634 | 0
[6 empty] | 51 | 3 | C50 16sep24 0.1m | 0 | 0 | -3.7206 | -25.0116 | 0
[6 empty] | 52 | 4 | C50 16sep24 0.1m | 0 | 0 | -3.7393 | -25.1534 | 0
[6 empty] | 53 | 1 | CC2 19jun24 0.1m | 0 | 0 | -4.7405 | -28.8199 | 0
[6 empty] | 54 | 2 | CC2 19jun24 0.1m | 0 | 0 | -4.7076 | -28.8932 | 0
[6 empty] | 55 | 3 | CC2 19jun24 0.1m | 0 | 0 | -4.6903 | -28.7595 | 0
[6 empty] | 56 | 4 | CC2 19jun24 0.1m | 0 | 0 | -4.763 | -29.1326 | 0
[6 empty] | 57 | 1 | Blacksburg | 0 | 0 | -7.617 | -47.6197 | 0
[6 empty] | 58 | 2 | Blacksburg | 0 | 0 | -7.5695 | -47.9973 | 0
[6 empty] | 59 | 3 | Blacksburg | 0 | 0 | -7.5948 | -48.0662 | 0
[6 empty] | 60 | 4 | Blacksburg | 0 | 0 | -7.6477 | -48.0081 | 0
[6 empty] | 61 | 1 | CP1 28oct24 0.1m | 0 | 0 | -6.0603 | -34.3719 | 0
[6 empty] | 62 | 2 | CP1 28oct24 0.1m | 0 | 0 | -6.0962 | -34.6381 | 0
[6 empty] | 63 | 3 | CP1 28oct24 0.1m | 0 | 0 | -6.1109 | -34.6468 | 0
[6 empty] | 64 | 4 | CP1 28oct24 0.1m | 0 | 0 | -6.0159 | -34.4305 | 0
[6 empty] | 65 | 1 | CC4 11jul24 9m | 0 | 0 | -5.319 | -32.4899 | 0
[6 empty] | 66 | 2 | CC4 11jul24 9m | 0 | 0 | -5.3211 | -32.5382 | 0
[6 empty] | 67 | 3 | CC4 11jul24 9m | 0 | 0 | -5.3597 | -32.3124 | 0
[6 empty] | 68 | 4 | CC4 11jul24 9m | 0 | 0 | -5.2745 | -32.1915 | 0
[6 empty] | 69 | 1 | CC4 15aug24 BOT | 0 | 0 | -5.1358 | -32.2543 | 0
[6 empty] | 70 | 2 | CC4 15aug24 BOT | 0 | 0 | -5.1862 | -31.8829 | 0
[6 empty] | 71 | 3 | CC4 15aug24 BOT | 0 | 0 | -5.2365 | -32.1375 | 0
[6 empty] | 72 | 4 | CC4 15aug24 BOT | 0 | 0 | -5.1427 | -31.7556 | 0
[6 empty] | 73 | 1 | C50 16apr25 0.1m | 0 | 0 | -5.0393 | -30.7259 | 0
[6 empty] | 74 | 2 | C50 16apr25 0.1m | 0 | 0 | -4.9931 | -30.9021 | 0
[6 empty] | 75 | 3 | C50 16apr25 0.1m | 0 | 0 | -5.1203 | -31.0145 | 0
[6 empty] | 76 | 4 | C50 16apr25 0.1m | 0 | 0 | -5.0722 | -30.7674 | 0
[6 empty] | 77 | 1 | C50 17feb25 1.5m | 0 | 0 | -5.036 | -31.2523 | 0
[6 empty] | 78 | 2 | C50 17feb25 1.5m | 0 | 0 | -5.1169 | -30.868 | 0
[6 empty] | 79 | 3 | C50 17feb25 1.5m | 0 | 0 | -5.1021 | -30.5122 | 0
[6 empty] | 80 | 4 | C50 17feb25 1.5m | 0 | 0 | -5.0779 | -31.1835 | 0
[6 empty] | 81 | 1 | C50 11jul24 9m | 0 | 0 | -5.1374 | -31.7708 | 0
[6 empty] | 82 | 2 | C50 11jul24 9m | 0 | 0 | -5.1883 | -31.4701 | 0
[6 empty] | 83 | 3 | C50 11jul24 9m | 0 | 0 | -5.1981 | -31.6933 | 0
[6 empty] | 84 | 4 | C50 11jul24 9m | 0 | 0 | -5.0947 | -31.56 | 0
[6 empty] | 85 | 1 | CC4 22may24 0.1m | 0 | 0 | -4.8948 | -30.4602 | 0
[6 empty] | 86 | 2 | CC4 22may24 0.1m | 0 | 0 | -4.925 | -30.6988 | 0
[6 empty] | 87 | 3 | CC4 22may24 0.1m | 0 | 0 | -4.9419 | -30.5008 | 0
[6 empty] | 88 | 4 | CC4 22may24 0.1m | 0 | 0 | -4.9839 | -30.5047 | 0
[6 empty] | 89 | 1 | C50 22may24 0.1m | 0 | 0 | -5.0221 | -30.4491 | 0
[6 empty] | 90 | 2 | C50 22may24 0.1m | 0 | 0 | -4.9637 | -30.4418 | 0
[6 empty] | 91 | 3 | C50 22may24 0.1m | 0 | 0 | -5.0259 | -30.447 | 0
[6 empty] | 92 | 4 | C50 22may24 0.1m | 0 | 0 | -4.9705 | -30.5442 | 0
[6 empty] | 93 | 1 | Blacksburg | 0 | 0 | -7.6316 | -48.3622 | 0
[6 empty] | 94 | 2 | Blacksburg | 0 | 0 | -7.5706 | -48.035 | 0
[6 empty] | 95 | 3 | Blacksburg | 0 | 0 | -7.5464 | -48.2563 | 0
[6 empty] | 96 | 4 | Blacksburg | 0 | 0 | -7.5658 | -48.1399 | 0
[6 empty] | 97 | 1 | CS2 22may24 0.1m | 0 | 0 | -6.7101 | -38.9763 | 0
[6 empty] | 98 | 2 | CS2 22may24 0.1m | 0 | 0 | -6.7717 | -39.0321 | 0
[6 empty] | 99 | 3 | CS2 22may24 0.1m | 0 | 0 | -6.7886 | -39.078 | 0
[6 empty] | 100 | 4 | CS2 22may24 0.1m | 0 | 0 | -6.7651 | -38.9081 | 0
[6 empty] | 101 | 1 | CCR raindeploy - 11jul24 R1 | 0 | 0 | -2.1815 | -6.1626 | 0
[6 empty] | 102 | 2 | CCR raindeploy - 11jul24 R1 | 0 | 0 | -2.1879 | -6.7169 | 0
[6 empty] | 103 | 3 | CCR raindeploy - 11jul24 R1 | 0 | 0 | -2.1558 | -6.7594 | 0
[6 empty] | 104 | 4 | CCR raindeploy - 11jul24 R1 | 0 | 0 | -2.1469 | -6.7805 | 0
[6 empty] | 105 | 1 | C50 16apr25 6m | 0 | 0 | -5.1385 | -31.0974 | 0
[6 empty] | 106 | 2 | C50 16apr25 6m | 0 | 0 | -5.0831 | -31.0294 | 0
[6 empty] | 107 | 3 | C50 16apr25 6m | 0 | 0 | -5.0986 | -31.01 | 0
[6 empty] | 108 | 4 | C50 16apr25 6m | 0 | 0 | -5.1199 | -30.9368 | 0
[6 empty] | 109 | 1 | CC4 16sep24 1.5m | 0 | 0 | -3.8021 | -24.5901 | 0
[6 empty] | 110 | 2 | CC4 16sep24 1.5m | 0 | 0 | -3.7531 | -24.534 | 0
[6 empty] | 111 | 3 | CC4 16sep24 1.5m | 0 | 0 | -3.705 | -24.6865 | 0
[6 empty] | 112 | 4 | CC4 16sep24 1.5m | 0 | 0 | -3.6181 | -24.9371 | 0
[6 empty] | 113 | 1 | C50 19jun24 1.5m | 0 | 0 | -4.6241 | -28.6112 | 0
[6 empty] | 114 | 2 | C50 19jun24 1.5m | 0 | 0 | -4.5568 | -28.5947 | 0
[6 empty] | 115 | 3 | C50 19jun24 1.5m | 0 | 0 | -4.5369 | -28.773 | 0
[6 empty] | 116 | 4 | C50 19jun24 1.5m | 0 | 0 | -4.5679 | -28.804 | 0
[6 empty] | 117 | 1 | CP2 22may24 0.1m | 0 | 0 | -4.8739 | -30.7979 | 0
[6 empty] | 118 | 2 | CP2 22may24 0.1m | 0 | 0 | -4.9241 | -30.7198 | 0
[6 empty] | 119 | 3 | CP2 22may24 0.1m | 0 | 0 | -4.8558 | -30.2565 | 0
[6 empty] | 120 | 4 | CP2 22may24 0.1m | 0 | 0 | -4.9195 | -30.3731 | 0
[6 empty] | 121 | 1 | CC4 22may24 0.1m - DUP | 0 | 0 | -4.8809 | -30.214 | 0
[6 empty] | 122 | 2 | CC4 22may24 0.1m - DUP | 0 | 0 | -4.7975 | -30.6247 | 0
[6 empty] | 123 | 3 | CC4 22may24 0.1m - DUP | 0 | 0 | -4.8794 | -30.5175 | 0
[6 empty] | 124 | 4 | CC4 22may24 0.1m - DUP | 0 | 0 | -4.8936 | -30.0687 | 0
[6 empty] | 125 | 1 | Hawaii | 0 | 0 | -1.7919 | -8.2848 | 0
[6 empty] | 126 | 2 | Hawaii | 0 | 0 | -1.7914 | -8.3051 | 0
[6 empty] | 127 | 3 | Hawaii | 0 | 0 | -1.8053 | -8.811 | 0
[6 empty] | 128 | 4 | Hawaii | 0 | 0 | -1.8498 | -8.0195 | 0
[6 empty] | 129 | 1 | Blacksburg | 0 | 0 | -7.6269 | -48.4319 | 0
[6 empty] | 130 | 2 | Blacksburg | 0 | 0 | -7.6376 | -48.0131 | 0
[6 empty] | 131 | 3 | Blacksburg | 0 | 0 | -7.5719 | -47.7128 | 0
[6 empty] | 132 | 4 | Blacksburg | 0 | 0 | -7.6646 | -47.933 | 0