## Instructions
Picarro Isotope Analyzer
Data Processor 3.0 (KAP)
How to Use this Spreadsheet:
Step 1: Retrieve raw result file from Piccaro CPU (C:Piccaro\IsotopeData). Add initials and project name to end
of file name. Copy file into your folder on Water Server.
Step 2: On syncing computer, move raw file into Raw Data folder on Google Drive:
Instruments and Analyses --> Picarro Water Isotope Analyzer --> Raw Data
Step 3: Save this spreadsheet in the Processed Data folder on Google Drive:
Instruments and Analyses --> Picarro Water Isotope Analyzer --> Processed Data
Use file name: YYYYMMDD_Initials_(anything else to identify ex. project, run number, etc.)
Step 4: Enter run information and sample positions in the "Tray Configuration" tab. (orange cells)
(It is important to use the same names for the standards throughout this spreadsheet)
Step 5: Copy text from raw result file and paste into "Picarro Output" tab.
Step 6: In Run Summary tab, check that H2O Mean is within "good" bounds. (between 17,000 - 23,000)
Step 7: In 18O tab, verify that the reference defined values match the references used in run (cells U39 - 42). Run Solver and
verify stdev memory corrected are reduced (column K values less than corresponding column I values). Verify stdev of
drift corrected reference is lower than memory corrected. Verify normalization R2 is 0.99+. Verify QC checks are within
0.15 tolerance.
Step 8: In dD tab, verify that the reference defined values match the references used in run (cells U39 - 42. Run Solver and
verify stdev memory corrected are reduced (column K values less than corresponding column I values). Verify stdev of
drift corrected reference is lower than memory corrected. Verify normalization R2 is 0.99+. Verify QC checks are within
1.5 tolerance.
Step 9: Fill in Quality Assurance Verification chart on RESULTS tab. Dr. Scott or Kelly Peeler will fill in as reviewer 2.
Step 10: Add reference results to Water Isotope Analyzer run log.

## Tray Configuration
| Analyst |  | DWH
| PI |  | Durelle Scott
| Project |  | CCR
| Sampling Date |  | 2024 - 2025
| Sampling Location |  | Carvins Cove Reservoir
| Analysis Date |  | 2025-05-29
Notes
Rear Rack
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
1 | Blacksburg | Myrtle | Homer | Blacksburg | Hawaii | C50 22may24 1.5m | CC3 19jun24 0.1m | C50 16sep24 6m | Blacksburg | 9
10 | CC5 22may24 1.5m | CS2 19jun24 0.1m | CC4 19jun24 9m | CC5 22may24 6m | C50 16apr25 1.5m | CC5 22may24 0.1m | CC5 22may24 9m | CC4 22may24 BOT | Blacksburg | 18
19 | CC3 16apr25 1.5m | CP2 22may24 BOT | CC4 16apr25 BOT | CP1 22may24 0.1m | CP2 19jun24 0.1m | CS1 19jun24 0.1m | CC5 22may24 9m DUP | Hawaii | Blacksburg | 27
28 | [9 empty] | 36
37 | [9 empty] | 45
46 | [9 empty] | 54
Rear Rack
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
1 | Drift | High reference | Low Reference | Drift | QC Check | Sample | Sample | Sample | Drift | 9
10 | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Drift | 18
19 | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Drift | 27
28 | [9 empty] | 36
37 | [9 empty] | 45
46 | [9 empty] | 54
Standards
|  | Standard ID | Actual d(18_16) | Actual d(D_H)
| #1 | Denali | -19.9287 | -153.7172 | [6 empty] | -3.37
| #2 | Alberta | -18.8392 | -146.8245 | [6 empty] | -20.92
| #3 | MT BHB | -18.431 | -140.5778 | [6 empty] | -17.35
| #4 | Anchorage | -16.4204 | -129.9028 | [6 empty] | -7.39
| #5 | Homer | -14.5399 | -110.5169
| #6 | MT Creek | -13.3562 | -97.7428
| #7 | Nickel Mines | -7.876 | -48.7951
| #8 | Blacksburg | -7.5615 | -47.446
| #9 | Denton's Well | -6.9583 | -42.6011
| #10 | Hawaii | -3.6108 | -13.2294
| #11 | Myrtle Beach | -2.2908 | -10.453
| #12 | Heavy Oxygen | 3.1753 | -15.7881
| #13 | Heavy Hydrogen | 11.1372 | 4.0332
| #14 | Andres Heavy Water | 17.5664 | 20.5382
| #15 | Picarro ZERO | 0.3 | 1.8
| #16 | Picarro MID | -20.6 | -159
| #17 | Picarro DEPL | -29.6 | -235
#18
#19
#20
|  | DI | -6 | -36
Note on C34: Nate: The Standard ID must match the Standard ID used above in the trey protocol. For example, if you are using the Boulder Colorado Standard, and you labeled it "B" above, it must be labled "B" in this collum.

## Timing
| Enter start time and date --> | 2019-01-10 09:30
Tray Position | Sample ID | Start Time | # Injections
1 | Blacksburg | 2019-01-10 10:58 | 10
2 | Myrtle | 2019-01-10 12:27 | 10
3 | Homer | 2019-01-10 13:55 | 10
4 | Blacksburg | 2019-01-10 15:23 | 10
5 | Hawaii | 2019-01-10 15:59 | 4
6 | C50 22may24 1.5m | 2019-01-10 16:34 | 4
7 | CC3 19jun24 0.1m | 2019-01-10 17:09 | 4
8 | C50 16sep24 6m | 2019-01-10 17:45 | 4
9 | Blacksburg | 2019-01-10 18:20 | 4
10 | CC5 22may24 1.5m | 2019-01-10 18:55 | 4
11 | CS2 19jun24 0.1m | 2019-01-10 19:31 | 4
12 | CC4 19jun24 9m | 2019-01-10 20:06 | 4
13 | CC5 22may24 6m | 2019-01-10 20:41 | 4
14 | C50 16apr25 1.5m | 2019-01-10 21:17 | 4
15 | CC5 22may24 0.1m | 2019-01-10 21:52 | 4
16 | CC5 22may24 9m | 2019-01-10 22:27 | 4
17 | CC4 22may24 BOT | 2019-01-10 23:03 | 4
18 | Blacksburg | 2019-01-10 23:38 | 4
19 | CC3 16apr25 1.5m | 2019-01-11 00:13 | 4
20 | CP2 22may24 BOT | 2019-01-11 00:49 | 4
21 | CC4 16apr25 BOT | 2019-01-11 01:24 | 4
22 | CP1 22may24 0.1m | 2019-01-11 01:59 | 4
23 | CP2 19jun24 0.1m | 2019-01-11 02:35 | 4
24 | CS1 19jun24 0.1m | 2019-01-11 03:10 | 4
25 | CC5 22may24 9m DUP | 2019-01-11 03:45 | 4
26 | Hawaii | 2019-01-11 04:21 | 4
27 | Blacksburg | 2019-01-11 04:56 | 4
Formulas:
Start Time = [Start Time, row above]+"0:08:50"*10

## Picarro Output
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | days_Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | BASELINE_Mean | BASELINE_SD | DAS Temp | Tray | Sample | Job | Method
1 | P-11755 | 2025/05/29 10:48:10 | MT1-Rear-01 | 1 | -9.037 | -49.993 | 20064 | -1 | 1 |  |  | H2O | 148.427 | 0.12 | 0.299 | 106 | -0.029 | -0.037 | 31.92 | 845 | 0 | 30.75 | MT1-Rear | 1 | 66 | BSE
2 | P-11755 | 2025/05/29 10:56:57 | MT1-Rear-01 | 2 | -8.856 | -45.757 | 19486 | -1 | 1 |  |  | H2O | 148.433 | 0.136 | 0.344 | 92 | -0.033 | -0.072 | 30.34 | 844 | 0 | 30.75 | MT1-Rear | 1 | 66 | BSE
3 | P-11755 | 2025/05/29 11:05:43 | MT1-Rear-01 | 3 | -8.719 | -44.86 | 19391 | -1 | 1 |  |  | H2O | 148.4391 | 0.148 | 0.402 | 99 | -0.019 | -0.041 | 29.86 | 844 | 0 | 30.563 | MT1-Rear | 1 | 66 | BSE
4 | P-11755 | 2025/05/29 11:14:29 | MT1-Rear-01 | 4 | -8.813 | -44.233 | 19438 | 0 | 1 |  |  | H2O | 148.4452 | 0.169 | 0.256 | 101 | -0.021 | -0.009 | 30.43 | 844 | 0 | 30.688 | MT1-Rear | 1 | 66 | BSE
5 | P-11755 | 2025/05/29 11:23:16 | MT1-Rear-01 | 5 | -8.793 | -44.072 | 19461 | 0 | 1 |  |  | H2O | 148.4514 | 0.151 | 0.275 | 101 | -0.038 | 0.03 | 30.56 | 844 | 0 | 30.875 | MT1-Rear | 1 | 66 | BSE
6 | P-11755 | 2025/05/29 11:32:02 | MT1-Rear-01 | 6 | -8.726 | -44.093 | 19726 | 0 | 1 |  |  | H2O | 148.4574 | 0.112 | 0.305 | 99 | 0.003 | 0.03 | 29.78 | 844 | 0 | 30.875 | MT1-Rear | 1 | 66 | BSE
7 | P-11755 | 2025/05/29 11:40:49 | MT1-Rear-01 | 7 | -8.84 | -44.062 | 19324 | 0 | 1 |  |  | H2O | 148.4635 | 0.134 | 0.371 | 99 | -0.007 | -0.052 | 29.76 | 844 | 0 | 30.813 | MT1-Rear | 1 | 66 | BSE
8 | P-11755 | 2025/05/29 11:49:36 | MT1-Rear-01 | 8 | -8.853 | -43.927 | 19254 | 0 | 1 |  |  | H2O | 148.4696 | 0.134 | 0.272 | 101 | -0.025 | -0.025 | 30.43 | 844 | 0 | 30.75 | MT1-Rear | 1 | 66 | BSE
9 | P-11755 | 2025/05/29 11:58:22 | MT1-Rear-01 | 9 | -8.816 | -43.805 | 19179 | 0 | 1 |  |  | H2O | 148.4757 | 0.132 | 0.319 | 91 | -0.021 | -0.017 | 29.98 | 844 | 0 | 30.625 | MT1-Rear | 1 | 66 | BSE
10 | P-11755 | 2025/05/29 12:07:08 | MT1-Rear-01 | 10 | -8.689 | -43.968 | 19081 | 0 | 1 |  |  | H2O | 148.4818 | 0.129 | 0.234 | 99 | -0.024 | -0.021 | 29.81 | 844 | 0 | 31 | MT1-Rear | 1 | 66 | BSE
11 | P-11756 | 2025/05/29 12:15:55 | MT1-Rear-02 | 1 | -3.619 | -9.741 | 19659 | -1 | 1 |  |  | H2O | 148.4879 | 0.103 | 0.371 | 101 | -0.008 | -0.058 | 30.42 | 844 | 0 | 30.813 | MT1-Rear | 2 | 66 | BSE
12 | P-11756 | 2025/05/29 12:24:41 | MT1-Rear-02 | 2 | -3.344 | -5.886 | 19607 | -1 | 1 |  |  | H2O | 148.494 | 0.148 | 0.417 | 101 | -0.032 | -0.092 | 30.47 | 844 | 0 | 30.875 | MT1-Rear | 2 | 66 | BSE
13 | P-11756 | 2025/05/29 12:33:28 | MT1-Rear-02 | 3 | -3.265 | -4.837 | 19472 | -1 | 1 |  |  | H2O | 148.5 | 0.128 | 0.322 | 97 | -0.024 | -0.011 | 29.26 | 844 | 0 | 30.688 | MT1-Rear | 2 | 66 | BSE
14 | P-11756 | 2025/05/29 12:42:14 | MT1-Rear-02 | 4 | -3.174 | -4.663 | 19577 | 0 | 1 |  |  | H2O | 148.5062 | 0.124 | 0.347 | 99 | -0.01 | -0.035 | 29.86 | 844 | 0 | 30.688 | MT1-Rear | 2 | 66 | BSE
15 | P-11756 | 2025/05/29 12:51:00 | MT1-Rear-02 | 5 | -3.193 | -4.292 | 19751 | 0 | 1 |  |  | H2O | 148.5122 | 0.129 | 0.449 | 88 | -0.021 | -0.014 | 29.04 | 844 | 0 | 30.75 | MT1-Rear | 2 | 66 | BSE
16 | P-11756 | 2025/05/29 12:59:46 | MT1-Rear-02 | 6 | -3.125 | -4.4 | 19509 | 0 | 1 |  |  | H2O | 148.5183 | 0.101 | 0.293 | 99 | -0.018 | -0.055 | 29.9 | 844 | 0 | 30.688 | MT1-Rear | 2 | 66 | BSE
17 | P-11756 | 2025/05/29 13:08:33 | MT1-Rear-02 | 7 | -3.159 | -4.4 | 19533 | 0 | 1 |  |  | H2O | 148.5245 | 0.126 | 0.335 | 99 | -0.02 | -0.044 | 29.71 | 844 | 0 | 30.938 | MT1-Rear | 2 | 66 | BSE
18 | P-11756 | 2025/05/29 13:17:19 | MT1-Rear-02 | 8 | -3.177 | -4.353 | 19435 | 0 | 1 |  |  | H2O | 148.5305 | 0.11 | 0.422 | 92 | -0.018 | 0.004 | 30.35 | 844 | 0 | 30.813 | MT1-Rear | 2 | 66 | BSE
19 | P-11756 | 2025/05/29 13:26:07 | MT1-Rear-02 | 9 | -3.159 | -4.001 | 19488 | 0 | 1 |  |  | H2O | 148.5367 | 0.121 | 0.212 | 93 | -0.016 | -0.005 | 30.69 | 844 | 0 | 30.875 | MT1-Rear | 2 | 66 | BSE
20 | P-11756 | 2025/05/29 13:34:53 | MT1-Rear-02 | 10 | -3.029 | -4.242 | 19468 | 0 | 1 |  |  | H2O | 148.5427 | 0.181 | 0.267 | 97 | -0.017 | -0.044 | 29.34 | 844 | 0 | 30.688 | MT1-Rear | 2 | 66 | BSE
21 | P-11757 | 2025/05/29 13:43:40 | MT1-Rear-03 | 1 | -14.828 | -94.703 | 19343 | -1 | 1 |  |  | H2O | 148.5488 | 0.074 | 0.461 | 99 | -0.008 | 0.103 | 29.91 | 844 | 0 | 31 | MT1-Rear | 3 | 66 | BSE
22 | P-11757 | 2025/05/29 13:52:28 | MT1-Rear-03 | 2 | -15.563 | -105.109 | 19615 | -1 | 1 |  |  | H2O | 148.555 | 0.104 | 0.438 | 101 | 0.005 | 0.093 | 30.37 | 844 | 0 | 30.813 | MT1-Rear | 3 | 66 | BSE
23 | P-11757 | 2025/05/29 14:01:15 | MT1-Rear-03 | 3 | -15.627 | -107.067 | 19512 | -1 | 1 |  |  | H2O | 148.561 | 0.075 | 0.413 | 89 | -0.009 | 0.112 | 29.53 | 844 | 0 | 30.938 | MT1-Rear | 3 | 66 | BSE
24 | P-11757 | 2025/05/29 14:10:02 | MT1-Rear-03 | 4 | -15.776 | -107.801 | 19476 | 0 | 1 |  |  | H2O | 148.5671 | 0.145 | 0.299 | 96 | -0.026 | 0.069 | 28.79 | 844 | 0 | 30.938 | MT1-Rear | 3 | 66 | BSE
25 | P-11757 | 2025/05/29 14:18:48 | MT1-Rear-03 | 5 | -15.792 | -108.31 | 19495 | 0 | 1 |  |  | H2O | 148.5732 | 0.119 | 0.368 | 99 | -0.024 | 0.05 | 29.79 | 844 | 0 | 30.563 | MT1-Rear | 3 | 66 | BSE
26 | P-11757 | 2025/05/29 14:27:34 | MT1-Rear-03 | 6 | -15.846 | -108.557 | 19516 | 0 | 1 |  |  | H2O | 148.5794 | 0.104 | 0.477 | 100 | 0.001 | 0.088 | 30.04 | 844 | 0 | 30.563 | MT1-Rear | 3 | 66 | BSE
27 | P-11757 | 2025/05/29 14:36:21 | MT1-Rear-03 | 7 | -15.782 | -108.687 | 19505 | 0 | 1 |  |  | H2O | 148.5854 | 0.132 | 0.324 | 92 | -0.025 | 0.083 | 30.29 | 844 | 0 | 30.813 | MT1-Rear | 3 | 66 | BSE
28 | P-11757 | 2025/05/29 14:45:10 | MT1-Rear-03 | 8 | -15.789 | -109.189 | 19877 | 0 | 1 |  |  | H2O | 148.5915 | 0.092 | 0.346 | 100 | -0.018 | 0.052 | 30 | 844 | 0 | 31.063 | MT1-Rear | 3 | 66 | BSE
29 | P-11757 | 2025/05/29 14:53:57 | MT1-Rear-03 | 9 | -15.81 | -108.907 | 19506 | 0 | 1 |  |  | H2O | 148.5976 | 0.098 | 0.274 | 100 | -0.02 | 0.02 | 30.13 | 844 | 0 | 30.813 | MT1-Rear | 3 | 66 | BSE
30 | P-11757 | 2025/05/29 15:02:44 | MT1-Rear-03 | 10 | -15.847 | -109.21 | 19573 | 0 | 1 |  |  | H2O | 148.6038 | 0.092 | 0.306 | 99 | -0.009 | 0.061 | 29.9 | 844 | 0 | 30.875 | MT1-Rear | 3 | 66 | BSE
31 | P-11758 | 2025/05/29 15:11:31 | MT1-Rear-04 | 1 | -9.211 | -52.491 | 19712 | -1 | 1 |  |  | H2O | 148.6098 | 0.154 | 0.303 | 89 | -0.011 | -0.057 | 29.51 | 844 | 0 | 30.75 | MT1-Rear | 4 | 66 | BSE
32 | P-11758 | 2025/05/29 15:20:18 | MT1-Rear-04 | 2 | -8.755 | -46.633 | 19602 | -1 | 1 |  |  | H2O | 148.6159 | 0.119 | 0.365 | 97 | -0.017 | -0.057 | 29.3 | 844 | 0 | 30.75 | MT1-Rear | 4 | 66 | BSE
33 | P-11758 | 2025/05/29 15:29:05 | MT1-Rear-04 | 3 | -8.788 | -45.174 | 19641 | -1 | 1 |  |  | H2O | 148.622 | 0.143 | 0.296 | 101 | -0.026 | 0.042 | 30.41 | 844 | 0 | 30.938 | MT1-Rear | 4 | 66 | BSE
34 | P-11758 | 2025/05/29 15:37:53 | MT1-Rear-04 | 4 | -8.656 | -44.927 | 19606 | 0 | 1 |  |  | H2O | 148.6282 | 0.12 | 0.273 | 99 | -0.014 | -0.035 | 29.86 | 844 | 0 | 30.875 | MT1-Rear | 4 | 66 | BSE
35 | P-11758 | 2025/05/29 15:46:40 | MT1-Rear-04 | 5 | -8.649 | -44.37 | 20029 | 0 | 1 |  |  | H2O | 148.6343 | 0.097 | 0.272 | 89 | -0.019 | -0.078 | 29.55 | 844 | 0 | 31.188 | MT1-Rear | 4 | 66 | BSE
36 | P-11758 | 2025/05/29 15:55:26 | MT1-Rear-04 | 6 | -8.648 | -44.474 | 19636 | 0 | 1 |  |  | H2O | 148.6403 | 0.119 | 0.322 | 98 | -0.016 | -0.067 | 29.42 | 844 | 0 | 30.75 | MT1-Rear | 4 | 66 | BSE
37 | P-11758 | 2025/05/29 16:04:13 | MT1-Rear-04 | 7 | -8.666 | -44.201 | 19630 | 0 | 1 |  |  | H2O | 148.6464 | 0.141 | 0.318 | 98 | -0.004 | 0.064 | 29.6 | 844 | 0 | 31 | MT1-Rear | 4 | 66 | BSE
38 | P-11758 | 2025/05/29 16:13:02 | MT1-Rear-04 | 8 | -8.602 | -44.398 | 19653 | 0 | 1 |  |  | H2O | 148.6526 | 0.109 | 0.233 | 98 | -0.008 | -0.003 | 29.56 | 844 | 0 | 30.938 | MT1-Rear | 4 | 66 | BSE
39 | P-11758 | 2025/05/29 16:21:49 | MT1-Rear-04 | 9 | -8.666 | -44.137 | 19506 | 0 | 1 |  |  | H2O | 148.6587 | 0.148 | 0.279 | 99 | -0.033 | 0.027 | 29.73 | 844 | 0 | 30.938 | MT1-Rear | 4 | 66 | BSE
40 | P-11758 | 2025/05/29 16:30:37 | MT1-Rear-04 | 10 | -8.616 | -44.072 | 20084 | 0 | 1 |  |  | H2O | 148.6648 | 0.094 | 0.29 | 88 | -0.016 | -0.025 | 29.09 | 844 | 0 | 30.75 | MT1-Rear | 4 | 66 | BSE
41 | P-11759 | 2025/05/29 16:39:25 | MT1-Rear-05 | 1 | -4.876 | -12.658 | 19765 | -1 | 1 |  |  | H2O | 148.6708 | 0.203 | 0.379 | 97 | -0.046 | -0.001 | 29.13 | 844 | 0 | 30.875 | MT1-Rear | 5 | 67 | BSE
42 | P-11759 | 2025/05/29 16:48:13 | MT1-Rear-05 | 2 | -4.618 | -9.388 | 19396 | -1 | 1 |  |  | H2O | 148.677 | 0.201 | 0.289 | 96 | -0.043 | -0.038 | 29.03 | 844 | 0 | 30.75 | MT1-Rear | 5 | 67 | BSE
43 | P-11759 | 2025/05/29 16:57:01 | MT1-Rear-05 | 3 | -4.675 | -8.253 | 19480 | -1 | 1 |  |  | H2O | 148.6831 | 0.15 | 0.269 | 97 | -0.017 | -0.034 | 29.26 | 844 | 0 | 30.75 | MT1-Rear | 5 | 67 | BSE
44 | P-11759 | 2025/05/29 17:05:50 | MT1-Rear-05 | 4 | -4.567 | -8.06 | 19884 | 0 | 1 |  |  | H2O | 148.6892 | 0.142 | 0.382 | 97 | -0.026 | -0.083 | 29.24 | 844 | 0 | 30.938 | MT1-Rear | 5 | 67 | BSE
45 | P-11760 | 2025/05/29 17:14:38 | MT1-Rear-06 | 1 | -6.029 | -23.486 | 19392 | -1 | 1 |  |  | H2O | 148.6954 | 0.152 | 0.356 | 99 | -0.028 | 0.013 | 29.71 | 844 | 0 | 30.75 | MT1-Rear | 6 | 67 | BSE
46 | P-11760 | 2025/05/29 17:23:27 | MT1-Rear-06 | 2 | -6.013 | -25.544 | 19727 | -1 | 1 |  |  | H2O | 148.7015 | 0.132 | 0.259 | 100 | -0.023 | -0.036 | 30.11 | 844 | 0 | 30.813 | MT1-Rear | 6 | 67 | BSE
47 | P-11760 | 2025/05/29 17:32:15 | MT1-Rear-06 | 3 | -5.995 | -25.756 | 19593 | -1 | 1 |  |  | H2O | 148.7076 | 0.124 | 0.445 | 99 | -0.007 | -0.029 | 29.71 | 844 | 0 | 31.125 | MT1-Rear | 6 | 67 | BSE
48 | P-11760 | 2025/05/29 17:41:03 | MT1-Rear-06 | 4 | -5.988 | -25.881 | 20102 | 0 | 1 |  |  | H2O | 148.7138 | 0.1 | 0.207 | 101 | -0.016 | -0.001 | 30.35 | 844 | 0 | 30.813 | MT1-Rear | 6 | 67 | BSE
49 | P-11761 | 2025/05/29 17:49:52 | MT1-Rear-07 | 1 | -5.636 | -24.078 | 20110 | -1 | 1 |  |  | H2O | 148.7198 | 0.171 | 0.368 | 97 | -0.023 | 0.017 | 29.34 | 844 | 0 | 30.813 | MT1-Rear | 7 | 67 | BSE
50 | P-11761 | 2025/05/29 17:58:40 | MT1-Rear-07 | 2 | -5.61 | -23.83 | 19621 | -1 | 1 |  |  | H2O | 148.7259 | 0.183 | 0.202 | 99 | -0.038 | -0.019 | 29.84 | 844 | 0 | 30.813 | MT1-Rear | 7 | 67 | BSE
51 | P-11761 | 2025/05/29 18:07:28 | MT1-Rear-07 | 3 | -5.613 | -23.905 | 19479 | -1 | 1 |  |  | H2O | 148.732 | 0.148 | 0.316 | 96 | -0.024 | -0.011 | 28.85 | 844 | 0 | 30.813 | MT1-Rear | 7 | 67 | BSE
52 | P-11761 | 2025/05/29 18:16:16 | MT1-Rear-07 | 4 | -5.678 | -23.849 | 19820 | 0 | 1 |  |  | H2O | 148.7381 | 0.136 | 0.277 | 99 | -0.016 | -0.004 | 29.94 | 844 | 0 | 30.875 | MT1-Rear | 7 | 67 | BSE
53 | P-11762 | 2025/05/29 18:25:04 | MT1-Rear-08 | 1 | -4.714 | -20.541 | 20263 | -1 | 1 |  |  | H2O | 148.7443 | 0.121 | 0.409 | 97 | -0.017 | -0.023 | 29.16 | 844 | 0 | 30.938 | MT1-Rear | 8 | 67 | BSE
54 | P-11762 | 2025/05/29 18:33:52 | MT1-Rear-08 | 2 | -4.687 | -20.249 | 20060 | -1 | 1 |  |  | H2O | 148.7504 | 0.087 | 0.309 | 101 | 0.001 | 0 | 30.51 | 844 | 0 | 30.938 | MT1-Rear | 8 | 67 | BSE
55 | P-11762 | 2025/05/29 18:42:40 | MT1-Rear-08 | 3 | -4.688 | -19.427 | 19780 | -1 | 1 |  |  | H2O | 148.7565 | 0.179 | 0.271 | 100 | -0.043 | -0.028 | 30.06 | 844 | 0 | 30.625 | MT1-Rear | 8 | 67 | BSE
56 | P-11762 | 2025/05/29 18:51:27 | MT1-Rear-08 | 4 | -4.65 | -19.78 | 20073 | 0 | 1 |  |  | H2O | 148.7625 | 0.146 | 0.149 | 96 | -0.035 | -0.014 | 29.02 | 844 | 0 | 30.813 | MT1-Rear | 8 | 67 | BSE
57 | P-11763 | 2025/05/29 19:00:16 | MT1-Rear-09 | 1 | -8.315 | -40.483 | 19702 | -1 | 1 |  |  | H2O | 148.7687 | 0.152 | 0.232 | 99 | -0.026 | 0.026 | 29.68 | 844 | 0 | 30.688 | MT1-Rear | 9 | 67 | BSE
58 | P-11763 | 2025/05/29 19:09:05 | MT1-Rear-09 | 2 | -8.536 | -42.795 | 20124 | -1 | 1 |  |  | H2O | 148.7748 | 0.122 | 0.288 | 97 | -0.012 | 0.019 | 29.22 | 844 | 0 | 30.75 | MT1-Rear | 9 | 67 | BSE
59 | P-11763 | 2025/05/29 19:17:53 | MT1-Rear-09 | 3 | -8.647 | -43.076 | 19740 | -1 | 1 |  |  | H2O | 148.7809 | 0.113 | 0.272 | 99 | 0.004 | -0.021 | 29.79 | 844 | 0 | 30.688 | MT1-Rear | 9 | 67 | BSE
60 | P-11763 | 2025/05/29 19:26:42 | MT1-Rear-09 | 4 | -8.611 | -43.265 | 19934 | 0 | 1 |  |  | H2O | 148.7871 | 0.102 | 0.42 | 99 | -0.011 | 0.05 | 29.68 | 844 | 0 | 30.75 | MT1-Rear | 9 | 67 | BSE
61 | P-11764 | 2025/05/29 19:35:30 | MT1-Rear-10 | 1 | -6.184 | -28.182 | 19727 | -1 | 1 |  |  | H2O | 148.7932 | 0.125 | 0.373 | 98 | -0.012 | -0.004 | 29.43 | 844 | 0 | 30.688 | MT1-Rear | 10 | 67 | BSE
62 | P-11764 | 2025/05/29 19:44:18 | MT1-Rear-10 | 2 | -6.038 | -26.671 | 20162 | -1 | 1 |  |  | H2O | 148.7993 | 0.15 | 0.199 | 100 | -0.028 | 0.015 | 30.09 | 844 | 0 | 30.688 | MT1-Rear | 10 | 67 | BSE
63 | P-11764 | 2025/05/29 19:53:08 | MT1-Rear-10 | 3 | -5.924 | -25.878 | 20150 | -1 | 1 |  |  | H2O | 148.8054 | 0.203 | 0.284 | 99 | -0.048 | -0.032 | 29.94 | 844 | 0 | 30.563 | MT1-Rear | 10 | 67 | BSE
64 | P-11764 | 2025/05/29 20:01:57 | MT1-Rear-10 | 4 | -6.069 | -25.92 | 19702 | 0 | 1 |  |  | H2O | 148.8116 | 0.108 | 0.229 | 101 | 0 | -0.032 | 30.31 | 844 | 0 | 29.813 | MT1-Rear | 10 | 67 | BSE
65 | P-11765 | 2025/05/29 20:10:46 | MT1-Rear-11 | 1 | -7.348 | -31.022 | 20002 | -1 | 1 |  |  | H2O | 148.8177 | 0.13 | 0.412 | 100 | -0.016 | -0.03 | 30.08 | 844 | 0 | 30.688 | MT1-Rear | 11 | 67 | BSE
66 | P-11765 | 2025/05/29 20:19:35 | MT1-Rear-11 | 2 | -7.361 | -31.66 | 20098 | -1 | 1 |  |  | H2O | 148.8238 | 0.166 | 0.286 | 98 | -0.033 | 0.001 | 29.54 | 844 | 0 | 30.5 | MT1-Rear | 11 | 67 | BSE
67 | P-11765 | 2025/05/29 20:28:25 | MT1-Rear-11 | 3 | -7.439 | -31.797 | 19819 | -1 | 1 |  |  | H2O | 148.83 | 0.149 | 0.265 | 100 | -0.028 | 0.011 | 30.19 | 844 | 0 | 30.313 | MT1-Rear | 11 | 67 | BSE
68 | P-11765 | 2025/05/29 20:37:14 | MT1-Rear-11 | 4 | -7.498 | -31.793 | 20057 | 0 | 1 |  |  | H2O | 148.8361 | 0.142 | 0.296 | 99 | -0.012 | -0.021 | 29.94 | 844 | 0 | 30.313 | MT1-Rear | 11 | 67 | BSE
69 | P-11766 | 2025/05/29 20:46:03 | MT1-Rear-12 | 1 | -6.319 | -27.677 | 19754 | -1 | 1 |  |  | H2O | 148.8422 | 0.061 | 0.315 | 98 | -0.007 | -0.01 | 29.56 | 844 | 0 | 30.375 | MT1-Rear | 12 | 67 | BSE
70 | P-11766 | 2025/05/29 20:54:52 | MT1-Rear-12 | 2 | -6.296 | -27.098 | 20013 | -1 | 1 |  |  | H2O | 148.8484 | 0.171 | 0.341 | 103 | -0.036 | -0.033 | 30.99 | 844 | 0 | 30.375 | MT1-Rear | 12 | 67 | BSE
71 | P-11766 | 2025/05/29 21:03:41 | MT1-Rear-12 | 3 | -6.221 | -27.359 | 19728 | -1 | 1 |  |  | H2O | 148.8545 | 0.167 | 0.3 | 98 | -0.023 | 0.068 | 29.64 | 844 | 0 | 30.75 | MT1-Rear | 12 | 67 | BSE
72 | P-11766 | 2025/05/29 21:12:30 | MT1-Rear-12 | 4 | -6.299 | -27.191 | 19816 | 0 | 1 |  |  | H2O | 148.8606 | 0.197 | 0.264 | 99 | -0.052 | 0.018 | 29.9 | 844 | 0 | 30.375 | MT1-Rear | 12 | 67 | BSE
73 | P-11767 | 2025/05/29 21:21:19 | MT1-Rear-13 | 1 | -6.223 | -26.504 | 19631 | -1 | 1 |  |  | H2O | 148.8667 | 0.086 | 0.291 | 100 | -0.001 | -0.016 | 30.24 | 844 | 0 | 30.813 | MT1-Rear | 13 | 67 | BSE
74 | P-11767 | 2025/05/29 21:30:08 | MT1-Rear-13 | 2 | -6.221 | -26.236 | 19564 | -1 | 1 |  |  | H2O | 148.8728 | 0.139 | 0.356 | 90 | -0.034 | 0.036 | 29.7 | 844 | 0 | 31.063 | MT1-Rear | 13 | 67 | BSE
75 | P-11767 | 2025/05/29 21:38:59 | MT1-Rear-13 | 3 | -6.242 | -26.334 | 19845 | -1 | 1 |  |  | H2O | 148.879 | 0.152 | 0.332 | 100 | -0.021 | -0.008 | 30 | 844 | 0 | 30.813 | MT1-Rear | 13 | 67 | BSE
76 | P-11767 | 2025/05/29 21:47:49 | MT1-Rear-13 | 4 | -6.266 | -26.274 | 19781 | 0 | 1 |  |  | H2O | 148.8851 | 0.119 | 0.368 | 97 | -0.021 | 0.027 | 29.31 | 844 | 0 | 30.938 | MT1-Rear | 13 | 67 | BSE
77 | P-11768 | 2025/05/29 21:56:38 | MT1-Rear-14 | 1 | -6.194 | -26.06 | 20116 | -1 | 1 |  |  | H2O | 148.8912 | 0.164 | 0.236 | 101 | -0.033 | 0.01 | 30.36 | 844 | 0 | 30.813 | MT1-Rear | 14 | 67 | BSE
78 | P-11768 | 2025/05/29 22:05:28 | MT1-Rear-14 | 2 | -6.235 | -25.595 | 19974 | -1 | 1 |  |  | H2O | 148.8974 | 0.102 | 0.205 | 100 | 0.004 | 0.031 | 30.24 | 844 | 0 | 30.625 | MT1-Rear | 14 | 67 | BSE
79 | P-11768 | 2025/05/29 22:14:17 | MT1-Rear-14 | 3 | -6.071 | -25.667 | 20178 | -1 | 1 |  |  | H2O | 148.9035 | 0.116 | 0.281 | 99 | -0.023 | -0.016 | 29.94 | 844 | 0 | 30.688 | MT1-Rear | 14 | 67 | BSE
80 | P-11768 | 2025/05/29 22:23:07 | MT1-Rear-14 | 4 | -6.164 | -25.893 | 20193 | 0 | 1 |  |  | H2O | 148.9096 | 0.099 | 0.545 | 108 | -0.005 | 0.121 | 32.41 | 844 | 0 | 30.625 | MT1-Rear | 14 | 67 | BSE
81 | P-11769 | 2025/05/29 22:31:55 | MT1-Rear-15 | 1 | -5.988 | -25.389 | 19862 | -1 | 1 |  |  | H2O | 148.9157 | 0.101 | 0.378 | 99 | -0.019 | -0.033 | 29.69 | 844 | 0 | 30.375 | MT1-Rear | 15 | 67 | BSE
82 | P-11769 | 2025/05/29 22:40:45 | MT1-Rear-15 | 2 | -5.953 | -25.457 | 20071 | -1 | 1 |  |  | H2O | 148.9219 | 0.113 | 0.381 | 99 | -0.019 | -0.052 | 29.88 | 844 | 0 | 30.063 | MT1-Rear | 15 | 67 | BSE
83 | P-11769 | 2025/05/29 22:49:35 | MT1-Rear-15 | 3 | -6.071 | -25.527 | 19776 | -1 | 1 |  |  | H2O | 148.928 | 0.073 | 0.328 | 97 | -0.008 | -0.046 | 29.33 | 844 | 0 | 30.438 | MT1-Rear | 15 | 67 | BSE
84 | P-11769 | 2025/05/29 22:58:24 | MT1-Rear-15 | 4 | -5.997 | -25.755 | 20072 | 0 | 1 |  |  | H2O | 148.9341 | 0.145 | 0.154 | 101 | -0.022 | -0.006 | 30.38 | 844 | 0 | 30.313 | MT1-Rear | 15 | 67 | BSE
85 | P-11770 | 2025/05/29 23:07:14 | MT1-Rear-16 | 1 | -6.353 | -26.89 | 19727 | -1 | 1 |  |  | H2O | 148.9402 | 0.136 | 0.315 | 99 | -0.029 | -0.012 | 29.86 | 844 | 0 | 30.625 | MT1-Rear | 16 | 67 | BSE
86 | P-11770 | 2025/05/29 23:16:04 | MT1-Rear-16 | 2 | -6.4 | -27.399 | 19822 | -1 | 1 |  |  | H2O | 148.9464 | 0.158 | 0.38 | 98 | -0.015 | 0.06 | 29.39 | 844 | 0 | 30.625 | MT1-Rear | 16 | 67 | BSE
87 | P-11770 | 2025/05/29 23:24:54 | MT1-Rear-16 | 3 | -6.313 | -27.199 | 19742 | -1 | 1 |  |  | H2O | 148.9525 | 0.123 | 0.321 | 98 | -0.03 | 0.017 | 29.46 | 844 | 0 | 30.688 | MT1-Rear | 16 | 67 | BSE
88 | P-11770 | 2025/05/29 23:33:44 | MT1-Rear-16 | 4 | -6.409 | -27.289 | 19952 | 0 | 1 |  |  | H2O | 148.9586 | 0.133 | 0.26 | 96 | -0.029 | -0.026 | 29.03 | 844 | 0 | 30.938 | MT1-Rear | 16 | 67 | BSE
89 | P-11771 | 2025/05/29 23:42:35 | MT1-Rear-17 | 1 | -6.317 | -26.762 | 19759 | -1 | 1 |  |  | H2O | 148.9647 | 0.11 | 0.279 | 98 | -0.015 | 0.057 | 29.62 | 844 | 0 | 30.938 | MT1-Rear | 17 | 67 | BSE
90 | P-11771 | 2025/05/29 23:51:25 | MT1-Rear-17 | 2 | -6.283 | -26.968 | 19770 | -1 | 1 |  |  | H2O | 148.9709 | 0.181 | 0.252 | 97 | -0.017 | 0.004 | 29.11 | 844 | 0 | 30.813 | MT1-Rear | 17 | 67 | BSE
91 | P-11771 | 2025/05/30 00:00:15 | MT1-Rear-17 | 3 | -6.292 | -26.835 | 19737 | -1 | 1 |  |  | H2O | 148.9771 | 0.128 | 0.296 | 98 | -0.026 | -0.03 | 29.47 | 844 | 0 | 30.625 | MT1-Rear | 17 | 67 | BSE
92 | P-11771 | 2025/05/30 00:09:06 | MT1-Rear-17 | 4 | -6.262 | -27.023 | 19835 | 0 | 1 |  |  | H2O | 148.9832 | 0.097 | 0.328 | 99 | -0.018 | -0.037 | 29.94 | 844 | 0 | 30.625 | MT1-Rear | 17 | 67 | BSE
93 | P-11772 | 2025/05/30 00:17:56 | MT1-Rear-18 | 1 | -8.523 | -41.252 | 19810 | -1 | 1 |  |  | H2O | 148.9894 | 0.164 | 0.232 | 97 | -0.035 | -0.011 | 29.19 | 844 | 0 | 30.5 | MT1-Rear | 18 | 67 | BSE
94 | P-11772 | 2025/05/30 00:26:47 | MT1-Rear-18 | 2 | -8.685 | -42.606 | 19815 | -1 | 1 |  |  | H2O | 148.9955 | 0.097 | 0.338 | 96 | -0.006 | -0.028 | 28.84 | 844 | 0 | 30.125 | MT1-Rear | 18 | 67 | BSE
95 | P-11772 | 2025/05/30 00:35:37 | MT1-Rear-18 | 3 | -8.636 | -43.048 | 20113 | -1 | 1 |  |  | H2O | 149.0016 | 0.127 | 0.327 | 96 | -0.022 | -0.011 | 28.96 | 844 | 0 | 30 | MT1-Rear | 18 | 67 | BSE
96 | P-11772 | 2025/05/30 00:44:28 | MT1-Rear-18 | 4 | -8.714 | -43.008 | 19818 | 0 | 1 |  |  | H2O | 149.0077 | 0.126 | 0.345 | 98 | -0.024 | -0.011 | 29.48 | 844 | 0 | 29.813 | MT1-Rear | 18 | 67 | BSE
97 | P-11773 | 2025/05/30 00:53:18 | MT1-Rear-19 | 1 | -6.366 | -28.26 | 19840 | -1 | 1 |  |  | H2O | 149.0138 | 0.082 | 0.451 | 97 | 0.004 | 0.029 | 29.26 | 844 | 0 | 30.438 | MT1-Rear | 19 | 67 | BSE
98 | P-11773 | 2025/05/30 01:02:09 | MT1-Rear-19 | 2 | -6.262 | -26.284 | 19605 | -1 | 1 |  |  | H2O | 149.0201 | 0.139 | 0.265 | 100 | -0.021 | 0.011 | 30.23 | 844 | 0 | 30.125 | MT1-Rear | 19 | 67 | BSE
99 | P-11773 | 2025/05/30 01:10:59 | MT1-Rear-19 | 3 | -6.2 | -26.089 | 19555 | -1 | 1 |  |  | H2O | 149.0262 | 0.114 | 0.327 | 97 | -0.011 | -0.043 | 29.25 | 844 | 0 | 30.313 | MT1-Rear | 19 | 67 | BSE
100 | P-11773 | 2025/05/30 01:19:50 | MT1-Rear-19 | 4 | -6.246 | -25.611 | 19727 | 0 | 1 |  |  | H2O | 149.0323 | 0.106 | 0.306 | 101 | -0.013 | 0.054 | 30.29 | 844 | 0 | 30.563 | MT1-Rear | 19 | 67 | BSE
101 | P-11774 | 2025/05/30 01:28:41 | MT1-Rear-20 | 1 | -6.072 | -25.602 | 19876 | -1 | 1 |  |  | H2O | 149.0385 | 0.087 | 0.217 | 97 | -0.003 | 0.009 | 29.2 | 844 | 0 | 30.813 | MT1-Rear | 20 | 67 | BSE
102 | P-11774 | 2025/05/30 01:37:33 | MT1-Rear-20 | 2 | -6.09 | -25.546 | 19846 | -1 | 1 |  |  | H2O | 149.0446 | 0.127 | 0.31 | 95 | -0.026 | -0.031 | 28.66 | 844 | 0 | 31 | MT1-Rear | 20 | 67 | BSE
103 | P-11774 | 2025/05/30 01:46:23 | MT1-Rear-20 | 3 | -6.165 | -25.12 | 20217 | -1 | 1 |  |  | H2O | 149.0507 | 0.119 | 0.329 | 100 | 0.004 | 0.013 | 30.06 | 844 | 0 | 30.438 | MT1-Rear | 20 | 67 | BSE
104 | P-11774 | 2025/05/30 01:55:14 | MT1-Rear-20 | 4 | -6.159 | -25.197 | 19625 | 0 | 1 |  |  | H2O | 149.057 | 0.151 | 0.22 | 99 | -0.031 | -0.012 | 29.82 | 844 | 0 | 30.563 | MT1-Rear | 20 | 67 | BSE
105 | P-11775 | 2025/05/30 02:04:04 | MT1-Rear-21 | 1 | -6.304 | -25.871 | 20115 | -1 | 1 |  |  | H2O | 149.0631 | 0.107 | 0.356 | 100 | -0.004 | 0.086 | 29.98 | 844 | 0 | 30.188 | MT1-Rear | 21 | 67 | BSE
106 | P-11775 | 2025/05/30 02:12:55 | MT1-Rear-21 | 2 | -6.292 | -26.169 | 19968 | -1 | 1 |  |  | H2O | 149.0692 | 0.142 | 0.242 | 99 | -0.016 | 0.011 | 29.75 | 844 | 0 | 30.625 | MT1-Rear | 21 | 67 | BSE
107 | P-11775 | 2025/05/30 02:21:46 | MT1-Rear-21 | 3 | -6.257 | -26.151 | 20074 | -1 | 1 |  |  | H2O | 149.0753 | 0.129 | 0.272 | 97 | -0.018 | -0.034 | 29.31 | 844 | 0 | 30.563 | MT1-Rear | 21 | 67 | BSE
108 | P-11775 | 2025/05/30 02:30:37 | MT1-Rear-21 | 4 | -6.358 | -25.944 | 19506 | 0 | 1 |  |  | H2O | 149.0815 | 0.166 | 0.292 | 96 | -0.031 | -0.029 | 28.8 | 844 | 0 | 30.75 | MT1-Rear | 21 | 67 | BSE
109 | P-11776 | 2025/05/30 02:39:27 | MT1-Rear-22 | 1 | -6.175 | -25.578 | 20079 | -1 | 1 |  |  | H2O | 149.0876 | 0.122 | 0.324 | 96 | -0.019 | -0.024 | 28.8 | 844 | 0 | 30.5 | MT1-Rear | 22 | 67 | BSE
110 | P-11776 | 2025/05/30 02:48:20 | MT1-Rear-22 | 2 | -6.153 | -25.504 | 19760 | -1 | 1 |  |  | H2O | 149.0938 | 0.134 | 0.348 | 100 | -0.026 | 0.041 | 30.22 | 844 | 0 | 30.938 | MT1-Rear | 22 | 67 | BSE
111 | P-11776 | 2025/05/30 02:57:10 | MT1-Rear-22 | 3 | -6.108 | -25.547 | 19613 | -1 | 1 |  |  | H2O | 149.1 | 0.132 | 0.47 | 98 | -0.024 | 0.014 | 29.51 | 844 | 0 | 30.438 | MT1-Rear | 22 | 67 | BSE
112 | P-11776 | 2025/05/30 03:06:02 | MT1-Rear-22 | 4 | -6.097 | -25.485 | 19745 | 0 | 1 |  |  | H2O | 149.1061 | 0.101 | 0.267 | 94 | -0.013 | -0.036 | 28.35 | 844 | 0 | 30.563 | MT1-Rear | 22 | 67 | BSE
113 | P-11777 | 2025/05/30 03:14:53 | MT1-Rear-23 | 1 | -5.69 | -23.711 | 19706 | -1 | 1 |  |  | H2O | 149.1122 | 0.142 | 0.238 | 95 | -0.032 | 0.006 | 28.76 | 844 | 0 | 29.5 | MT1-Rear | 23 | 67 | BSE
114 | P-11777 | 2025/05/30 03:23:45 | MT1-Rear-23 | 2 | -5.745 | -23.247 | 19835 | -1 | 1 |  |  | H2O | 149.1184 | 0.12 | 0.252 | 89 | -0.028 | -0.018 | 29.24 | 844 | 0 | 30.375 | MT1-Rear | 23 | 67 | BSE
115 | P-11777 | 2025/05/30 03:32:37 | MT1-Rear-23 | 3 | -5.661 | -23.486 | 20136 | -1 | 1 |  |  | H2O | 149.1245 | 0.123 | 0.393 | 99 | -0.02 | -0.054 | 29.74 | 844 | 0 | 30.375 | MT1-Rear | 23 | 67 | BSE
116 | P-11777 | 2025/05/30 03:41:29 | MT1-Rear-23 | 4 | -5.737 | -23.171 | 19800 | 0 | 1 |  |  | H2O | 149.1307 | 0.099 | 0.282 | 95 | -0.022 | -0.013 | 28.7 | 844 | 0 | 30.25 | MT1-Rear | 23 | 67 | BSE
117 | P-11778 | 2025/05/30 03:50:20 | MT1-Rear-24 | 1 | -7.43 | -30.362 | 20283 | -1 | 1 |  |  | H2O | 149.1369 | 0.082 | 0.39 | 91 | 0 | 0.041 | 30.09 | 844 | 0 | 30.25 | MT1-Rear | 24 | 67 | BSE
118 | P-11778 | 2025/05/30 03:59:12 | MT1-Rear-24 | 2 | -7.584 | -31.308 | 19889 | -1 | 1 |  |  | H2O | 149.143 | 0.11 | 0.445 | 107 | 0.008 | 0.114 | 32.21 | 844 | 0 | 29.938 | MT1-Rear | 24 | 67 | BSE
119 | P-11778 | 2025/05/30 04:08:03 | MT1-Rear-24 | 3 | -7.498 | -31.134 | 19741 | -1 | 1 |  |  | H2O | 149.1492 | 0.171 | 0.367 | 97 | -0.033 | 0.041 | 29.25 | 844 | 0 | 29.813 | MT1-Rear | 24 | 67 | BSE
120 | P-11778 | 2025/05/30 04:16:54 | MT1-Rear-24 | 4 | -7.593 | -31.363 | 19743 | 0 | 1 |  |  | H2O | 149.1553 | 0.101 | 0.296 | 99 | -0.008 | 0.053 | 29.93 | 844 | 0 | 30.125 | MT1-Rear | 24 | 67 | BSE
121 | P-11779 | 2025/05/30 04:25:46 | MT1-Rear-25 | 1 | -6.54 | -27.115 | 20256 | -1 | 1 |  |  | H2O | 149.1614 | 0.119 | 0.408 | 98 | -0.012 | 0.096 | 29.66 | 844 | 0 | 30.563 | MT1-Rear | 25 | 67 | BSE
122 | P-11779 | 2025/05/30 04:34:37 | MT1-Rear-25 | 2 | -6.523 | -26.764 | 19814 | -1 | 1 |  |  | H2O | 149.1676 | 0.123 | 0.348 | 92 | -0.023 | 0.052 | 30.51 | 844 | 0 | 30.625 | MT1-Rear | 25 | 67 | BSE
123 | P-11779 | 2025/05/30 04:43:29 | MT1-Rear-25 | 3 | -6.439 | -26.992 | 19850 | -1 | 1 |  |  | H2O | 149.1738 | 0.086 | 0.269 | 97 | 0.006 | -0.023 | 29.1 | 844 | 0 | 30.438 | MT1-Rear | 25 | 67 | BSE
124 | P-11779 | 2025/05/30 04:52:21 | MT1-Rear-25 | 4 | -6.425 | -27.064 | 19794 | 0 | 1 |  |  | H2O | 149.1799 | 0.106 | 0.309 | 96 | -0.003 | -0.018 | 29 | 844 | 0 | 30.813 | MT1-Rear | 25 | 67 | BSE
125 | P-11780 | 2025/05/30 05:01:12 | MT1-Rear-26 | 1 | -4.852 | -9.663 | 20142 | -1 | 1 |  |  | H2O | 149.186 | 0.086 | 0.393 | 99 | -0.012 | 0.004 | 29.98 | 844 | 0 | 30.625 | MT1-Rear | 26 | 67 | BSE
126 | P-11780 | 2025/05/30 05:10:03 | MT1-Rear-26 | 2 | -4.711 | -7.801 | 19851 | -1 | 1 |  |  | H2O | 149.1923 | 0.096 | 0.31 | 97 | 0 | -0.034 | 29.06 | 844 | 0 | 30.438 | MT1-Rear | 26 | 67 | BSE
127 | P-11780 | 2025/05/30 05:18:54 | MT1-Rear-26 | 3 | -4.716 | -7.216 | 20197 | -1 | 1 |  |  | H2O | 149.1984 | 0.189 | 0.369 | 100 | -0.051 | -0.077 | 30.09 | 844 | 0 | 30.125 | MT1-Rear | 26 | 67 | BSE
128 | P-11780 | 2025/05/30 05:27:45 | MT1-Rear-26 | 4 | -4.703 | -7.093 | 19825 | 0 | 1 |  |  | H2O | 149.2045 | 0.139 | 0.401 | 98 | -0.006 | 0.033 | 29.54 | 844 | 0 | 30.5 | MT1-Rear | 26 | 67 | BSE
129 | P-11781 | 2025/05/30 05:36:37 | MT1-Rear-27 | 1 | -8.464 | -38.288 | 19853 | -1 | 1 |  |  | H2O | 149.2106 | 0.12 | 0.309 | 97 | -0.026 | 0.079 | 29.19 | 844 | 0 | 30.375 | MT1-Rear | 27 | 67 | BSE
130 | P-11781 | 2025/05/30 05:45:29 | MT1-Rear-27 | 2 | -8.607 | -41.39 | 20210 | -1 | 1 |  |  | H2O | 149.2169 | 0.113 | 0.453 | 103 | -0.008 | 0.107 | 30.98 | 844 | 0 | 29.563 | MT1-Rear | 27 | 67 | BSE
131 | P-11781 | 2025/05/30 05:54:21 | MT1-Rear-27 | 3 | -8.683 | -42.528 | 19746 | -1 | 1 |  |  | H2O | 149.223 | 0.134 | 0.391 | 98 | -0.007 | 0.059 | 29.46 | 844 | 0 | 29.375 | MT1-Rear | 27 | 67 | BSE
132 | P-11781 | 2025/05/30 06:03:14 | MT1-Rear-27 | 4 | -8.732 | -42.878 | 19821 | 0 | 1 |  |  | H2O | 149.2291 | 0.138 | 0.328 | 97 | -0.035 | -0.008 | 29.25 | 844 | 0 | 29.688 | MT1-Rear | 27 | 67 | BSE

## Run Summary
Line | Sample | Inj Nr | Sample Name | H2O_Mean | Good | Ignore | Time Code | d(18_16)Mean | d(D_H)Mean
1 | 1 | 1 | Blacksburg | 20064 | 1 | 0 | 2025/05/29 10:48:10 | -9.037 | -49.993 |  | Min | Max
2 | 1 | 2 | Blacksburg | 19486 | 1 | 0 | 2025/05/29 10:56:57 | -8.856 | -45.757 | 0 | 17000 | 23000
3 | 1 | 3 | Blacksburg | 19391 | 1 | 0 | 2025/05/29 11:05:43 | -8.719 | -44.86 | 500 | 17000 | 23000
4 | 1 | 4 | Blacksburg | 19438 | 1 | 0 | 2025/05/29 11:14:29 | -8.813 | -44.233
5 | 1 | 5 | Blacksburg | 19461 | 1 | 0 | 2025/05/29 11:23:16 | -8.793 | -44.072
6 | 1 | 6 | Blacksburg | 19726 | 1 | 0 | 2025/05/29 11:32:02 | -8.726 | -44.093
7 | 1 | 7 | Blacksburg | 19324 | 1 | 0 | 2025/05/29 11:40:49 | -8.84 | -44.062
8 | 1 | 8 | Blacksburg | 19254 | 1 | 0 | 2025/05/29 11:49:36 | -8.853 | -43.927
9 | 1 | 9 | Blacksburg | 19179 | 1 | 0 | 2025/05/29 11:58:22 | -8.816 | -43.805
10 | 1 | 10 | Blacksburg | 19081 | 1 | 0 | 2025/05/29 12:07:08 | -8.689 | -43.968
11 | 2 | 1 | Myrtle | 19659 | 1 | 0 | 2025/05/29 12:15:55 | -3.619 | -9.741
12 | 2 | 2 | Myrtle | 19607 | 1 | 0 | 2025/05/29 12:24:41 | -3.344 | -5.886
13 | 2 | 3 | Myrtle | 19472 | 1 | 0 | 2025/05/29 12:33:28 | -3.265 | -4.837
14 | 2 | 4 | Myrtle | 19577 | 1 | 0 | 2025/05/29 12:42:14 | -3.174 | -4.663
15 | 2 | 5 | Myrtle | 19751 | 1 | 0 | 2025/05/29 12:51:00 | -3.193 | -4.292
16 | 2 | 6 | Myrtle | 19509 | 1 | 0 | 2025/05/29 12:59:46 | -3.125 | -4.4
17 | 2 | 7 | Myrtle | 19533 | 1 | 0 | 2025/05/29 13:08:33 | -3.159 | -4.4 | Concentration of water vapor should be about 20,000 ppmV (roughly +/- 1000; but not outside 17,000 and 23,000 ppmV).
18 | 2 | 8 | Myrtle | 19435 | 1 | 0 | 2025/05/29 13:17:19 | -3.177 | -4.353 | Large scatter may indicate a problem with the syringe or the septa.
19 | 2 | 9 | Myrtle | 19488 | 1 | 0 | 2025/05/29 13:26:07 | -3.159 | -4.001
20 | 2 | 10 | Myrtle | 19468 | 1 | 0 | 2025/05/29 13:34:53 | -3.029 | -4.242
21 | 3 | 1 | Homer | 19343 | 1 | 0 | 2025/05/29 13:43:40 | -14.828 | -94.703
22 | 3 | 2 | Homer | 19615 | 1 | 0 | 2025/05/29 13:52:28 | -15.563 | -105.109
23 | 3 | 3 | Homer | 19512 | 1 | 0 | 2025/05/29 14:01:15 | -15.627 | -107.067
24 | 3 | 4 | Homer | 19476 | 1 | 0 | 2025/05/29 14:10:02 | -15.776 | -107.801
25 | 3 | 5 | Homer | 19495 | 1 | 0 | 2025/05/29 14:18:48 | -15.792 | -108.31
26 | 3 | 6 | Homer | 19516 | 1 | 0 | 2025/05/29 14:27:34 | -15.846 | -108.557
27 | 3 | 7 | Homer | 19505 | 1 | 0 | 2025/05/29 14:36:21 | -15.782 | -108.687
28 | 3 | 8 | Homer | 19877 | 1 | 0 | 2025/05/29 14:45:10 | -15.789 | -109.189
29 | 3 | 9 | Homer | 19506 | 1 | 0 | 2025/05/29 14:53:57 | -15.81 | -108.907
30 | 3 | 10 | Homer | 19573 | 1 | 0 | 2025/05/29 15:02:44 | -15.847 | -109.21
31 | 4 | 1 | Blacksburg | 19712 | 1 | 0 | 2025/05/29 15:11:31 | -9.211 | -52.491
32 | 4 | 2 | Blacksburg | 19602 | 1 | 0 | 2025/05/29 15:20:18 | -8.755 | -46.633
33 | 4 | 3 | Blacksburg | 19641 | 1 | 0 | 2025/05/29 15:29:05 | -8.788 | -45.174
34 | 4 | 4 | Blacksburg | 19606 | 1 | 0 | 2025/05/29 15:37:53 | -8.656 | -44.927
35 | 4 | 5 | Blacksburg | 20029 | 1 | 0 | 2025/05/29 15:46:40 | -8.649 | -44.37
36 | 4 | 6 | Blacksburg | 19636 | 1 | 0 | 2025/05/29 15:55:26 | -8.648 | -44.474
37 | 4 | 7 | Blacksburg | 19630 | 1 | 0 | 2025/05/29 16:04:13 | -8.666 | -44.201
38 | 4 | 8 | Blacksburg | 19653 | 1 | 0 | 2025/05/29 16:13:02 | -8.602 | -44.398
39 | 4 | 9 | Blacksburg | 19506 | 1 | 0 | 2025/05/29 16:21:49 | -8.666 | -44.137
40 | 4 | 10 | Blacksburg | 20084 | 1 | 0 | 2025/05/29 16:30:37 | -8.616 | -44.072
41 | 5 | 1 | Hawaii | 19765 | 1 | 0 | 2025/05/29 16:39:25 | -4.876 | -12.658
42 | 5 | 2 | Hawaii | 19396 | 1 | 0 | 2025/05/29 16:48:13 | -4.618 | -9.388
43 | 5 | 3 | Hawaii | 19480 | 1 | 0 | 2025/05/29 16:57:01 | -4.675 | -8.253
44 | 5 | 4 | Hawaii | 19884 | 1 | 0 | 2025/05/29 17:05:50 | -4.567 | -8.06
45 | 6 | 1 | C50 22may24 1.5m | 19392 | 1 | 0 | 2025/05/29 17:14:38 | -6.029 | -23.486
46 | 6 | 2 | C50 22may24 1.5m | 19727 | 1 | 0 | 2025/05/29 17:23:27 | -6.013 | -25.544
47 | 6 | 3 | C50 22may24 1.5m | 19593 | 1 | 0 | 2025/05/29 17:32:15 | -5.995 | -25.756
48 | 6 | 4 | C50 22may24 1.5m | 20102 | 1 | 0 | 2025/05/29 17:41:03 | -5.988 | -25.881
49 | 7 | 1 | CC3 19jun24 0.1m | 20110 | 1 | 0 | 2025/05/29 17:49:52 | -5.636 | -24.078
50 | 7 | 2 | CC3 19jun24 0.1m | 19621 | 1 | 0 | 2025/05/29 17:58:40 | -5.61 | -23.83
51 | 7 | 3 | CC3 19jun24 0.1m | 19479 | 1 | 0 | 2025/05/29 18:07:28 | -5.613 | -23.905
52 | 7 | 4 | CC3 19jun24 0.1m | 19820 | 1 | 0 | 2025/05/29 18:16:16 | -5.678 | -23.849
53 | 8 | 1 | C50 16sep24 6m | 20263 | 1 | 0 | 2025/05/29 18:25:04 | -4.714 | -20.541
54 | 8 | 2 | C50 16sep24 6m | 20060 | 1 | 0 | 2025/05/29 18:33:52 | -4.687 | -20.249
55 | 8 | 3 | C50 16sep24 6m | 19780 | 1 | 0 | 2025/05/29 18:42:40 | -4.688 | -19.427
56 | 8 | 4 | C50 16sep24 6m | 20073 | 1 | 0 | 2025/05/29 18:51:27 | -4.65 | -19.78
57 | 9 | 1 | Blacksburg | 19702 | 1 | 0 | 2025/05/29 19:00:16 | -8.315 | -40.483
58 | 9 | 2 | Blacksburg | 20124 | 1 | 0 | 2025/05/29 19:09:05 | -8.536 | -42.795
59 | 9 | 3 | Blacksburg | 19740 | 1 | 0 | 2025/05/29 19:17:53 | -8.647 | -43.076
60 | 9 | 4 | Blacksburg | 19934 | 1 | 0 | 2025/05/29 19:26:42 | -8.611 | -43.265
61 | 10 | 1 | CC5 22may24 1.5m | 19727 | 1 | 0 | 2025/05/29 19:35:30 | -6.184 | -28.182
62 | 10 | 2 | CC5 22may24 1.5m | 20162 | 1 | 0 | 2025/05/29 19:44:18 | -6.038 | -26.671
63 | 10 | 3 | CC5 22may24 1.5m | 20150 | 1 | 0 | 2025/05/29 19:53:08 | -5.924 | -25.878
64 | 10 | 4 | CC5 22may24 1.5m | 19702 | 1 | 0 | 2025/05/29 20:01:57 | -6.069 | -25.92
65 | 11 | 1 | CS2 19jun24 0.1m | 20002 | 1 | 0 | 2025/05/29 20:10:46 | -7.348 | -31.022
66 | 11 | 2 | CS2 19jun24 0.1m | 20098 | 1 | 0 | 2025/05/29 20:19:35 | -7.361 | -31.66
67 | 11 | 3 | CS2 19jun24 0.1m | 19819 | 1 | 0 | 2025/05/29 20:28:25 | -7.439 | -31.797
68 | 11 | 4 | CS2 19jun24 0.1m | 20057 | 1 | 0 | 2025/05/29 20:37:14 | -7.498 | -31.793
69 | 12 | 1 | CC4 19jun24 9m | 19754 | 1 | 0 | 2025/05/29 20:46:03 | -6.319 | -27.677
70 | 12 | 2 | CC4 19jun24 9m | 20013 | 1 | 0 | 2025/05/29 20:54:52 | -6.296 | -27.098
71 | 12 | 3 | CC4 19jun24 9m | 19728 | 1 | 0 | 2025/05/29 21:03:41 | -6.221 | -27.359
72 | 12 | 4 | CC4 19jun24 9m | 19816 | 1 | 0 | 2025/05/29 21:12:30 | -6.299 | -27.191
73 | 13 | 1 | CC5 22may24 6m | 19631 | 1 | 0 | 2025/05/29 21:21:19 | -6.223 | -26.504
74 | 13 | 2 | CC5 22may24 6m | 19564 | 1 | 0 | 2025/05/29 21:30:08 | -6.221 | -26.236
75 | 13 | 3 | CC5 22may24 6m | 19845 | 1 | 0 | 2025/05/29 21:38:59 | -6.242 | -26.334
76 | 13 | 4 | CC5 22may24 6m | 19781 | 1 | 0 | 2025/05/29 21:47:49 | -6.266 | -26.274
77 | 14 | 1 | C50 16apr25 1.5m | 20116 | 1 | 0 | 2025/05/29 21:56:38 | -6.194 | -26.06
78 | 14 | 2 | C50 16apr25 1.5m | 19974 | 1 | 0 | 2025/05/29 22:05:28 | -6.235 | -25.595
79 | 14 | 3 | C50 16apr25 1.5m | 20178 | 1 | 0 | 2025/05/29 22:14:17 | -6.071 | -25.667
80 | 14 | 4 | C50 16apr25 1.5m | 20193 | 1 | 0 | 2025/05/29 22:23:07 | -6.164 | -25.893
81 | 15 | 1 | CC5 22may24 0.1m | 19862 | 1 | 0 | 2025/05/29 22:31:55 | -5.988 | -25.389
82 | 15 | 2 | CC5 22may24 0.1m | 20071 | 1 | 0 | 2025/05/29 22:40:45 | -5.953 | -25.457
83 | 15 | 3 | CC5 22may24 0.1m | 19776 | 1 | 0 | 2025/05/29 22:49:35 | -6.071 | -25.527
84 | 15 | 4 | CC5 22may24 0.1m | 20072 | 1 | 0 | 2025/05/29 22:58:24 | -5.997 | -25.755
85 | 16 | 1 | CC5 22may24 9m | 19727 | 1 | 0 | 2025/05/29 23:07:14 | -6.353 | -26.89
86 | 16 | 2 | CC5 22may24 9m | 19822 | 1 | 0 | 2025/05/29 23:16:04 | -6.4 | -27.399
87 | 16 | 3 | CC5 22may24 9m | 19742 | 1 | 0 | 2025/05/29 23:24:54 | -6.313 | -27.199
88 | 16 | 4 | CC5 22may24 9m | 19952 | 1 | 0 | 2025/05/29 23:33:44 | -6.409 | -27.289
89 | 17 | 1 | CC4 22may24 BOT | 19759 | 1 | 0 | 2025/05/29 23:42:35 | -6.317 | -26.762
90 | 17 | 2 | CC4 22may24 BOT | 19770 | 1 | 0 | 2025/05/29 23:51:25 | -6.283 | -26.968
91 | 17 | 3 | CC4 22may24 BOT | 19737 | 1 | 0 | 2025/05/30 00:00:15 | -6.292 | -26.835
92 | 17 | 4 | CC4 22may24 BOT | 19835 | 1 | 0 | 2025/05/30 00:09:06 | -6.262 | -27.023
93 | 18 | 1 | Blacksburg | 19810 | 1 | 0 | 2025/05/30 00:17:56 | -8.523 | -41.252
94 | 18 | 2 | Blacksburg | 19815 | 1 | 0 | 2025/05/30 00:26:47 | -8.685 | -42.606
95 | 18 | 3 | Blacksburg | 20113 | 1 | 0 | 2025/05/30 00:35:37 | -8.636 | -43.048
96 | 18 | 4 | Blacksburg | 19818 | 1 | 0 | 2025/05/30 00:44:28 | -8.714 | -43.008
97 | 19 | 1 | CC3 16apr25 1.5m | 19840 | 1 | 0 | 2025/05/30 00:53:18 | -6.366 | -28.26
98 | 19 | 2 | CC3 16apr25 1.5m | 19605 | 1 | 0 | 2025/05/30 01:02:09 | -6.262 | -26.284
99 | 19 | 3 | CC3 16apr25 1.5m | 19555 | 1 | 0 | 2025/05/30 01:10:59 | -6.2 | -26.089
100 | 19 | 4 | CC3 16apr25 1.5m | 19727 | 1 | 0 | 2025/05/30 01:19:50 | -6.246 | -25.611
101 | 20 | 1 | CP2 22may24 BOT | 19876 | 1 | 0 | 2025/05/30 01:28:41 | -6.072 | -25.602
102 | 20 | 2 | CP2 22may24 BOT | 19846 | 1 | 0 | 2025/05/30 01:37:33 | -6.09 | -25.546
103 | 20 | 3 | CP2 22may24 BOT | 20217 | 1 | 0 | 2025/05/30 01:46:23 | -6.165 | -25.12
104 | 20 | 4 | CP2 22may24 BOT | 19625 | 1 | 0 | 2025/05/30 01:55:14 | -6.159 | -25.197
105 | 21 | 1 | CC4 16apr25 BOT | 20115 | 1 | 0 | 2025/05/30 02:04:04 | -6.304 | -25.871
106 | 21 | 2 | CC4 16apr25 BOT | 19968 | 1 | 0 | 2025/05/30 02:12:55 | -6.292 | -26.169
107 | 21 | 3 | CC4 16apr25 BOT | 20074 | 1 | 0 | 2025/05/30 02:21:46 | -6.257 | -26.151
108 | 21 | 4 | CC4 16apr25 BOT | 19506 | 1 | 0 | 2025/05/30 02:30:37 | -6.358 | -25.944
109 | 22 | 1 | CP1 22may24 0.1m | 20079 | 1 | 0 | 2025/05/30 02:39:27 | -6.175 | -25.578
110 | 22 | 2 | CP1 22may24 0.1m | 19760 | 1 | 0 | 2025/05/30 02:48:20 | -6.153 | -25.504
111 | 22 | 3 | CP1 22may24 0.1m | 19613 | 1 | 0 | 2025/05/30 02:57:10 | -6.108 | -25.547
112 | 22 | 4 | CP1 22may24 0.1m | 19745 | 1 | 0 | 2025/05/30 03:06:02 | -6.097 | -25.485
113 | 23 | 1 | CP2 19jun24 0.1m | 19706 | 1 | 0 | 2025/05/30 03:14:53 | -5.69 | -23.711
114 | 23 | 2 | CP2 19jun24 0.1m | 19835 | 1 | 0 | 2025/05/30 03:23:45 | -5.745 | -23.247
115 | 23 | 3 | CP2 19jun24 0.1m | 20136 | 1 | 0 | 2025/05/30 03:32:37 | -5.661 | -23.486
116 | 23 | 4 | CP2 19jun24 0.1m | 19800 | 1 | 0 | 2025/05/30 03:41:29 | -5.737 | -23.171
117 | 24 | 1 | CS1 19jun24 0.1m | 20283 | 1 | 0 | 2025/05/30 03:50:20 | -7.43 | -30.362
118 | 24 | 2 | CS1 19jun24 0.1m | 19889 | 1 | 0 | 2025/05/30 03:59:12 | -7.584 | -31.308
119 | 24 | 3 | CS1 19jun24 0.1m | 19741 | 1 | 0 | 2025/05/30 04:08:03 | -7.498 | -31.134
120 | 24 | 4 | CS1 19jun24 0.1m | 19743 | 1 | 0 | 2025/05/30 04:16:54 | -7.593 | -31.363
121 | 25 | 1 | CC5 22may24 9m DUP | 20256 | 1 | 0 | 2025/05/30 04:25:46 | -6.54 | -27.115
122 | 25 | 2 | CC5 22may24 9m DUP | 19814 | 1 | 0 | 2025/05/30 04:34:37 | -6.523 | -26.764
123 | 25 | 3 | CC5 22may24 9m DUP | 19850 | 1 | 0 | 2025/05/30 04:43:29 | -6.439 | -26.992
124 | 25 | 4 | CC5 22may24 9m DUP | 19794 | 1 | 0 | 2025/05/30 04:52:21 | -6.425 | -27.064
125 | 26 | 1 | Hawaii | 20142 | 1 | 0 | 2025/05/30 05:01:12 | -4.852 | -9.663
126 | 26 | 2 | Hawaii | 19851 | 1 | 0 | 2025/05/30 05:10:03 | -4.711 | -7.801
127 | 26 | 3 | Hawaii | 20197 | 1 | 0 | 2025/05/30 05:18:54 | -4.716 | -7.216
128 | 26 | 4 | Hawaii | 19825 | 1 | 0 | 2025/05/30 05:27:45 | -4.703 | -7.093
129 | 27 | 1 | Blacksburg | 19853 | 1 | 0 | 2025/05/30 05:36:37 | -8.464 | -38.288
130 | 27 | 2 | Blacksburg | 20210 | 1 | 0 | 2025/05/30 05:45:29 | -8.607 | -41.39
131 | 27 | 3 | Blacksburg | 19746 | 1 | 0 | 2025/05/30 05:54:21 | -8.683 | -42.528
132 | 27 | 4 | Blacksburg | 19821 | 1 | 0 | 2025/05/30 06:03:14 | -8.732 | -42.878

## 18O
Line | Sample | Injection | Sample Name | Correction Type | Good | Ignore | d(18_16) Picarro Raw | Stdev Raw | d(18_16) Memory Corrected | Stdev Memory Corrected | d(18_16) Memory Corrected Drifts | d(18_16) Drift Corrected | Drift Corrected Drifts | d(18_16) Drift Corrected High, Low, and Drift | d(18_16) Reference Defined | d18O Final Corrected Value | Post run corrections:
1 | 1 | 1 | Blacksburg | Drift | 1 | -1 | -9.037 | [9 empty] | (1) Memory correction |  |  | (1) Graphs for memory correction
2 | 1 | 2 | Blacksburg | Drift | 1 | -1 | -8.856 | [9 empty] | Combined Stdev | 0.056323
3 | 1 | 3 | Blacksburg | Drift | 1 | -1 | -8.719 | [9 empty] | Date:
4 | 1 | 4 | Blacksburg | Drift | 1 | -1 | -8.813 | [9 empty] | injection | 18O memory coef
5 | 1 | 5 | Blacksburg | Drift | 1 | -1 | -8.793 | [9 empty] | 1 | 0.90931
6 | 1 | 6 | Blacksburg | Drift | 1 | -1 | -8.726 | [9 empty] | 2 | 0.97364
7 | 1 | 7 | Blacksburg | Drift | 1 | 0 | -8.84 |  | -8.84 |  | -8.84 | -8.8416 | -8.8416 |  |  | -7.7753 | 3 | 0.97793
8 | 1 | 8 | Blacksburg | Drift | 1 | 0 | -8.853 |  | -8.853 |  | -8.853 | -8.8548 | -8.8548 |  |  | -7.788 | 4 | 0.99192
9 | 1 | 9 | Blacksburg | Drift | 1 | 0 | -8.816 |  | -8.816 |  | -8.816 | -8.8181 | -8.8181 |  |  | -7.7527 | 5 | 0.99254
10 | 1 | 10 | Blacksburg | Drift | 1 | 0 | -8.689 | 0.075244 | -8.689 | 0.075244 | -8.689 | -8.6913 | -8.6913 |  |  | -7.631 | 6 | 0.99428
11 | 2 | 1 | Myrtle | High Reference | 1 | 0 | -3.619 |  | -3.1592 |  |  | -3.1617 |  | -3.1617 | -2.2908 | -2.3209 | 7 | 0.99428
12 | 2 | 2 | Myrtle | High Reference | 1 | 0 | -3.344 |  | -3.2031 |  |  | -3.2058 |  | -3.2058 | -2.2908 | -2.3633 | 8 | 0.99428
13 | 2 | 3 | Myrtle | High Reference | 1 | 0 | -3.265 |  | -3.1453 |  |  | -3.1483 |  | -3.1483 | -2.2908 | -2.308 | 9 | 0.99428
14 | 2 | 4 | Myrtle | High Reference | 1 | 0 | -3.174 |  | -3.1295 |  |  | -3.1327 |  | -3.1327 | -2.2908 | -2.293 | 10 | 1
15 | 2 | 5 | Myrtle | High Reference | 1 | 0 | -3.193 |  | -3.152 |  |  | -3.1554 |  | -3.1554 | -2.2908 | -2.3149
16 | 2 | 6 | Myrtle | High Reference | 1 | 0 | -3.125 |  | -3.0932 |  |  | -3.0968 |  | -3.0968 | -2.2908 | -2.2587 | (2) Drift correction
17 | 2 | 7 | Myrtle | High Reference | 1 | 0 | -3.159 |  | -3.1274 |  |  | -3.1313 |  | -3.1313 | -2.2908 | -2.2917 | Slope m: | 0.00022797
18 | 2 | 8 | Myrtle | High Reference | 1 | 0 | -3.177 |  | -3.1455 |  |  | -3.1496 |  | -3.1496 | -2.2908 | -2.3093 | Stdev Mem Corr. | 0.067842
19 | 2 | 9 | Myrtle | High Reference | 1 | 0 | -3.159 |  | -3.1274 |  |  | -3.1317 |  | -3.1317 | -2.2908 | -2.2921 | Stdev Drift Corr. | 0.067005
20 | 2 | 10 | Myrtle | High Reference | 1 | 0 | -3.029 | 0.16126 | -3.029 | 0.045594 |  | -3.0336 |  | -3.0336 | -2.2908 | -2.1979
21 | 3 | 1 | Homer | Low Reference | 1 | 0 | -14.828 |  | -15.8981 |  |  | -15.9029 |  | -15.9029 | -14.5399 | -14.5563 | (3) Normalization to VSMOW scale
22 | 3 | 2 | Homer | Low Reference | 1 | 0 | -15.563 |  | -15.8934 |  |  | -15.8984 |  | -15.8984 | -14.5399 | -14.552 | slope m | 0.9603
23 | 3 | 3 | Homer | Low Reference | 1 | 0 | -15.627 |  | -15.905 |  |  | -15.9103 |  | -15.9103 | -14.5399 | -14.5634 | interception t | 0.71524
24 | 3 | 4 | Homer | Low Reference | 1 | 0 | -15.776 |  | -15.8789 |  |  | -15.8844 |  | -15.8844 | -14.5399 | -14.5385
25 | 3 | 5 | Homer | Low Reference | 1 | 0 | -15.792 |  | -15.8872 |  |  | -15.8929 |  | -15.8929 | -14.5399 | -14.5467 | (4) Quality control |  | offset
26 | 3 | 6 | Homer | Low Reference | 1 | 0 | -15.846 |  | -15.9193 |  |  | -15.9252 |  | -15.9252 | -14.5399 | -14.5777 | QC Check Defined | -3.6108 | accepted ± 0.15
27 | 3 | 7 | Homer | Low Reference | 1 | 0 | -15.782 |  | -15.8549 |  |  | -15.8611 |  | -15.8611 | -14.5399 | -14.5161 | QC Check Vial 5 | -3.6566 | -0.04582
28 | 3 | 8 | Homer | Low Reference | 1 | 0 | -15.789 |  | -15.8619 |  |  | -15.8683 |  | -15.8683 | -14.5399 | -14.5231 | QC Check Vial 26 | -3.8121 | -0.20124
29 | 3 | 9 | Homer | Low Reference | 1 | 0 | -15.81 |  | -15.8831 |  |  | -15.8897 |  | -15.8897 | -14.5399 | -14.5436
30 | 3 | 10 | Homer | Low Reference | 1 | 0 | -15.847 | 0.30855 | -15.847 | 0.02285 |  | -15.8538 |  | -15.8538 | -14.5399 | -14.5092 | Drift Defined | -7.5615 | accepted ± 0.15
31 | 4 | 1 | Blacksburg | Drift | 1 | 0 | -9.211 |  | -8.6092 |  |  | -8.6162 |  | -8.6162 | -7.5615 | -7.5589 | Drift Average | -7.6653 | -0.10377
32 | 4 | 2 | Blacksburg | Drift | 1 | 0 | -8.755 |  | -8.568 |  |  | -8.5753 |  | -8.5753 | -7.5615 | -7.5196 | Drift Vial 4 | -7.5581 | 0.0033683
33 | 4 | 3 | Blacksburg | Drift | 1 | 0 | -8.788 |  | -8.6322 |  |  | -8.6397 |  | -8.6397 | -7.5615 | -7.5815 | Drift Vial 9 | -7.6195 | -0.058036
34 | 4 | 4 | Blacksburg | Drift | 1 | 0 | -8.656 |  | -8.5979 |  |  | -8.6057 |  | -8.6057 | -7.5615 | -7.5488 | Drift Vial 18 | -7.6838 | -0.12236
35 | 4 | 5 | Blacksburg | Drift | 1 | 0 | -8.649 |  | -8.5953 |  |  | -8.6033 |  | -8.6033 | -7.5615 | -7.5465 | Drift Vial 27 | -7.728 | -0.16655
36 | 4 | 6 | Blacksburg | Drift | 1 | 0 | -8.648 |  | -8.6069 |  |  | -8.6151 |  | -8.6151 | -7.5615 | -7.5578
37 | 4 | 7 | Blacksburg | Drift | 1 | 0 | -8.666 |  | -8.625 |  |  | -8.6334 |  | -8.6334 | -7.5615 | -7.5754 | Reference Defined Values |  | d18O
38 | 4 | 8 | Blacksburg | Drift | 1 | 0 | -8.602 |  | -8.5606 |  |  | -8.5693 |  | -8.5693 | -7.5615 | -7.5138 | High Reference | Myrtle Beach | -2.2908
39 | 4 | 9 | Blacksburg | Drift | 1 | 0 | -8.666 |  | -8.625 |  |  | -8.6338 |  | -8.6338 | -7.5615 | -7.5758 | Low Reference | Homer | -14.5399
40 | 4 | 10 | Blacksburg | Drift | 1 | 0 | -8.616 | 0.17997 | -8.616 | 0.023903 |  | -8.6251 |  | -8.6251 | -7.5615 | -7.5675 | Drift | Blacksburg | -7.5615
41 | 5 | 1 | Hawaii | QC Check | 1 | 0 | -4.876 |  | -4.5368 |  |  | -4.5462 |  |  |  | -3.6504 | QC Check | Hawaii | -3.6108
42 | 5 | 2 | Hawaii | QC Check | 1 | 0 | -4.618 |  | -4.5126 |  |  | -4.5222 |  |  |  | -3.6274 | Calibration Date:
43 | 5 | 3 | Hawaii | QC Check | 1 | 0 | -4.675 |  | -4.588 |  |  | -4.5978 |  |  |  | -3.7001
44 | 5 | 4 | Hawaii | QC Check | 1 | 0 | -4.567 | 0.13539 | -4.5343 | 0.031967 |  | -4.5443 |  |  |  | -3.6487
45 | 6 | 1 | C50 22may24 1.5m |  | 1 | 0 | -6.029 |  | -6.1616 |  |  | -6.1719 |  |  |  | -5.2116
46 | 6 | 2 | C50 22may24 1.5m |  | 1 | 0 | -6.013 |  | -6.0511 |  |  | -6.0616 |  |  |  | -5.1057 | Denali | -19.9287
47 | 6 | 3 | C50 22may24 1.5m |  | 1 | 0 | -5.995 |  | -6.0265 |  |  | -6.0372 |  |  |  | -5.0823 | Alberta | -18.8392
48 | 6 | 4 | C50 22may24 1.5m |  | 1 | 0 | -5.988 | 0.018464 | -5.9995 | 0.071142 |  | -6.0104 |  |  |  | -5.0566 | MT BHB | -18.431
49 | 7 | 1 | CC3 19jun24 0.1m |  | 1 | 0 | -5.636 |  | -5.6041 |  |  | -5.6152 |  |  |  | -4.6771 | Anchorage | -16.4204
50 | 7 | 2 | CC3 19jun24 0.1m |  | 1 | 0 | -5.61 |  | -5.6 |  |  | -5.6114 |  |  |  | -4.6734 | Homer | -14.5399
51 | 7 | 3 | CC3 19jun24 0.1m |  | 1 | 0 | -5.613 |  | -5.6047 |  |  | -5.6164 |  |  |  | -4.6781 | MT Creek | -13.3562
52 | 7 | 4 | CC3 19jun24 0.1m |  | 1 | 0 | -5.678 | 0.031394 | -5.6755 | 0.036335 |  | -5.6874 |  |  |  | -4.7463 | Nickel Mines | -7.876
53 | 8 | 1 | C50 16sep24 6m |  | 1 | 0 | -4.714 |  | -4.6266 |  |  | -4.6387 |  |  |  | -3.7393 | Blacksburg | -7.5615
54 | 8 | 2 | C50 16sep24 6m |  | 1 | 0 | -4.687 |  | -4.6609 |  |  | -4.6732 |  |  |  | -3.7724 | Denton's Well | -6.87
55 | 8 | 3 | C50 16sep24 6m |  | 1 | 0 | -4.688 |  | -4.6662 |  |  | -4.6787 |  |  |  | -3.7777 | Hawaii | -3.6108
56 | 8 | 4 | C50 16sep24 6m |  | 1 | 0 | -4.65 | 0.026323 | -4.6417 | 0.018178 |  | -4.6545 |  |  |  | -3.7544 | Myrtle Beach | -2.2908
57 | 9 | 1 | Blacksburg | Drift | 1 | 0 | -8.315 |  | -8.6474 |  | -8.6474 | -8.6604 | -8.6604 |  |  | -7.6013 | Heavy Oxygen | 3.1753
58 | 9 | 2 | Blacksburg | Drift | 1 | 0 | -8.536 |  | -8.6384 |  | -8.6384 | -8.6517 | -8.6517 |  |  | -7.5929 | Heavy Hydrogen | 11.1372
59 | 9 | 3 | Blacksburg | Drift | 1 | 0 | -8.647 |  | -8.7352 |  | -8.7352 | -8.7487 | -8.7487 |  |  | -7.6861 | Andres Heavy Water | 17.5664
60 | 9 | 4 | Blacksburg | Drift | 1 | 0 | -8.611 | 0.14886 | -8.643 | 0.046278 | -8.643 | -8.6567 | -8.6567 |  |  | -7.5977 | Picarro ZERO | 0.3
61 | 10 | 1 | CC5 22may24 1.5m |  | 1 | 0 | -6.184 |  | -5.9639 |  |  | -5.9778 |  |  |  | -5.0252 | Picarro MID | -20.6
62 | 10 | 2 | CC5 22may24 1.5m |  | 1 | 0 | -6.038 |  | -5.9702 |  |  | -5.9843 |  |  |  | -5.0315 | Picarro DEPL | -29.6
63 | 10 | 3 | CC5 22may24 1.5m |  | 1 | 0 | -5.924 |  | -5.8647 |  |  | -5.8791 |  |  |  | -4.9304
64 | 10 | 4 | CC5 22may24 1.5m |  | 1 | 0 | -6.069 | 0.1069 | -6.0485 | 0.075309 |  | -6.0631 |  |  |  | -5.1071
65 | 11 | 1 | CS2 19jun24 0.1m |  | 1 | 0 | -7.348 |  | -7.464 |  |  | -7.4788 |  |  |  | -6.4667
66 | 11 | 2 | CS2 19jun24 0.1m |  | 1 | 0 | -7.361 |  | -7.3951 |  |  | -7.4101 |  |  |  | -6.4007
67 | 11 | 3 | CS2 19jun24 0.1m |  | 1 | 0 | -7.439 |  | -7.4692 |  |  | -7.4845 |  |  |  | -6.4721
68 | 11 | 4 | CS2 19jun24 0.1m |  | 1 | 0 | -7.498 | 0.070288 | -7.5095 | 0.047509 |  | -7.525 |  |  |  | -6.511
69 | 12 | 1 | CC4 19jun24 9m |  | 1 | 0 | -6.319 |  | -6.2121 |  |  | -6.2278 |  |  |  | -5.2653
70 | 12 | 2 | CC4 19jun24 9m |  | 1 | 0 | -6.296 |  | -6.2643 |  |  | -6.2803 |  |  |  | -5.3157
71 | 12 | 3 | CC4 19jun24 9m |  | 1 | 0 | -6.221 |  | -6.1928 |  |  | -6.209 |  |  |  | -5.2473
72 | 12 | 4 | CC4 19jun24 9m |  | 1 | 0 | -6.299 | 0.043061 | -6.2893 | 0.044828 |  | -6.3057 |  |  |  | -5.3401
73 | 13 | 1 | CC5 22may24 6m |  | 1 | 0 | -6.223 |  | -6.2161 |  |  | -6.2327 |  |  |  | -5.2701
74 | 13 | 2 | CC5 22may24 6m |  | 1 | 0 | -6.221 |  | -6.2189 |  |  | -6.2358 |  |  |  | -5.273
75 | 13 | 3 | CC5 22may24 6m |  | 1 | 0 | -6.242 |  | -6.2407 |  |  | -6.2578 |  |  |  | -5.2942
76 | 13 | 4 | CC5 22may24 6m |  | 1 | 0 | -6.266 | 0.020928 | -6.2657 | 0.023034 |  | -6.2831 |  |  |  | -5.3184
77 | 14 | 1 | C50 16apr25 1.5m |  | 1 | 0 | -6.194 |  | -6.1875 |  |  | -6.205 |  |  |  | -5.2434
78 | 14 | 2 | C50 16apr25 1.5m |  | 1 | 0 | -6.235 |  | -6.2342 |  |  | -6.252 |  |  |  | -5.2885
79 | 14 | 3 | C50 16apr25 1.5m |  | 1 | 0 | -6.071 |  | -6.0667 |  |  | -6.0847 |  |  |  | -5.1279
80 | 14 | 4 | C50 16apr25 1.5m |  | 1 | 0 | -6.164 | 0.069699 | -6.1632 | 0.070569 |  | -6.1814 |  |  |  | -5.2208
81 | 15 | 1 | CC5 22may24 0.1m |  | 1 | 0 | -5.988 |  | -5.972 |  |  | -5.9905 |  |  |  | -5.0374
82 | 15 | 2 | CC5 22may24 0.1m |  | 1 | 0 | -5.953 |  | -5.9474 |  |  | -5.9661 |  |  |  | -5.014
83 | 15 | 3 | CC5 22may24 0.1m |  | 1 | 0 | -6.071 |  | -6.0689 |  |  | -6.0879 |  |  |  | -5.1309
84 | 15 | 4 | CC5 22may24 0.1m |  | 1 | 0 | -5.997 | 0.049608 | -5.9957 | 0.052453 |  | -6.0148 |  |  |  | -5.0608
85 | 16 | 1 | CC5 22may24 9m |  | 1 | 0 | -6.353 |  | -6.3853 |  |  | -6.4047 |  |  |  | -5.4352
86 | 16 | 2 | CC5 22may24 9m |  | 1 | 0 | -6.4 |  | -6.4106 |  |  | -6.4302 |  |  |  | -5.4597
87 | 16 | 3 | CC5 22may24 9m |  | 1 | 0 | -6.313 |  | -6.32 |  |  | -6.3398 |  |  |  | -5.3729
88 | 16 | 4 | CC5 22may24 9m |  | 1 | 0 | -6.409 | 0.044545 | -6.4123 | 0.043194 |  | -6.4324 |  |  |  | -5.4618
89 | 17 | 1 | CC4 22may24 BOT |  | 1 | 0 | -6.317 |  | -6.3087 |  |  | -6.3289 |  |  |  | -5.3624
90 | 17 | 2 | CC4 22may24 BOT |  | 1 | 0 | -6.283 |  | -6.2797 |  |  | -6.3002 |  |  |  | -5.3348
91 | 17 | 3 | CC4 22may24 BOT |  | 1 | 0 | -6.292 |  | -6.2894 |  |  | -6.3102 |  |  |  | -5.3444
92 | 17 | 4 | CC4 22may24 BOT |  | 1 | 0 | -6.262 | 0.022782 | -6.2608 | 0.019933 |  | -6.2818 |  |  |  | -5.3172
93 | 18 | 1 | Blacksburg | Drift | 1 | 0 | -8.523 |  | -8.7281 |  | -8.7281 | -8.7493 | -8.7493 |  |  | -7.6867
94 | 18 | 2 | Blacksburg | Drift | 1 | 0 | -8.685 |  | -8.7489 |  | -8.7489 | -8.7703 | -8.7703 |  |  | -7.7069
95 | 18 | 3 | Blacksburg | Drift | 1 | 0 | -8.636 |  | -8.6884 |  | -8.6884 | -8.71 | -8.71 |  |  | -7.649
96 | 18 | 4 | Blacksburg | Drift | 1 | 0 | -8.714 | 0.084073 | -8.7338 | 0.0258 | -8.7338 | -8.7557 | -8.7557 |  |  | -7.6928
97 | 19 | 1 | CC3 16apr25 1.5m |  | 1 | 0 | -6.366 |  | -6.1531 |  |  | -6.1752 |  |  |  | -5.2148
98 | 19 | 2 | CC3 16apr25 1.5m |  | 1 | 0 | -6.262 |  | -6.1974 |  |  | -6.2197 |  |  |  | -5.2575
99 | 19 | 3 | CC3 16apr25 1.5m |  | 1 | 0 | -6.2 |  | -6.1445 |  |  | -6.1671 |  |  |  | -5.207
100 | 19 | 4 | CC3 16apr25 1.5m |  | 1 | 0 | -6.246 | 0.070112 | -6.2261 | 0.038333 |  | -6.2489 |  |  |  | -5.2855
101 | 20 | 1 | CP2 22may24 BOT |  | 1 | 0 | -6.072 |  | -6.0562 |  |  | -6.0792 |  |  |  | -5.1227
102 | 20 | 2 | CP2 22may24 BOT |  | 1 | 0 | -6.09 |  | -6.0859 |  |  | -6.1091 |  |  |  | -5.1514
103 | 20 | 3 | CP2 22may24 BOT |  | 1 | 0 | -6.165 |  | -6.1632 |  |  | -6.1867 |  |  |  | -5.2258
104 | 20 | 4 | CP2 22may24 BOT |  | 1 | 0 | -6.159 | 0.047403 | -6.1583 | 0.053224 |  | -6.182 |  |  |  | -5.2213
105 | 21 | 1 | CC4 16apr25 BOT |  | 1 | 0 | -6.304 |  | -6.3172 |  |  | -6.3411 |  |  |  | -5.3741
106 | 21 | 2 | CC4 16apr25 BOT |  | 1 | 0 | -6.292 |  | -6.2955 |  |  | -6.3197 |  |  |  | -5.3535
107 | 21 | 3 | CC4 16apr25 BOT |  | 1 | 0 | -6.257 |  | -6.2592 |  |  | -6.2836 |  |  |  | -5.3189
108 | 21 | 4 | CC4 16apr25 BOT |  | 1 | 0 | -6.358 | 0.041884 | -6.3596 | 0.041985 |  | -6.3842 |  |  |  | -5.4155
109 | 22 | 1 | CP1 22may24 0.1m |  | 1 | 0 | -6.175 |  | -6.1584 |  |  | -6.1833 |  |  |  | -5.2225
110 | 22 | 2 | CP1 22may24 0.1m |  | 1 | 0 | -6.153 |  | -6.1476 |  |  | -6.1727 |  |  |  | -5.2124
111 | 22 | 3 | CP1 22may24 0.1m |  | 1 | 0 | -6.108 |  | -6.1025 |  |  | -6.1278 |  |  |  | -5.1693
112 | 22 | 4 | CP1 22may24 0.1m |  | 1 | 0 | -6.097 | 0.0369 | -6.0949 | 0.031817 |  | -6.1204 |  |  |  | -5.1622
113 | 23 | 1 | CP2 19jun24 0.1m |  | 1 | 0 | -5.69 |  | -5.6531 |  |  | -5.6788 |  |  |  | -4.7382
114 | 23 | 2 | CP2 19jun24 0.1m |  | 1 | 0 | -5.745 |  | -5.7357 |  |  | -5.7617 |  |  |  | -4.8177
115 | 23 | 3 | CP2 19jun24 0.1m |  | 1 | 0 | -5.661 |  | -5.6514 |  |  | -5.6776 |  |  |  | -4.7369
116 | 23 | 4 | CP2 19jun24 0.1m |  | 1 | 0 | -5.737 | 0.039761 | -5.7341 | 0.047741 |  | -5.7605 |  |  |  | -4.8166
117 | 24 | 1 | CS1 19jun24 0.1m |  | 1 | 0 | -7.43 |  | -7.5835 |  |  | -7.6102 |  |  |  | -6.5928
118 | 24 | 2 | CS1 19jun24 0.1m |  | 1 | 0 | -7.584 |  | -7.6327 |  |  | -7.6596 |  |  |  | -6.6403
119 | 24 | 3 | CS1 19jun24 0.1m |  | 1 | 0 | -7.498 |  | -7.5369 |  |  | -7.564 |  |  |  | -6.5485
120 | 24 | 4 | CS1 19jun24 0.1m |  | 1 | 0 | -7.593 | 0.077142 | -7.608 | 0.040871 |  | -7.6353 |  |  |  | -6.617
121 | 25 | 1 | CC5 22may24 9m DUP |  | 1 | 0 | -6.54 |  | -6.4445 |  |  | -6.4721 |  |  |  | -5.4999
122 | 25 | 2 | CC5 22may24 9m DUP |  | 1 | 0 | -6.523 |  | -6.4948 |  |  | -6.5226 |  |  |  | -5.5484
123 | 25 | 3 | CC5 22may24 9m DUP |  | 1 | 0 | -6.439 |  | -6.4135 |  |  | -6.4416 |  |  |  | -5.4706
124 | 25 | 4 | CC5 22may24 9m DUP |  | 1 | 0 | -6.425 | 0.058146 | -6.4156 | 0.037869 |  | -6.4438 |  |  |  | -5.4728
125 | 26 | 1 | Hawaii |  | 1 | 0 | -4.852 |  | -4.7093 |  |  | -4.7378 |  |  |  | -3.8345
126 | 26 | 2 | Hawaii |  | 1 | 0 | -4.711 |  | -4.6658 |  |  | -4.6945 |  |  |  | -3.7929
127 | 26 | 3 | Hawaii |  | 1 | 0 | -4.716 |  | -4.6783 |  |  | -4.7072 |  |  |  | -3.8051
128 | 26 | 4 | Hawaii |  | 1 | 0 | -4.703 | 0.071202 | -4.6891 | 0.018447 |  | -4.7183 |  |  |  | -3.8157
129 | 27 | 1 | Blacksburg | Drift | 1 | 0 | -8.464 |  | -8.8051 |  | -8.8051 | -8.8345 | -8.8345 |  |  | -7.7685
130 | 27 | 2 | Blacksburg | Drift | 1 | 0 | -8.607 |  | -8.7099 |  | -8.7099 | -8.7396 | -8.7396 |  |  | -7.6773
131 | 27 | 3 | Blacksburg | Drift | 1 | 0 | -8.683 |  | -8.7708 |  | -8.7708 | -8.8007 | -8.8007 |  |  | -7.7361
132 | 27 | 4 | Blacksburg | Drift | 1 | 0 | -8.732 | 0.11692 | -8.7645 | 0.039379 | -8.7645 | -8.7946 | -8.7946 |  |  | -7.7302
Note on G2: Manually change first six injections to ignore "-1".
Note on G3: Manually change first six injections to ignore "-1".
Note on G4: Manually change first six injections to ignore "-1".
Note on G5: Manually change first six injections to ignore "-1".
Note on G6: Manually change first six injections to ignore "-1".
Note on G7: Manually change first six injections to ignore "-1".

## 2H
Line | Sample | Injection | Sample Name | Correction Type | Good | Ignore | dD_H) Picarro Raw | Stdev Raw | d(D_H) Memory Corrected | Stdev Memory Corrected | d(D_H) Memory Corrected Drift | d(D_H) Drift Corrected | Drift Corrected Drift | d(D_H) Drift Corrected High, Low, Drift | d(D_H) Reference Defined | dD Final Corrected Value | Post run corrections:
1 | 1 | 1 | Blacksburg | Drift | 1 | -1 | -49.993 | [9 empty] | (1) Memory correction |  |  | (1) Graphs for memory correction
2 | 1 | 2 | Blacksburg | Drift | 1 | -1 | -45.757 | [9 empty] | Combined Stdev | 0.26078
3 | 1 | 3 | Blacksburg | Drift | 1 | -1 | -44.86 | [9 empty] | Date:
4 | 1 | 4 | Blacksburg | Drift | 1 | -1 | -44.233 | [9 empty] | injection | 18O memory coef
5 | 1 | 5 | Blacksburg | Drift | 1 | -1 | -44.072 | [9 empty] | 1 | 0.8427
6 | 1 | 6 | Blacksburg | Drift | 1 | -1 | -44.093 | [9 empty] | 2 | 0.95908
7 | 1 | 7 | Blacksburg | Drift | 1 | 0 | -44.062 |  | -44.062 |  | -44.062 | -44.1094 | -44.1094 |  |  | -48.025 | 3 | 0.98065
8 | 1 | 8 | Blacksburg | Drift | 1 | 0 | -43.927 |  | -43.927 |  | -43.927 | -43.9811 | -43.9811 |  |  | -47.9026 | 4 | 0.98678
9 | 1 | 9 | Blacksburg | Drift | 1 | 0 | -43.805 |  | -43.805 |  | -43.805 | -43.8659 | -43.8659 |  |  | -47.7926 | 5 | 0.99311
10 | 1 | 10 | Blacksburg | Drift | 1 | 0 | -43.968 | 0.10655 | -43.968 | 0.10655 | -43.968 | -44.0357 | -44.0357 |  |  | -47.9546 | 6 | 0.994
11 | 2 | 1 | Myrtle | High Reference | 1 | 0 | -9.741 |  | -4.357 |  |  | -4.4315 |  | -4.4315 | -10.453 | -10.158 | 7 | 0.99589
12 | 2 | 2 | Myrtle | High Reference | 1 | 0 | -5.886 |  | -4.3276 |  |  | -4.4089 |  | -4.4089 | -10.453 | -10.1365 | 8 | 0.99838
13 | 2 | 3 | Myrtle | High Reference | 1 | 0 | -4.837 |  | -4.0796 |  |  | -4.1676 |  | -4.1676 | -10.453 | -9.9062 | 9 | 0.99843
14 | 2 | 4 | Myrtle | High Reference | 1 | 0 | -4.663 |  | -4.1434 |  |  | -4.2381 |  | -4.2381 | -10.453 | -9.9735 | 10 | 1
15 | 2 | 5 | Myrtle | High Reference | 1 | 0 | -4.292 |  | -4.0187 |  |  | -4.1202 |  | -4.1202 | -10.453 | -9.861
16 | 2 | 6 | Myrtle | High Reference | 1 | 0 | -4.4 |  | -4.1627 |  |  | -4.271 |  | -4.271 | -10.453 | -10.0049 | (2) Drift correction
17 | 2 | 7 | Myrtle | High Reference | 1 | 0 | -4.4 |  | -4.2373 |  |  | -4.3524 |  | -4.3524 | -10.453 | -10.0825 | Slope m: | 0.0067682
18 | 2 | 8 | Myrtle | High Reference | 1 | 0 | -4.353 |  | -4.2889 |  |  | -4.4107 |  | -4.4107 | -10.453 | -10.1382 | Stdev Mem Corr. | 0.3453
19 | 2 | 9 | Myrtle | High Reference | 1 | 0 | -4.001 |  | -3.9381 |  |  | -4.0667 |  | -4.0667 | -10.453 | -9.8099 | Stdev Drift Corr. | 0.14051
20 | 2 | 10 | Myrtle | High Reference | 1 | 0 | -4.242 | 1.7175 | -4.242 | 0.13677 |  | -4.3774 |  | -4.3774 | -10.453 | -10.1064
21 | 3 | 1 | Homer | Low Reference | 1 | 0 | -94.703 |  | -108.9327 |  |  | -109.0749 |  | -109.0749 | -110.5169 | -110.0253 | (3) Normalization to VSMOW scale
22 | 3 | 2 | Homer | Low Reference | 1 | 0 | -105.109 |  | -109.2366 |  |  | -109.3855 |  | -109.3855 | -110.5169 | -110.3218 | slope m | 0.95436
23 | 3 | 3 | Homer | Low Reference | 1 | 0 | -107.067 |  | -109.0571 |  |  | -109.2128 |  | -109.2128 | -110.5169 | -110.1569 | interception t | -5.9288
24 | 3 | 4 | Homer | Low Reference | 1 | 0 | -107.801 |  | -109.17 |  |  | -109.3325 |  | -109.3325 | -110.5169 | -110.2711
25 | 3 | 5 | Homer | Low Reference | 1 | 0 | -108.31 |  | -109.0268 |  |  | -109.196 |  | -109.196 | -110.5169 | -110.1409 | (4) Quality control |  | offset
26 | 3 | 6 | Homer | Low Reference | 1 | 0 | -108.557 |  | -109.1827 |  |  | -109.3586 |  | -109.3586 | -110.5169 | -110.2961 | QC Check Defined | -13.2294 | accepted ±1.5
27 | 3 | 7 | Homer | Low Reference | 1 | 0 | -108.687 |  | -109.1165 |  |  | -109.2992 |  | -109.2992 | -110.5169 | -110.2394 | QC Check Vial 5 | -13.5588 | -0.32941
28 | 3 | 8 | Homer | Low Reference | 1 | 0 | -109.189 |  | -109.3589 |  |  | -109.5484 |  | -109.5484 | -110.5169 | -110.4772 | QC Check Vial 26 | -13.3309 | -0.10147
29 | 3 | 9 | Homer | Low Reference | 1 | 0 | -108.907 |  | -109.0717 |  |  | -109.268 |  | -109.268 | -110.5169 | -110.2096
30 | 3 | 10 | Homer | Low Reference | 1 | 0 | -109.21 | 4.4112 | -109.21 | 0.12137 |  | -109.413 |  | -109.413 | -110.5169 | -110.348 | Drift Defined | -47.446 | accepted ±1.5
31 | 4 | 1 | Blacksburg | Drift | 1 | 0 | -52.491 |  | -43.569 |  |  | -43.7788 |  | -43.7788 | -47.446 | -47.7095 | Drift Average | -47.9915 | -0.54551
32 | 4 | 2 | Blacksburg | Drift | 1 | 0 | -46.633 |  | -44.0723 |  |  | -44.2889 |  | -44.2889 | -47.446 | -48.1963 | Drift Vial 4 | -48.2488 | -0.8028
33 | 4 | 3 | Blacksburg | Drift | 1 | 0 | -45.174 |  | -43.9346 |  |  | -44.158 |  | -44.158 | -47.446 | -48.0713 | Drift Vial 9 | -47.9594 | -0.51335
34 | 4 | 4 | Blacksburg | Drift | 1 | 0 | -44.927 |  | -44.0772 |  |  | -44.3073 |  | -44.3073 | -47.446 | -48.2139 | Drift Vial 18 | -47.8895 | -0.4435
35 | 4 | 5 | Blacksburg | Drift | 1 | 0 | -44.37 |  | -43.9234 |  |  | -44.1603 |  | -44.1603 | -47.446 | -48.0735 | Drift Vial 27 | -47.9412 | -0.49519
36 | 4 | 6 | Blacksburg | Drift | 1 | 0 | -44.474 |  | -44.0857 |  |  | -44.3294 |  | -44.3294 | -47.446 | -48.2349
37 | 4 | 7 | Blacksburg | Drift | 1 | 0 | -44.201 |  | -43.9337 |  |  | -44.1841 |  | -44.1841 | -47.446 | -48.0963 | Reference Defined Values |  | d2H
38 | 4 | 8 | Blacksburg | Drift | 1 | 0 | -44.398 |  | -44.2931 |  |  | -44.5503 |  | -44.5503 | -47.446 | -48.4458 | High Reference | North Myrtle Beach, SC | -10.453
39 | 4 | 9 | Blacksburg | Drift | 1 | 0 | -44.137 |  | -44.0346 |  |  | -44.2986 |  | -44.2986 | -47.446 | -48.2055 | Low Reference | Homer, AK | -110.5169
40 | 4 | 10 | Blacksburg | Drift | 1 | 0 | -44.072 | 2.5759 | -44.072 | 0.18594 |  | -44.3427 |  | -44.3427 | -47.446 | -48.2477 | Drift | Blacksburg | -47.446
41 | 5 | 1 | Hawaii | QC Check | 1 | 0 | -12.658 |  | -7.7165 |  |  | -7.994 |  |  |  | -13.558 | QC Check | Hawaii | -13.2294
42 | 5 | 2 | Hawaii | QC Check | 1 | 0 | -9.388 |  | -7.9687 |  |  | -8.253 |  |  |  | -13.8051 | Calibration Date:
43 | 5 | 3 | Hawaii | QC Check | 1 | 0 | -8.253 |  | -7.5597 |  |  | -7.8508 |  |  |  | -13.4213
44 | 5 | 4 | Hawaii | QC Check | 1 | 0 | -8.06 | 2.1277 | -7.5839 | 0.18744 |  | -7.8817 |  |  |  | -13.4508
45 | 6 | 1 | C50 22may24 1.5m |  | 1 | 0 | -23.486 |  | -25.9125 |  |  | -26.2171 |  |  |  | -30.9493
46 | 6 | 2 | C50 22may24 1.5m |  | 1 | 0 | -25.544 |  | -26.2595 |  |  | -26.5708 |  |  |  | -31.2869 | Denali National Park | -153.7172
47 | 6 | 3 | C50 22may24 1.5m |  | 1 | 0 | -25.756 |  | -26.0985 |  |  | -26.4166 |  |  |  | -31.1397 | Edmonton, Alberta | -146.8245
48 | 6 | 4 | C50 22may24 1.5m |  | 1 | 0 | -25.881 | 1.1291 | -26.1166 | 0.14237 |  | -26.4415 |  |  |  | -31.1635 | MT BHB | -140.5778
49 | 7 | 1 | CC3 19jun24 0.1m |  | 1 | 0 | -24.078 |  | -23.7944 |  |  | -24.126 |  |  |  | -28.9537 | Anchorage, AK | -129.9028
50 | 7 | 2 | CC3 19jun24 0.1m |  | 1 | 0 | -23.83 |  | -23.7461 |  |  | -24.0845 |  |  |  | -28.9141 | Homer, AK | -110.5169
51 | 7 | 3 | CC3 19jun24 0.1m |  | 1 | 0 | -23.905 |  | -23.8668 |  |  | -24.2119 |  |  |  | -29.0357 | MT Creek | -97.7428
52 | 7 | 4 | CC3 19jun24 0.1m |  | 1 | 0 | -23.849 | 0.11291 | -23.8221 | 0.050567 |  | -24.1741 |  |  |  | -28.9996 | Nickel Mines, PA | -48.7951
53 | 8 | 1 | C50 16sep24 6m |  | 1 | 0 | -20.541 |  | -20.0206 |  |  | -20.3794 |  |  |  | -25.378 | Blacksburg | -47.446
54 | 8 | 2 | C50 16sep24 6m |  | 1 | 0 | -20.249 |  | -20.1017 |  |  | -20.4672 |  |  |  | -25.4618 | Denton's Well | -41.771
55 | 8 | 3 | C50 16sep24 6m |  | 1 | 0 | -19.427 |  | -19.3414 |  |  | -19.7137 |  |  |  | -24.7427 | Hawaii | -13.2294
56 | 8 | 4 | C50 16sep24 6m |  | 1 | 0 | -19.78 | 0.49376 | -19.7262 | 0.3442 |  | -20.1052 |  |  |  | -25.1164 | North Myrtle Beach, SC | -10.453
57 | 9 | 1 | Blacksburg | Drift | 1 | 0 | -40.483 |  | -43.7396 |  | -43.7396 | -44.1254 | -44.1254 |  |  | -48.0403 | WQL Heavy Oxygen | -15.7881
58 | 9 | 2 | Blacksburg | Drift | 1 | 0 | -42.795 |  | -43.7368 |  | -43.7368 | -44.1294 | -44.1294 |  |  | -48.044 | WQL Heavy Hydrogen | 4.0332
59 | 9 | 3 | Blacksburg | Drift | 1 | 0 | -43.076 |  | -43.5269 |  | -43.5269 | -43.9262 | -43.9262 |  |  | -47.8502 | Andres Heavy Water | 20.5382
60 | 9 | 4 | Blacksburg | Drift | 1 | 0 | -43.265 | 1.2956 | -43.5755 | 0.1098 | -43.5755 | -43.9816 | -43.9816 |  |  | -47.903
61 | 10 | 1 | CC5 22may24 1.5m |  | 1 | 0 | -28.182 |  | -25.8094 |  |  | -26.2223 |  |  |  | -30.9543
62 | 10 | 2 | CC5 22may24 1.5m |  | 1 | 0 | -26.671 |  | -25.992 |  |  | -26.4116 |  |  |  | -31.1349
63 | 10 | 3 | CC5 22may24 1.5m |  | 1 | 0 | -25.878 |  | -25.5415 |  |  | -25.9679 |  |  |  | -30.7115
64 | 10 | 4 | CC5 22may24 1.5m |  | 1 | 0 | -25.92 | 1.0764 | -25.6907 | 0.19042 |  | -26.1239 |  |  |  | -30.8604
65 | 11 | 1 | CS2 19jun24 0.1m |  | 1 | 0 | -31.022 |  | -31.8246 |  |  | -32.2645 |  |  |  | -36.7207
66 | 11 | 2 | CS2 19jun24 0.1m |  | 1 | 0 | -31.66 |  | -31.8949 |  |  | -32.3416 |  |  |  | -36.7943
67 | 11 | 3 | CS2 19jun24 0.1m |  | 1 | 0 | -31.797 |  | -31.9107 |  |  | -32.3642 |  |  |  | -36.8159
68 | 11 | 4 | CS2 19jun24 0.1m |  | 1 | 0 | -31.793 | 0.36952 | -31.8706 | 0.037579 |  | -32.3309 |  |  |  | -36.7841
69 | 12 | 1 | CC4 19jun24 9m |  | 1 | 0 | -27.677 |  | -27.0295 |  |  | -27.4966 |  |  |  | -32.1704
70 | 12 | 2 | CC4 19jun24 9m |  | 1 | 0 | -27.098 |  | -26.9059 |  |  | -27.3797 |  |  |  | -32.0588
71 | 12 | 3 | CC4 19jun24 9m |  | 1 | 0 | -27.359 |  | -27.2732 |  |  | -27.7537 |  |  |  | -32.4158
72 | 12 | 4 | CC4 19jun24 9m |  | 1 | 0 | -27.191 | 0.25455 | -27.1302 | 0.15558 |  | -27.6175 |  |  |  | -32.2858
73 | 13 | 1 | CC5 22may24 6m |  | 1 | 0 | -26.504 |  | -26.3959 |  |  | -26.89 |  |  |  | -31.5915
74 | 13 | 2 | CC5 22may24 6m |  | 1 | 0 | -26.236 |  | -26.1969 |  |  | -26.6978 |  |  |  | -31.4081
75 | 13 | 3 | CC5 22may24 6m |  | 1 | 0 | -26.334 |  | -26.3174 |  |  | -26.825 |  |  |  | -31.5295
76 | 13 | 4 | CC5 22may24 6m |  | 1 | 0 | -26.274 | 0.11842 | -26.2619 | 0.084442 |  | -26.7763 |  |  |  | -31.483
77 | 14 | 1 | C50 16apr25 1.5m |  | 1 | 0 | -26.06 |  | -26.0263 |  |  | -26.5475 |  |  |  | -31.2646
78 | 14 | 2 | C50 16apr25 1.5m |  | 1 | 0 | -25.595 |  | -25.5672 |  |  | -26.0951 |  |  |  | -30.8329
79 | 14 | 3 | C50 16apr25 1.5m |  | 1 | 0 | -25.667 |  | -25.6553 |  |  | -26.1899 |  |  |  | -30.9234
80 | 14 | 4 | C50 16apr25 1.5m |  | 1 | 0 | -25.893 | 0.21284 | -25.888 | 0.21064 |  | -26.4294 |  |  |  | -31.152
81 | 15 | 1 | CC5 22may24 0.1m |  | 1 | 0 | -25.389 |  | -25.3097 |  |  | -25.8579 |  |  |  | -30.6066
82 | 15 | 2 | CC5 22may24 0.1m |  | 1 | 0 | -25.457 |  | -25.4392 |  |  | -25.9942 |  |  |  | -30.7366
83 | 15 | 3 | CC5 22may24 0.1m |  | 1 | 0 | -25.527 |  | -25.5199 |  |  | -26.0817 |  |  |  | -30.8201
84 | 15 | 4 | CC5 22may24 0.1m |  | 1 | 0 | -25.755 | 0.15898 | -25.7532 | 0.18644 |  | -26.3217 |  |  |  | -31.0492
85 | 16 | 1 | CC5 22may24 9m |  | 1 | 0 | -26.89 |  | -27.0685 |  |  | -27.6438 |  |  |  | -32.311
86 | 16 | 2 | CC5 22may24 9m |  | 1 | 0 | -27.399 |  | -27.4663 |  |  | -28.0483 |  |  |  | -32.697
87 | 16 | 3 | CC5 22may24 9m |  | 1 | 0 | -27.199 |  | -27.2269 |  |  | -27.8158 |  |  |  | -32.475
88 | 16 | 4 | CC5 22may24 9m |  | 1 | 0 | -27.289 | 0.2187 | -27.3093 | 0.16582 |  | -27.9049 |  |  |  | -32.5601
89 | 17 | 1 | CC4 22may24 BOT |  | 1 | 0 | -26.762 |  | -26.6791 |  |  | -27.2815 |  |  |  | -31.9651
90 | 17 | 2 | CC4 22may24 BOT |  | 1 | 0 | -26.968 |  | -26.9549 |  |  | -27.564 |  |  |  | -32.2348
91 | 17 | 3 | CC4 22may24 BOT |  | 1 | 0 | -26.835 |  | -26.8262 |  |  | -27.4421 |  |  |  | -32.1184
92 | 17 | 4 | CC4 22may24 BOT |  | 1 | 0 | -27.023 | 0.1197 | -27.0195 | 0.15045 |  | -27.6422 |  |  |  | -32.3094
93 | 18 | 1 | Blacksburg | Drift | 1 | 0 | -41.252 |  | -43.4903 |  | -43.4903 | -44.1197 | -44.1197 |  |  | -48.0348
94 | 18 | 2 | Blacksburg | Drift | 1 | 0 | -42.606 |  | -43.2437 |  | -43.2437 | -43.8799 | -43.8799 |  |  | -47.806
95 | 18 | 3 | Blacksburg | Drift | 1 | 0 | -43.048 |  | -43.3582 |  | -43.3582 | -44.0011 | -44.0011 |  |  | -47.9217
96 | 18 | 4 | Blacksburg | Drift | 1 | 0 | -43.008 | 0.84168 | -43.2193 | 0.12404 | -43.2193 | -43.8691 | -43.8691 |  |  | -47.7956
97 | 19 | 1 | CC3 16apr25 1.5m |  | 1 | 0 | -28.26 |  | -25.9401 |  |  | -26.5966 |  |  |  | -31.3115
98 | 19 | 2 | CC3 16apr25 1.5m |  | 1 | 0 | -26.284 |  | -25.5996 |  |  | -26.2629 |  |  |  | -30.9931
99 | 19 | 3 | CC3 16apr25 1.5m |  | 1 | 0 | -26.089 |  | -25.7615 |  |  | -26.4316 |  |  |  | -31.154
100 | 19 | 4 | CC3 16apr25 1.5m |  | 1 | 0 | -25.611 | 1.1674 | -25.381 | 0.23791 |  | -26.0578 |  |  |  | -30.7973
101 | 20 | 1 | CP2 22may24 BOT |  | 1 | 0 | -25.602 |  | -25.6006 |  |  | -26.2842 |  |  |  | -31.0133
102 | 20 | 2 | CP2 22may24 BOT |  | 1 | 0 | -25.546 |  | -25.5433 |  |  | -26.2337 |  |  |  | -30.9652
103 | 20 | 3 | CP2 22may24 BOT |  | 1 | 0 | -25.12 |  | -25.1105 |  |  | -25.8076 |  |  |  | -30.5585
104 | 20 | 4 | CP2 22may24 BOT |  | 1 | 0 | -25.197 | 0.24302 | -25.1915 | 0.24639 |  | -25.8954 |  |  |  | -30.6423
105 | 21 | 1 | CC4 16apr25 BOT |  | 1 | 0 | -25.871 |  | -25.977 |  |  | -26.6877 |  |  |  | -31.3984
106 | 21 | 2 | CC4 16apr25 BOT |  | 1 | 0 | -26.169 |  | -26.2088 |  |  | -26.9262 |  |  |  | -31.6261
107 | 21 | 3 | CC4 16apr25 BOT |  | 1 | 0 | -26.151 |  | -26.1695 |  |  | -26.8937 |  |  |  | -31.595
108 | 21 | 4 | CC4 16apr25 BOT |  | 1 | 0 | -25.944 | 0.14898 | -25.9539 | 0.13047 |  | -26.6848 |  |  |  | -31.3957
109 | 22 | 1 | CP1 22may24 0.1m |  | 1 | 0 | -25.578 |  | -25.5204 |  |  | -26.2582 |  |  |  | -30.9885
110 | 22 | 2 | CP1 22may24 0.1m |  | 1 | 0 | -25.504 |  | -25.486 |  |  | -26.2305 |  |  |  | -30.9621
111 | 22 | 3 | CP1 22may24 0.1m |  | 1 | 0 | -25.547 |  | -25.5393 |  |  | -26.2906 |  |  |  | -31.0195
112 | 22 | 4 | CP1 22may24 0.1m |  | 1 | 0 | -25.485 | 0.041972 | -25.4789 | 0.028583 |  | -26.237 |  |  |  | -30.9683
113 | 23 | 1 | CP2 19jun24 0.1m |  | 1 | 0 | -23.711 |  | -23.4319 |  |  | -24.1968 |  |  |  | -29.0212
114 | 23 | 2 | CP2 19jun24 0.1m |  | 1 | 0 | -23.247 |  | -23.1554 |  |  | -23.927 |  |  |  | -28.7638
115 | 23 | 3 | CP2 19jun24 0.1m |  | 1 | 0 | -23.486 |  | -23.4473 |  |  | -24.2257 |  |  |  | -29.0488
116 | 23 | 4 | CP2 19jun24 0.1m |  | 1 | 0 | -23.171 | 0.24489 | -23.1404 | 0.16865 |  | -23.9255 |  |  |  | -28.7624
117 | 24 | 1 | CS1 19jun24 0.1m |  | 1 | 0 | -30.362 |  | -31.4932 |  |  | -32.285 |  |  |  | -36.7403
118 | 24 | 2 | CS1 19jun24 0.1m |  | 1 | 0 | -31.308 |  | -31.641 |  |  | -32.4396 |  |  |  | -36.8879
119 | 24 | 3 | CS1 19jun24 0.1m |  | 1 | 0 | -31.134 |  | -31.2881 |  |  | -32.0935 |  |  |  | -36.5576
120 | 24 | 4 | CS1 19jun24 0.1m |  | 1 | 0 | -31.363 | 0.46356 | -31.4713 | 0.14469 |  | -32.2835 |  |  |  | -36.7388
121 | 25 | 1 | CC5 22may24 9m DUP |  | 1 | 0 | -27.115 |  | -26.4468 |  |  | -27.2657 |  |  |  | -31.9501
122 | 25 | 2 | CC5 22may24 9m DUP |  | 1 | 0 | -26.764 |  | -26.5758 |  |  | -27.4015 |  |  |  | -32.0797
123 | 25 | 3 | CC5 22may24 9m DUP |  | 1 | 0 | -26.992 |  | -26.9074 |  |  | -27.7399 |  |  |  | -32.4026
124 | 25 | 4 | CC5 22may24 9m DUP |  | 1 | 0 | -27.064 | 0.15495 | -27.0072 | 0.26596 |  | -27.8464 |  |  |  | -32.5043
125 | 26 | 1 | Hawaii |  | 1 | 0 | -9.663 |  | -6.9258 |  |  | -7.7718 |  |  |  | -13.3459
126 | 26 | 2 | Hawaii |  | 1 | 0 | -7.801 |  | -7.0127 |  |  | -7.8655 |  |  |  | -13.4354
127 | 26 | 3 | Hawaii |  | 1 | 0 | -7.216 |  | -6.8319 |  |  | -7.6914 |  |  |  | -13.2692
128 | 26 | 4 | Hawaii |  | 1 | 0 | -7.093 | 1.1874 | -6.829 | 0.087675 |  | -7.6953 |  |  |  | -13.2729
129 | 27 | 1 | Blacksburg | Drift | 1 | 0 | -38.288 |  | -43.1951 |  | -43.1951 | -44.0682 | -44.0682 |  |  | -47.9856
130 | 27 | 2 | Blacksburg | Drift | 1 | 0 | -41.39 |  | -42.7935 |  | -42.7935 | -43.6733 | -43.6733 |  |  | -47.6088
131 | 27 | 3 | Blacksburg | Drift | 1 | 0 | -42.528 |  | -43.2138 |  | -43.2138 | -44.1005 | -44.1005 |  |  | -48.0165
132 | 27 | 4 | Blacksburg | Drift | 1 | 0 | -42.878 | 2.0877 | -43.3511 | 0.24021 | -43.3511 | -44.2445 | -44.2445 |  |  | -48.1539
Note on G2: Manually change first six injections to ignore "-1".
Note on G3: Manually change first six injections to ignore "-1".
Note on G4: Manually change first six injections to ignore "-1".
Note on G5: Manually change first six injections to ignore "-1".
Note on G6: Manually change first six injections to ignore "-1".
Note on G7: Manually change first six injections to ignore "-1".

## RESULTS
| Analyst: | DWH |  |  |  | Complete Sequence |  |  | Run Summary | 18O | 18O | 2H | 2H
| Analysis Date: | 2025-05-29 |  |  |  | Line | Inj Number | Identifier | Ignore | Ignore | d18O Final Value | d2H Final Value | Ignore
| Project: | CCR |  |  |  | 1 | 1 | Blacksburg | 0 | -1 |  |  | -1
[6 empty] | 2 | 2 | Blacksburg | 0 | -1 |  |  | -1
[6 empty] | 3 | 3 | Blacksburg | 0 | -1 |  |  | -1
Vial | Sample Name | d18O | 1 s.d. | d2H | 1 s.d. | 4 | 4 | Blacksburg | 0 | -1 |  |  | -1
|  | ‰ vs. VSMOW |  | ‰ vs. VSMOW |  | 5 | 5 | Blacksburg | 0 | -1 |  |  | -1
1 | Blacksburg | -7.7368 | 0.072019 | -47.9187 | 0.097879 | 6 | 6 | Blacksburg | 0 | -1 |  |  | -1
2 | Myrtle | -2.295 | 0.043315 | -10.0177 | 0.1259 | 7 | 7 | Blacksburg | 0 | 0 | -7.7753 | -48.025 | 0
3 | Homer | -14.5427 | 0.02153 | -110.2486 | 0.12549 | 8 | 8 | Blacksburg | 0 | 0 | -7.788 | -47.9026 | 0
4 | Blacksburg | -7.5546 | 0.023062 | -48.1495 | 0.1893 | 9 | 9 | Blacksburg | 0 | 0 | -7.7527 | -47.7926 | 0
5 | Hawaii | -3.6566 | 0.030776 | -13.5588 | 0.17439 | 10 | 10 | Blacksburg | 0 | 0 | -7.631 | -47.9546 | 0
6 | C50 22may24 1.5m | -5.114 | 0.068056 | -31.1349 | 0.13949 | 11 | 1 | Myrtle | 0 | 0 | -2.3209 | -10.158 | 0
7 | CC3 19jun24 0.1m | -4.6937 | 0.035113 | -28.9758 | 0.05308 | 12 | 2 | Myrtle | 0 | 0 | -2.3633 | -10.1365 | 0
8 | C50 16sep24 6m | -3.761 | 0.01756 | -25.1748 | 0.32341 | 13 | 3 | Myrtle | 0 | 0 | -2.308 | -9.9062 | 0
9 | Blacksburg | -7.6195 | 0.044507 | -47.9594 | 0.098016 | 14 | 4 | Myrtle | 0 | 0 | -2.293 | -9.9735 | 0
10 | CC5 22may24 1.5m | -5.0236 | 0.072391 | -30.9153 | 0.17731 | 15 | 5 | Myrtle | 0 | 0 | -2.3149 | -9.861 | 0
11 | CS2 19jun24 0.1m | -6.4626 | 0.045786 | -36.7787 | 0.040895 | 16 | 6 | Myrtle | 0 | 0 | -2.2587 | -10.0049 | 0
12 | CC4 19jun24 9m | -5.2921 | 0.043179 | -32.2327 | 0.15326 | 17 | 7 | Myrtle | 0 | 0 | -2.2917 | -10.0825 | 0
13 | CC5 22may24 6m | -5.2889 | 0.02239 | -31.503 | 0.077363 | 18 | 8 | Myrtle | 0 | 0 | -2.3093 | -10.1382 | 0
14 | C50 16apr25 1.5m | -5.2202 | 0.067644 | -31.0432 | 0.19952 | 19 | 9 | Myrtle | 0 | 0 | -2.2921 | -9.8099 | 0
15 | CC5 22may24 0.1m | -5.0608 | 0.050505 | -30.8031 | 0.18609 | 20 | 10 | Myrtle | 0 | 0 | -2.1979 | -10.1064 | 0
16 | CC5 22may24 9m | -5.4324 | 0.041472 | -32.5108 | 0.16157 | 21 | 1 | Homer | 0 | 0 | -14.5563 | -110.0253 | 0
17 | CC4 22may24 BOT | -5.3397 | 0.018897 | -32.1569 | 0.15007 | 22 | 2 | Homer | 0 | 0 | -14.552 | -110.3218 | 0
18 | Blacksburg | -7.6838 | 0.024717 | -47.8895 | 0.11246 | 23 | 3 | Homer | 0 | 0 | -14.5634 | -110.1569 | 0
19 | CC3 16apr25 1.5m | -5.2412 | 0.03697 | -31.064 | 0.22024 | 24 | 4 | Homer | 0 | 0 | -14.5385 | -110.2711 | 0
20 | CP2 22may24 BOT | -5.1803 | 0.051374 | -30.7949 | 0.22793 | 25 | 5 | Homer | 0 | 0 | -14.5467 | -110.1409 | 0
21 | CC4 16apr25 BOT | -5.3655 | 0.040398 | -31.5038 | 0.1239 | 26 | 6 | Homer | 0 | 0 | -14.5777 | -110.2961 | 0
22 | CP1 22may24 0.1m | -5.1916 | 0.030284 | -30.9846 | 0.025834 | 27 | 7 | Homer | 0 | 0 | -14.5161 | -110.2394 | 0
23 | CP2 19jun24 0.1m | -4.7774 | 0.045968 | -28.899 | 0.15741 | 28 | 8 | Homer | 0 | 0 | -14.5231 | -110.4772 | 0
24 | CS1 19jun24 0.1m | -6.5996 | 0.039229 | -36.7311 | 0.13519 | 29 | 9 | Homer | 0 | 0 | -14.5436 | -110.2096 | 0
25 | CC5 22may24 9m DUP | -5.4979 | 0.036205 | -32.2342 | 0.26198 | 30 | 10 | Homer | 0 | 0 | -14.5092 | -110.348 | 0
26 | Hawaii | -3.8121 | 0.017621 | -13.3309 | 0.078118 | 31 | 1 | Blacksburg | 0 | 0 | -7.5589 | -47.7095 | 0
27 | Blacksburg | -7.728 | 0.037761 | -47.9412 | 0.23334 | 32 | 2 | Blacksburg | 0 | 0 | -7.5196 | -48.1963 | 0
[6 empty] | 33 | 3 | Blacksburg | 0 | 0 | -7.5815 | -48.0713 | 0
[6 empty] | 34 | 4 | Blacksburg | 0 | 0 | -7.5488 | -48.2139 | 0
[6 empty] | 35 | 5 | Blacksburg | 0 | 0 | -7.5465 | -48.0735 | 0
| Quality Assurance Verification |  |  |  |  | 36 | 6 | Blacksburg | 0 | 0 | -7.5578 | -48.2349 | 0
| Metric |  | Reviewer 1 | Reviewer 2 |  | 37 | 7 | Blacksburg | 0 | 0 | -7.5754 | -48.0963 | 0
| d18O Cal Curve R^2 | 0.999+ | pass |  |  | 38 | 8 | Blacksburg | 0 | 0 | -7.5138 | -48.4458 | 0
| d18O QC checks | < 0.15 | pass |  |  | 39 | 9 | Blacksburg | 0 | 0 | -7.5758 | -48.2055 | 0
| dD Cal Curve R^2 | 0.999+ | pass |  |  | 40 | 10 | Blacksburg | 0 | 0 | -7.5675 | -48.2477 | 0
| dD QC checks | < 1.5 | pass |  |  | 41 | 1 | Hawaii | 0 | 0 | -3.6504 | -13.558 | 0
| Initials |  | DWH |  |  | 42 | 2 | Hawaii | 0 | 0 | -3.6274 | -13.8051 | 0
| Date |  | 2025-06-03 |  |  | 43 | 3 | Hawaii | 0 | 0 | -3.7001 | -13.4213 | 0
| Comments: |  |  |  |  | 44 | 4 | Hawaii | 0 | 0 | -3.6487 | -13.4508 | 0
[6 empty] | 45 | 1 | C50 22may24 1.5m | 0 | 0 | -5.2116 | -30.9493 | 0
[6 empty] | 46 | 2 | C50 22may24 1.5m | 0 | 0 | -5.1057 | -31.2869 | 0
[6 empty] | 47 | 3 | C50 22may24 1.5m | 0 | 0 | -5.0823 | -31.1397 | 0
[6 empty] | 48 | 4 | C50 22may24 1.5m | 0 | 0 | -5.0566 | -31.1635 | 0
[6 empty] | 49 | 1 | CC3 19jun24 0.1m | 0 | 0 | -4.6771 | -28.9537 | 0
[6 empty] | 50 | 2 | CC3 19jun24 0.1m | 0 | 0 | -4.6734 | -28.9141 | 0
[6 empty] | 51 | 3 | CC3 19jun24 0.1m | 0 | 0 | -4.6781 | -29.0357 | 0
[6 empty] | 52 | 4 | CC3 19jun24 0.1m | 0 | 0 | -4.7463 | -28.9996 | 0
[6 empty] | 53 | 1 | C50 16sep24 6m | 0 | 0 | -3.7393 | -25.378 | 0
[6 empty] | 54 | 2 | C50 16sep24 6m | 0 | 0 | -3.7724 | -25.4618 | 0
[6 empty] | 55 | 3 | C50 16sep24 6m | 0 | 0 | -3.7777 | -24.7427 | 0
[6 empty] | 56 | 4 | C50 16sep24 6m | 0 | 0 | -3.7544 | -25.1164 | 0
[6 empty] | 57 | 1 | Blacksburg | 0 | 0 | -7.6013 | -48.0403 | 0
[6 empty] | 58 | 2 | Blacksburg | 0 | 0 | -7.5929 | -48.044 | 0
[6 empty] | 59 | 3 | Blacksburg | 0 | 0 | -7.6861 | -47.8502 | 0
[6 empty] | 60 | 4 | Blacksburg | 0 | 0 | -7.5977 | -47.903 | 0
[6 empty] | 61 | 1 | CC5 22may24 1.5m | 0 | 0 | -5.0252 | -30.9543 | 0
[6 empty] | 62 | 2 | CC5 22may24 1.5m | 0 | 0 | -5.0315 | -31.1349 | 0
[6 empty] | 63 | 3 | CC5 22may24 1.5m | 0 | 0 | -4.9304 | -30.7115 | 0
[6 empty] | 64 | 4 | CC5 22may24 1.5m | 0 | 0 | -5.1071 | -30.8604 | 0
[6 empty] | 65 | 1 | CS2 19jun24 0.1m | 0 | 0 | -6.4667 | -36.7207 | 0
[6 empty] | 66 | 2 | CS2 19jun24 0.1m | 0 | 0 | -6.4007 | -36.7943 | 0
[6 empty] | 67 | 3 | CS2 19jun24 0.1m | 0 | 0 | -6.4721 | -36.8159 | 0
[6 empty] | 68 | 4 | CS2 19jun24 0.1m | 0 | 0 | -6.511 | -36.7841 | 0
[6 empty] | 69 | 1 | CC4 19jun24 9m | 0 | 0 | -5.2653 | -32.1704 | 0
[6 empty] | 70 | 2 | CC4 19jun24 9m | 0 | 0 | -5.3157 | -32.0588 | 0
[6 empty] | 71 | 3 | CC4 19jun24 9m | 0 | 0 | -5.2473 | -32.4158 | 0
[6 empty] | 72 | 4 | CC4 19jun24 9m | 0 | 0 | -5.3401 | -32.2858 | 0
[6 empty] | 73 | 1 | CC5 22may24 6m | 0 | 0 | -5.2701 | -31.5915 | 0
[6 empty] | 74 | 2 | CC5 22may24 6m | 0 | 0 | -5.273 | -31.4081 | 0
[6 empty] | 75 | 3 | CC5 22may24 6m | 0 | 0 | -5.2942 | -31.5295 | 0
[6 empty] | 76 | 4 | CC5 22may24 6m | 0 | 0 | -5.3184 | -31.483 | 0
[6 empty] | 77 | 1 | C50 16apr25 1.5m | 0 | 0 | -5.2434 | -31.2646 | 0
[6 empty] | 78 | 2 | C50 16apr25 1.5m | 0 | 0 | -5.2885 | -30.8329 | 0
[6 empty] | 79 | 3 | C50 16apr25 1.5m | 0 | 0 | -5.1279 | -30.9234 | 0
[6 empty] | 80 | 4 | C50 16apr25 1.5m | 0 | 0 | -5.2208 | -31.152 | 0
[6 empty] | 81 | 1 | CC5 22may24 0.1m | 0 | 0 | -5.0374 | -30.6066 | 0
[6 empty] | 82 | 2 | CC5 22may24 0.1m | 0 | 0 | -5.014 | -30.7366 | 0
[6 empty] | 83 | 3 | CC5 22may24 0.1m | 0 | 0 | -5.1309 | -30.8201 | 0
[6 empty] | 84 | 4 | CC5 22may24 0.1m | 0 | 0 | -5.0608 | -31.0492 | 0
[6 empty] | 85 | 1 | CC5 22may24 9m | 0 | 0 | -5.4352 | -32.311 | 0
[6 empty] | 86 | 2 | CC5 22may24 9m | 0 | 0 | -5.4597 | -32.697 | 0
[6 empty] | 87 | 3 | CC5 22may24 9m | 0 | 0 | -5.3729 | -32.475 | 0
[6 empty] | 88 | 4 | CC5 22may24 9m | 0 | 0 | -5.4618 | -32.5601 | 0
[6 empty] | 89 | 1 | CC4 22may24 BOT | 0 | 0 | -5.3624 | -31.9651 | 0
[6 empty] | 90 | 2 | CC4 22may24 BOT | 0 | 0 | -5.3348 | -32.2348 | 0
[6 empty] | 91 | 3 | CC4 22may24 BOT | 0 | 0 | -5.3444 | -32.1184 | 0
[6 empty] | 92 | 4 | CC4 22may24 BOT | 0 | 0 | -5.3172 | -32.3094 | 0
[6 empty] | 93 | 1 | Blacksburg | 0 | 0 | -7.6867 | -48.0348 | 0
[6 empty] | 94 | 2 | Blacksburg | 0 | 0 | -7.7069 | -47.806 | 0
[6 empty] | 95 | 3 | Blacksburg | 0 | 0 | -7.649 | -47.9217 | 0
[6 empty] | 96 | 4 | Blacksburg | 0 | 0 | -7.6928 | -47.7956 | 0
[6 empty] | 97 | 1 | CC3 16apr25 1.5m | 0 | 0 | -5.2148 | -31.3115 | 0
[6 empty] | 98 | 2 | CC3 16apr25 1.5m | 0 | 0 | -5.2575 | -30.9931 | 0
[6 empty] | 99 | 3 | CC3 16apr25 1.5m | 0 | 0 | -5.207 | -31.154 | 0
[6 empty] | 100 | 4 | CC3 16apr25 1.5m | 0 | 0 | -5.2855 | -30.7973 | 0
[6 empty] | 101 | 1 | CP2 22may24 BOT | 0 | 0 | -5.1227 | -31.0133 | 0
[6 empty] | 102 | 2 | CP2 22may24 BOT | 0 | 0 | -5.1514 | -30.9652 | 0
[6 empty] | 103 | 3 | CP2 22may24 BOT | 0 | 0 | -5.2258 | -30.5585 | 0
[6 empty] | 104 | 4 | CP2 22may24 BOT | 0 | 0 | -5.2213 | -30.6423 | 0
[6 empty] | 105 | 1 | CC4 16apr25 BOT | 0 | 0 | -5.3741 | -31.3984 | 0
[6 empty] | 106 | 2 | CC4 16apr25 BOT | 0 | 0 | -5.3535 | -31.6261 | 0
[6 empty] | 107 | 3 | CC4 16apr25 BOT | 0 | 0 | -5.3189 | -31.595 | 0
[6 empty] | 108 | 4 | CC4 16apr25 BOT | 0 | 0 | -5.4155 | -31.3957 | 0
[6 empty] | 109 | 1 | CP1 22may24 0.1m | 0 | 0 | -5.2225 | -30.9885 | 0
[6 empty] | 110 | 2 | CP1 22may24 0.1m | 0 | 0 | -5.2124 | -30.9621 | 0
[6 empty] | 111 | 3 | CP1 22may24 0.1m | 0 | 0 | -5.1693 | -31.0195 | 0
[6 empty] | 112 | 4 | CP1 22may24 0.1m | 0 | 0 | -5.1622 | -30.9683 | 0
[6 empty] | 113 | 1 | CP2 19jun24 0.1m | 0 | 0 | -4.7382 | -29.0212 | 0
[6 empty] | 114 | 2 | CP2 19jun24 0.1m | 0 | 0 | -4.8177 | -28.7638 | 0
[6 empty] | 115 | 3 | CP2 19jun24 0.1m | 0 | 0 | -4.7369 | -29.0488 | 0
[6 empty] | 116 | 4 | CP2 19jun24 0.1m | 0 | 0 | -4.8166 | -28.7624 | 0
[6 empty] | 117 | 1 | CS1 19jun24 0.1m | 0 | 0 | -6.5928 | -36.7403 | 0
[6 empty] | 118 | 2 | CS1 19jun24 0.1m | 0 | 0 | -6.6403 | -36.8879 | 0
[6 empty] | 119 | 3 | CS1 19jun24 0.1m | 0 | 0 | -6.5485 | -36.5576 | 0
[6 empty] | 120 | 4 | CS1 19jun24 0.1m | 0 | 0 | -6.617 | -36.7388 | 0
[6 empty] | 121 | 1 | CC5 22may24 9m DUP | 0 | 0 | -5.4999 | -31.9501 | 0
[6 empty] | 122 | 2 | CC5 22may24 9m DUP | 0 | 0 | -5.5484 | -32.0797 | 0
[6 empty] | 123 | 3 | CC5 22may24 9m DUP | 0 | 0 | -5.4706 | -32.4026 | 0
[6 empty] | 124 | 4 | CC5 22may24 9m DUP | 0 | 0 | -5.4728 | -32.5043 | 0
[6 empty] | 125 | 1 | Hawaii | 0 | 0 | -3.8345 | -13.3459 | 0
[6 empty] | 126 | 2 | Hawaii | 0 | 0 | -3.7929 | -13.4354 | 0
[6 empty] | 127 | 3 | Hawaii | 0 | 0 | -3.8051 | -13.2692 | 0
[6 empty] | 128 | 4 | Hawaii | 0 | 0 | -3.8157 | -13.2729 | 0
[6 empty] | 129 | 1 | Blacksburg | 0 | 0 | -7.7685 | -47.9856 | 0
[6 empty] | 130 | 2 | Blacksburg | 0 | 0 | -7.6773 | -47.6088 | 0
[6 empty] | 131 | 3 | Blacksburg | 0 | 0 | -7.7361 | -48.0165 | 0
[6 empty] | 132 | 4 | Blacksburg | 0 | 0 | -7.7302 | -48.1539 | 0