## Instructions
Picarro Isotope Analyzer
Data Processor 3.0 (KAP)
How to Use this Spreadsheet:
Step 1: Retrieve raw result file from Piccaro CPU (C:Piccaro\IsotopeData). Add initials and project name to end
of file name. Copy file into your folder on Water Server.
Step 2: On syncing computer, move raw file into Raw Data folder on Google Drive:
Instruments and Analyses --> Picarro Water Isotope Analyzer --> Raw Data
Step 3: Save this spreadsheet in the Processed Data folder on Google Drive:
Instruments and Analyses --> Picarro Water Isotope Analyzer --> Processed Data
Use file name: YYYYMMDD_Initials_(anything else to identify ex. project, run number, etc.)
Step 4: Enter run information and sample positions in the "Tray Configuration" tab. (orange cells)
(It is important to use the same names for the standards throughout this spreadsheet)
Step 5: Copy text from raw result file and paste into "Picarro Output" tab.
Step 6: In Run Summary tab, check that H2O Mean is within "good" bounds. (between 17,000 - 23,000)
Step 7: In 18O tab, verify that the reference defined values match the references used in run (cells U39 - 42). Run Solver and
verify stdev memory corrected are reduced (column K values less than corresponding column I values). Verify stdev of
drift corrected reference is lower than memory corrected. Verify normalization R2 is 0.99+. Verify QC checks are within
0.15 tolerance.
Step 8: In dD tab, verify that the reference defined values match the references used in run (cells U39 - 42. Run Solver and
verify stdev memory corrected are reduced (column K values less than corresponding column I values). Verify stdev of
drift corrected reference is lower than memory corrected. Verify normalization R2 is 0.99+. Verify QC checks are within
1.5 tolerance.
Step 9: Fill in Quality Assurance Verification chart on RESULTS tab. Dr. Scott or Kelly Peeler will fill in as reviewer 2.
Step 10: Add reference results to Water Isotope Analyzer run log.

## Tray Configuration
| Analyst |  | DWH
| PI |  | Durelle Scott
| Project |  | CCR
| Sampling Date |  | 2024 - 2025
| Sampling Location |  | Carvins Cove Reservoir
| Analysis Date |  | 2025-05-30
Notes
Rear Rack
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
1 | Blacksburg | Myrtle | Homer | Blacksburg | Hawaii | CC4 17feb25 9m | C50 22may24 20m | CC2 16apr25 0.1m | Blacksburg | 9
10 | CC4 16apr25 6m | CC3 19jun24 1.5m | CC2 28oct24 0.1m | C50 16sep24 9m | CC4 16apr25 9m | C50 16apr25 20m | CP2 19jun24 BOT | CS2 16apr25 0.1m | Blacksburg | 18
19 | CC2 19jun24 BOT | CC4 22may24 1.5m | C50 22may24 6m | CS1 16apr25 0.1m | CC4 19jun24 BOT | C50 19jun24 20m | CP2 19jun24 BOT - DUP | Hawaii | Blacksburg | 27
28 | [9 empty] | 36
37 | [9 empty] | 45
46 | [9 empty] | 54
Rear Rack
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
1 | Drift | High reference | Low Reference | Drift | QC Check | Sample | Sample | Sample | Drift | 9
10 | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Drift | 18
19 | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Drift | 27
28 | [9 empty] | 36
37 | [9 empty] | 45
46 | [9 empty] | 54
Standards
|  | Standard ID | Actual d(18_16) | Actual d(D_H)
| #1 | Denali | -19.9287 | -153.7172 | [6 empty] | -3.37
| #2 | Alberta | -18.8392 | -146.8245 | [6 empty] | -20.92
| #3 | MT BHB | -18.431 | -140.5778 | [6 empty] | -17.35
| #4 | Anchorage | -16.4204 | -129.9028 | [6 empty] | -7.39
| #5 | Homer | -14.5399 | -110.5169
| #6 | MT Creek | -13.3562 | -97.7428
| #7 | Nickel Mines | -7.876 | -48.7951
| #8 | Blacksburg | -7.5615 | -47.446
| #9 | Denton's Well | -6.9583 | -42.6011
| #10 | Hawaii | -3.6108 | -13.2294
| #11 | Myrtle Beach | -2.2908 | -10.453
| #12 | Heavy Oxygen | 3.1753 | -15.7881
| #13 | Heavy Hydrogen | 11.1372 | 4.0332
| #14 | Andres Heavy Water | 17.5664 | 20.5382
| #15 | Picarro ZERO | 0.3 | 1.8
| #16 | Picarro MID | -20.6 | -159
| #17 | Picarro DEPL | -29.6 | -235
#18
#19
#20
|  | DI | -6 | -36
Note on C34: Nate: The Standard ID must match the Standard ID used above in the trey protocol. For example, if you are using the Boulder Colorado Standard, and you labeled it "B" above, it must be labled "B" in this collum.

## Timing
| Enter start time and date --> | 2019-01-10 09:30
Tray Position | Sample ID | Start Time | # Injections
1 | Blacksburg | 2019-01-10 10:58 | 10
2 | Myrtle | 2019-01-10 12:27 | 10
3 | Homer | 2019-01-10 13:55 | 10
4 | Blacksburg | 2019-01-10 15:23 | 10
5 | Hawaii | 2019-01-10 15:59 | 4
6 | CC4 17feb25 9m | 2019-01-10 16:34 | 4
7 | C50 22may24 20m | 2019-01-10 17:09 | 4
8 | CC2 16apr25 0.1m | 2019-01-10 17:45 | 4
9 | Blacksburg | 2019-01-10 18:20 | 4
10 | CC4 16apr25 6m | 2019-01-10 18:55 | 4
11 | CC3 19jun24 1.5m | 2019-01-10 19:31 | 4
12 | CC2 28oct24 0.1m | 2019-01-10 20:06 | 4
13 | C50 16sep24 9m | 2019-01-10 20:41 | 4
14 | CC4 16apr25 9m | 2019-01-10 21:17 | 4
15 | C50 16apr25 20m | 2019-01-10 21:52 | 4
16 | CP2 19jun24 BOT | 2019-01-10 22:27 | 4
17 | CS2 16apr25 0.1m | 2019-01-10 23:03 | 4
18 | Blacksburg | 2019-01-10 23:38 | 4
19 | CC2 19jun24 BOT | 2019-01-11 00:13 | 4
20 | CC4 22may24 1.5m | 2019-01-11 00:49 | 4
21 | C50 22may24 6m | 2019-01-11 01:24 | 4
22 | CS1 16apr25 0.1m | 2019-01-11 01:59 | 4
23 | CC4 19jun24 BOT | 2019-01-11 02:35 | 4
24 | C50 19jun24 20m | 2019-01-11 03:10 | 4
25 | CP2 19jun24 BOT - DUP | 2019-01-11 03:45 | 4
26 | Hawaii | 2019-01-11 04:21 | 4
27 | Blacksburg | 2019-01-11 04:56 | 4
Formulas:
Start Time = [Start Time, row above]+"0:08:50"*10

## Picarro Output
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | days_Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | BASELINE_Mean | BASELINE_SD | DAS Temp | Tray | Sample | Job | Method
1 | P-11782 | 2025/05/30 09:21:18 | MT1-Rear-01 | 1 | -9.115 | -49.164 | 19818 | -1 | 1 |  |  | H2O | 149.3667 | 0.13 | 0.352 | 104 | -0.026 | -0.082 | 31.41 | 845 | 0 | 30.438 | MT1-Rear | 1 | 66 | BSE
2 | P-11782 | 2025/05/30 09:30:04 | MT1-Rear-01 | 2 | -8.833 | -44.946 | 20160 | -1 | 1 |  |  | H2O | 149.3728 | 0.087 | 0.362 | 102 | -0.013 | -0.092 | 30.79 | 845 | 0 | 30.438 | MT1-Rear | 1 | 66 | BSE
3 | P-11782 | 2025/05/30 09:38:49 | MT1-Rear-01 | 3 | -8.751 | -44.572 | 20062 | -1 | 1 |  |  | H2O | 149.3787 | 0.182 | 0.388 | 100 | -0.025 | -0.037 | 30.06 | 845 | 0 | 30 | MT1-Rear | 1 | 66 | BSE
4 | P-11782 | 2025/05/30 09:47:36 | MT1-Rear-01 | 4 | -8.79 | -44.16 | 19988 | 0 | 1 |  |  | H2O | 149.3848 | 0.157 | 0.284 | 100 | -0.04 | -0.042 | 29.99 | 844 | 0 | 30.813 | MT1-Rear | 1 | 66 | BSE
5 | P-11782 | 2025/05/30 09:56:23 | MT1-Rear-01 | 5 | -8.846 | -44.057 | 19540 | 0 | 1 |  |  | H2O | 149.391 | 0.093 | 0.27 | 99 | -0.023 | 0.041 | 29.97 | 844 | 0 | 31.063 | MT1-Rear | 1 | 66 | BSE
6 | P-11782 | 2025/05/30 10:05:09 | MT1-Rear-01 | 6 | -8.816 | -43.958 | 19461 | 0 | 1 |  |  | H2O | 149.3971 | 0.126 | 0.401 | 101 | -0.027 | 0.079 | 30.49 | 844 | 0 | 31.063 | MT1-Rear | 1 | 66 | BSE
7 | P-11782 | 2025/05/30 10:13:56 | MT1-Rear-01 | 7 | -8.789 | -43.858 | 19412 | 0 | 1 |  |  | H2O | 149.4032 | 0.137 | 0.269 | 97 | -0.016 | 0.018 | 29.36 | 844 | 0 | 30.938 | MT1-Rear | 1 | 66 | BSE
8 | P-11782 | 2025/05/30 10:22:42 | MT1-Rear-01 | 8 | -8.729 | -43.923 | 19504 | 0 | 1 |  |  | H2O | 149.4093 | 0.152 | 0.236 | 99 | -0.03 | -0.016 | 29.76 | 844 | 0 | 30.5 | MT1-Rear | 1 | 66 | BSE
9 | P-11782 | 2025/05/30 10:31:29 | MT1-Rear-01 | 9 | -8.719 | -44.118 | 19383 | 0 | 1 |  |  | H2O | 149.4154 | 0.088 | 0.269 | 97 | -0.018 | -0.014 | 29.36 | 844 | 0 | 30.688 | MT1-Rear | 1 | 66 | BSE
10 | P-11782 | 2025/05/30 10:40:15 | MT1-Rear-01 | 10 | -8.719 | -43.772 | 19691 | 0 | 1 |  |  | H2O | 149.4215 | 0.16 | 0.198 | 98 | -0.035 | -0.001 | 29.63 | 844 | 0 | 31.563 | MT1-Rear | 1 | 66 | BSE
11 | P-11783 | 2025/05/30 10:49:02 | MT1-Rear-02 | 1 | -3.698 | -9.712 | 19676 | -1 | 1 |  |  | H2O | 149.4276 | 0.115 | 0.404 | 97 | -0.016 | -0.069 | 29.27 | 845 | 0 | 30.75 | MT1-Rear | 2 | 66 | BSE
12 | P-11783 | 2025/05/30 10:57:49 | MT1-Rear-02 | 2 | -3.315 | -5.97 | 19667 | -1 | 1 |  |  | H2O | 149.4337 | 0.136 | 0.37 | 99 | -0.035 | -0.028 | 29.92 | 845 | 0 | 30.563 | MT1-Rear | 2 | 66 | BSE
13 | P-11783 | 2025/05/30 11:06:35 | MT1-Rear-02 | 3 | -3.27 | -5.228 | 19364 | -1 | 1 |  |  | H2O | 149.4398 | 0.106 | 0.402 | 96 | -0.002 | -0.084 | 28.82 | 845 | 0 | 30.5 | MT1-Rear | 2 | 66 | BSE
14 | P-11783 | 2025/05/30 11:15:22 | MT1-Rear-02 | 4 | -3.195 | -4.65 | 19657 | 0 | 1 |  |  | H2O | 149.4458 | 0.095 | 0.277 | 91 | -0.014 | -0.031 | 30.16 | 845 | 0 | 30.625 | MT1-Rear | 2 | 66 | BSE
15 | P-11783 | 2025/05/30 11:24:09 | MT1-Rear-02 | 5 | -3.216 | -4.594 | 19707 | 0 | 1 |  |  | H2O | 149.452 | 0.166 | 0.332 | 98 | -0.035 | 0.027 | 29.67 | 844 | 0 | 30.875 | MT1-Rear | 2 | 66 | BSE
16 | P-11783 | 2025/05/30 11:32:56 | MT1-Rear-02 | 6 | -3.226 | -4.396 | 19648 | 0 | 1 |  |  | H2O | 149.458 | 0.127 | 0.392 | 98 | -0.023 | -0.046 | 29.43 | 844 | 0 | 30.875 | MT1-Rear | 2 | 66 | BSE
17 | P-11783 | 2025/05/30 11:41:42 | MT1-Rear-02 | 7 | -3.186 | -4.446 | 19633 | 0 | 1 |  |  | H2O | 149.4641 | 0.124 | 0.365 | 96 | -0.005 | -0.058 | 29.03 | 844 | 0 | 31.188 | MT1-Rear | 2 | 66 | BSE
18 | P-11783 | 2025/05/30 11:50:29 | MT1-Rear-02 | 8 | -3.144 | -4.367 | 19707 | 0 | 1 |  |  | H2O | 149.4702 | 0.097 | 0.331 | 95 | 0.013 | -0.065 | 28.53 | 844 | 0 | 31.125 | MT1-Rear | 2 | 66 | BSE
19 | P-11783 | 2025/05/30 11:59:15 | MT1-Rear-02 | 9 | -3.2 | -4.179 | 19743 | 0 | 1 |  |  | H2O | 149.4763 | 0.173 | 0.305 | 99 | -0.025 | -0.026 | 29.75 | 844 | 0 | 31.063 | MT1-Rear | 2 | 66 | BSE
20 | P-11783 | 2025/05/30 12:08:02 | MT1-Rear-02 | 10 | -3.153 | -4.425 | 19723 | 0 | 1 |  |  | H2O | 149.4824 | 0.159 | 0.316 | 98 | -0.021 | 0.032 | 29.58 | 845 | 0 | 30.938 | MT1-Rear | 2 | 66 | BSE
21 | P-11784 | 2025/05/30 12:16:49 | MT1-Rear-03 | 1 | -14.826 | -94.75 | 19785 | -1 | 1 |  |  | H2O | 149.4885 | 0.124 | 0.424 | 95 | 0.013 | 0.094 | 28.75 | 844 | 0 | 30.875 | MT1-Rear | 3 | 66 | BSE
22 | P-11784 | 2025/05/30 12:25:35 | MT1-Rear-03 | 2 | -15.529 | -105.054 | 19525 | -1 | 1 |  |  | H2O | 149.4946 | 0.104 | 0.502 | 95 | 0.007 | 0.105 | 28.69 | 844 | 0 | 31.125 | MT1-Rear | 3 | 66 | BSE
23 | P-11784 | 2025/05/30 12:34:22 | MT1-Rear-03 | 3 | -15.726 | -106.917 | 19482 | -1 | 1 |  |  | H2O | 149.5007 | 0.104 | 0.575 | 96 | -0.017 | 0.148 | 28.88 | 844 | 0 | 30.938 | MT1-Rear | 3 | 66 | BSE
24 | P-11784 | 2025/05/30 12:43:08 | MT1-Rear-03 | 4 | -15.775 | -107.625 | 19482 | 0 | 1 |  |  | H2O | 149.5068 | 0.13 | 0.256 | 96 | -0.019 | 0.049 | 29.03 | 844 | 0 | 30.625 | MT1-Rear | 3 | 66 | BSE
25 | P-11784 | 2025/05/30 12:51:56 | MT1-Rear-03 | 5 | -15.836 | -108.105 | 19509 | 0 | 1 |  |  | H2O | 149.5129 | 0.149 | 0.41 | 98 | -0.017 | 0.074 | 29.46 | 844 | 0 | 30.375 | MT1-Rear | 3 | 66 | BSE
26 | P-11784 | 2025/05/30 13:00:43 | MT1-Rear-03 | 6 | -15.812 | -108.481 | 20068 | 0 | 1 |  |  | H2O | 149.519 | 0.099 | 0.513 | 97 | -0.006 | 0.118 | 29.36 | 844 | 0 | 30.438 | MT1-Rear | 3 | 66 | BSE
27 | P-11784 | 2025/05/30 13:09:30 | MT1-Rear-03 | 7 | -15.766 | -108.906 | 19447 | 0 | 1 |  |  | H2O | 149.5251 | 0.114 | 0.29 | 98 | -0.033 | 0.032 | 29.43 | 844 | 0 | 30.5 | MT1-Rear | 3 | 66 | BSE
28 | P-11784 | 2025/05/30 13:18:17 | MT1-Rear-03 | 8 | -15.791 | -109.01 | 19900 | 0 | 1 |  |  | H2O | 149.5312 | 0.084 | 0.532 | 98 | -0.002 | 0.117 | 29.39 | 844 | 0 | 30.813 | MT1-Rear | 3 | 66 | BSE
29 | P-11784 | 2025/05/30 13:27:04 | MT1-Rear-03 | 9 | -15.915 | -109.403 | 19452 | 0 | 1 |  |  | H2O | 149.5373 | 0.147 | 0.378 | 97 | -0.034 | 0.074 | 29.17 | 844 | 0 | 30.875 | MT1-Rear | 3 | 66 | BSE
30 | P-11784 | 2025/05/30 13:35:51 | MT1-Rear-03 | 10 | -15.796 | -108.936 | 19993 | 0 | 1 |  |  | H2O | 149.5434 | 0.164 | 0.287 | 94 | -0.021 | 0.051 | 28.23 | 844 | 0 | 30.875 | MT1-Rear | 3 | 66 | BSE
31 | P-11785 | 2025/05/30 13:44:38 | MT1-Rear-04 | 1 | -9.146 | -52.767 | 19820 | -1 | 1 |  |  | H2O | 149.5495 | 0.121 | 0.317 | 95 | -0.028 | -0.037 | 28.53 | 844 | 0 | 31 | MT1-Rear | 4 | 66 | BSE
32 | P-11785 | 2025/05/30 13:53:26 | MT1-Rear-04 | 2 | -8.892 | -46.38 | 19553 | -1 | 1 |  |  | H2O | 149.5556 | 0.087 | 0.193 | 94 | -0.016 | -0.032 | 28.3 | 844 | 0 | 31 | MT1-Rear | 4 | 66 | BSE
33 | P-11785 | 2025/05/30 14:02:14 | MT1-Rear-04 | 3 | -8.664 | -45.136 | 19788 | -1 | 1 |  |  | H2O | 149.5617 | 0.119 | 0.353 | 107 | -0.016 | -0.018 | 29.51 | 844 | 0 | 30.75 | MT1-Rear | 4 | 66 | BSE
34 | P-11785 | 2025/05/30 14:11:00 | MT1-Rear-04 | 4 | -8.709 | -44.613 | 19766 | 0 | 1 |  |  | H2O | 149.5678 | 0.092 | 0.341 | 95 | -0.007 | -0.042 | 28.48 | 844 | 0 | 30.938 | MT1-Rear | 4 | 66 | BSE
35 | P-11785 | 2025/05/30 14:19:48 | MT1-Rear-04 | 5 | -8.641 | -44.607 | 19755 | 0 | 1 |  |  | H2O | 149.5739 | 0.098 | 0.29 | 105 | -0.011 | 0.039 | 29.13 | 844 | 0 | 31 | MT1-Rear | 4 | 66 | BSE
36 | P-11785 | 2025/05/30 14:28:35 | MT1-Rear-04 | 6 | -8.726 | -44.295 | 19501 | 0 | 1 |  |  | H2O | 149.58 | 0.121 | 0.45 | 111 | 0.021 | 0.102 | 30.68 | 844 | 0 | 31 | MT1-Rear | 4 | 66 | BSE
37 | P-11785 | 2025/05/30 14:37:23 | MT1-Rear-04 | 7 | -8.591 | -44.044 | 19814 | 0 | 1 |  |  | H2O | 149.5862 | 0.098 | 0.294 | 101 | 0.009 | 0.049 | 30.35 | 844 | 0 | 30.875 | MT1-Rear | 4 | 66 | BSE
38 | P-11785 | 2025/05/30 14:46:10 | MT1-Rear-04 | 8 | -8.589 | -43.919 | 20180 | 0 | 1 |  |  | H2O | 149.5923 | 0.098 | 0.429 | 97 | -0.001 | -0.041 | 29.26 | 844 | 0 | 31 | MT1-Rear | 4 | 66 | BSE
39 | P-11785 | 2025/05/30 14:54:57 | MT1-Rear-04 | 9 | -8.611 | -44.061 | 19479 | 0 | 1 |  |  | H2O | 149.5984 | 0.131 | 0.283 | 95 | 0.004 | -0.015 | 28.74 | 844 | 0 | 30.875 | MT1-Rear | 4 | 66 | BSE
40 | P-11785 | 2025/05/30 15:03:45 | MT1-Rear-04 | 10 | -8.677 | -44.008 | 19479 | 0 | 1 |  |  | H2O | 149.6045 | 0.152 | 0.288 | 96 | -0.016 | -0.019 | 28.99 | 844 | 0 | 31 | MT1-Rear | 4 | 66 | BSE
41 | P-11786 | 2025/05/30 15:12:33 | MT1-Rear-05 | 1 | -4.878 | -12.652 | 19425 | -1 | 1 |  |  | H2O | 149.6106 | 0.093 | 0.277 | 94 | 0 | -0.053 | 28.34 | 844 | 0 | 31 | MT1-Rear | 5 | 67 | BSE
42 | P-11786 | 2025/05/30 15:21:20 | MT1-Rear-05 | 2 | -4.647 | -8.963 | 19913 | -1 | 1 |  |  | H2O | 149.6167 | 0.147 | 0.386 | 97 | -0.023 | -0.069 | 29.15 | 844 | 0 | 30.938 | MT1-Rear | 5 | 67 | BSE
43 | P-11786 | 2025/05/30 15:30:08 | MT1-Rear-05 | 3 | -4.595 | -8.292 | 19945 | -1 | 1 |  |  | H2O | 149.6228 | 0.121 | 0.307 | 97 | -0.025 | -0.068 | 29.11 | 844 | 0 | 30.938 | MT1-Rear | 5 | 67 | BSE
44 | P-11786 | 2025/05/30 15:38:56 | MT1-Rear-05 | 4 | -4.682 | -7.909 | 19762 | 0 | 1 |  |  | H2O | 149.6289 | 0.144 | 0.203 | 96 | -0.027 | 0.004 | 28.86 | 844 | 0 | 30.938 | MT1-Rear | 5 | 67 | BSE
45 | P-11787 | 2025/05/30 15:47:44 | MT1-Rear-06 | 1 | -5.99 | -24.127 | 20018 | -1 | 1 |  |  | H2O | 149.635 | 0.1 | 0.258 | 98 | -0.005 | -0.016 | 29.55 | 844 | 0 | 30.875 | MT1-Rear | 6 | 67 | BSE
46 | P-11787 | 2025/05/30 15:56:31 | MT1-Rear-06 | 2 | -6.186 | -25.301 | 19780 | -1 | 1 |  |  | H2O | 149.6411 | 0.097 | 0.333 | 96 | -0.016 | -0.015 | 28.78 | 844 | 0 | 31 | MT1-Rear | 6 | 67 | BSE
47 | P-11787 | 2025/05/30 16:05:18 | MT1-Rear-06 | 3 | -6.174 | -25.765 | 19955 | -1 | 1 |  |  | H2O | 149.6472 | 0.079 | 0.224 | 99 | -0.015 | 0.015 | 29.7 | 844 | 0 | 30.938 | MT1-Rear | 6 | 67 | BSE
48 | P-11787 | 2025/05/30 16:14:07 | MT1-Rear-06 | 4 | -6.211 | -25.554 | 19667 | 0 | 1 |  |  | H2O | 149.6533 | 0.106 | 0.384 | 94 | -0.021 | 0.018 | 28.29 | 844 | 0 | 31.125 | MT1-Rear | 6 | 67 | BSE
49 | P-11788 | 2025/05/30 16:22:55 | MT1-Rear-07 | 1 | -6.426 | -26.612 | 19892 | -1 | 1 |  |  | H2O | 149.6594 | 0.106 | 0.4 | 96 | -0.007 | 0.086 | 28.81 | 844 | 0 | 30.688 | MT1-Rear | 7 | 67 | BSE
50 | P-11788 | 2025/05/30 16:31:42 | MT1-Rear-07 | 2 | -6.429 | -26.738 | 19907 | -1 | 1 |  |  | H2O | 149.6655 | 0.09 | 0.252 | 95 | 0.001 | -0.032 | 28.49 | 844 | 0 | 30.438 | MT1-Rear | 7 | 67 | BSE
51 | P-11788 | 2025/05/30 16:40:30 | MT1-Rear-07 | 3 | -6.45 | -26.686 | 19811 | -1 | 1 |  |  | H2O | 149.6716 | 0.148 | 0.383 | 94 | -0.025 | -0.01 | 28.27 | 844 | 0 | 30.375 | MT1-Rear | 7 | 67 | BSE
52 | P-11788 | 2025/05/30 16:49:17 | MT1-Rear-07 | 4 | -6.545 | -26.607 | 19960 | 0 | 1 |  |  | H2O | 149.6777 | 0.127 | 0.201 | 96 | -0.027 | 0 | 28.84 | 844 | 0 | 30.313 | MT1-Rear | 7 | 67 | BSE
53 | P-11789 | 2025/05/30 16:58:05 | MT1-Rear-08 | 1 | -6.391 | -25.127 | 19919 | -1 | 1 |  |  | H2O | 149.6838 | 0.15 | 0.333 | 94 | -0.027 | 0.014 | 28.21 | 844 | 0 | 30.625 | MT1-Rear | 8 | 67 | BSE
54 | P-11789 | 2025/05/30 17:06:53 | MT1-Rear-08 | 2 | -6.473 | -24.991 | 20051 | -1 | 1 |  |  | H2O | 149.6899 | 0.13 | 0.234 | 97 | -0.013 | -0.028 | 29.15 | 844 | 0 | 30.938 | MT1-Rear | 8 | 67 | BSE
55 | P-11789 | 2025/05/30 17:15:42 | MT1-Rear-08 | 3 | -6.488 | -24.864 | 19993 | -1 | 1 |  |  | H2O | 149.6961 | 0.104 | 0.219 | 95 | -0.002 | -0.011 | 28.78 | 844 | 0 | 31.313 | MT1-Rear | 8 | 67 | BSE
56 | P-11789 | 2025/05/30 17:24:30 | MT1-Rear-08 | 4 | -6.465 | -24.833 | 20058 | 0 | 1 |  |  | H2O | 149.7022 | 0.105 | 0.167 | 96 | -0.016 | 0.018 | 28.82 | 844 | 0 | 31.438 | MT1-Rear | 8 | 67 | BSE
57 | P-11790 | 2025/05/30 17:33:19 | MT1-Rear-09 | 1 | -8.815 | -39.96 | 20364 | -1 | 1 |  |  | H2O | 149.7083 | 0.096 | 0.217 | 96 | -0.01 | 0.046 | 28.85 | 844 | 0 | 31.375 | MT1-Rear | 9 | 67 | BSE
58 | P-11790 | 2025/05/30 17:42:07 | MT1-Rear-09 | 2 | -8.972 | -41.66 | 19979 | -1 | 1 |  |  | H2O | 149.7144 | 0.097 | 0.345 | 97 | -0.017 | -0.041 | 29.37 | 844 | 0 | 30.938 | MT1-Rear | 9 | 67 | BSE
59 | P-11790 | 2025/05/30 17:50:54 | MT1-Rear-09 | 3 | -8.971 | -41.945 | 19919 | -1 | 1 |  |  | H2O | 149.7205 | 0.094 | 0.358 | 97 | -0.006 | 0.023 | 29.23 | 844 | 0 | 30.688 | MT1-Rear | 9 | 67 | BSE
60 | P-11790 | 2025/05/30 17:59:42 | MT1-Rear-09 | 4 | -8.92 | -42.316 | 19462 | 0 | 1 |  |  | H2O | 149.7266 | 0.157 | 0.344 | 97 | -0.022 | 0.004 | 29.17 | 844 | 0 | 30.813 | MT1-Rear | 9 | 67 | BSE
61 | P-11791 | 2025/05/30 18:08:31 | MT1-Rear-10 | 1 | -6.609 | -27.453 | 19958 | -1 | 1 |  |  | H2O | 149.7327 | 0.135 | 0.351 | 92 | -0.028 | -0.041 | 27.82 | 844 | 0 | 31.063 | MT1-Rear | 10 | 67 | BSE
62 | P-11791 | 2025/05/30 18:17:20 | MT1-Rear-10 | 2 | -6.537 | -25.735 | 19940 | -1 | 1 |  |  | H2O | 149.7388 | 0.089 | 0.338 | 93 | -0.01 | -0.042 | 28.12 | 844 | 0 | 31.125 | MT1-Rear | 10 | 67 | BSE
63 | P-11791 | 2025/05/30 18:26:10 | MT1-Rear-10 | 3 | -6.517 | -25.13 | 20027 | -1 | 1 |  |  | H2O | 149.7451 | 0.163 | 0.364 | 97 | -0.034 | 0.005 | 29.14 | 844 | 0 | 30.938 | MT1-Rear | 10 | 67 | BSE
64 | P-11791 | 2025/05/30 18:35:00 | MT1-Rear-10 | 4 | -6.506 | -25.001 | 19929 | 0 | 1 |  |  | H2O | 149.7512 | 0.115 | 0.372 | 96 | -0.01 | -0.046 | 28.8 | 844 | 0 | 30.75 | MT1-Rear | 10 | 67 | BSE
65 | P-11792 | 2025/05/30 18:43:49 | MT1-Rear-11 | 1 | -6.016 | -23.125 | 20353 | -1 | 1 |  |  | H2O | 149.7573 | 0.124 | 0.143 | 97 | -0.015 | 0 | 29.37 | 844 | 0 | 31 | MT1-Rear | 11 | 67 | BSE
66 | P-11792 | 2025/05/30 18:52:38 | MT1-Rear-11 | 2 | -5.948 | -22.95 | 19942 | -1 | 1 |  |  | H2O | 149.7634 | 0.119 | 0.396 | 95 | 0.004 | 0.049 | 28.74 | 844 | 0 | 31 | MT1-Rear | 11 | 67 | BSE
67 | P-11792 | 2025/05/30 19:01:27 | MT1-Rear-11 | 3 | -5.851 | -23.03 | 20051 | -1 | 1 |  |  | H2O | 149.7695 | 0.142 | 0.296 | 94 | -0.013 | -0.01 | 28.17 | 844 | 0 | 31.125 | MT1-Rear | 11 | 67 | BSE
68 | P-11792 | 2025/05/30 19:10:16 | MT1-Rear-11 | 4 | -5.971 | -22.911 | 20046 | 0 | 1 |  |  | H2O | 149.7757 | 0.151 | 0.498 | 97 | -0.036 | 0 | 29.32 | 844 | 0 | 30.875 | MT1-Rear | 11 | 67 | BSE
69 | P-11793 | 2025/05/30 19:19:05 | MT1-Rear-12 | 1 | -5.507 | -19.774 | 19966 | -1 | 1 |  |  | H2O | 149.7818 | 0.104 | 0.238 | 95 | -0.012 | -0.024 | 28.6 | 844 | 0 | 30.688 | MT1-Rear | 12 | 67 | BSE
70 | P-11793 | 2025/05/30 19:27:55 | MT1-Rear-12 | 2 | -5.482 | -19.472 | 19895 | -1 | 1 |  |  | H2O | 149.7879 | 0.081 | 0.327 | 94 | 0.001 | -0.027 | 28.21 | 844 | 0 | 31 | MT1-Rear | 12 | 67 | BSE
71 | P-11793 | 2025/05/30 19:36:44 | MT1-Rear-12 | 3 | -5.497 | -19.697 | 19696 | -1 | 1 |  |  | H2O | 149.794 | 0.158 | 0.341 | 95 | -0.034 | 0.027 | 28.56 | 844 | 0 | 31.188 | MT1-Rear | 12 | 67 | BSE
72 | P-11793 | 2025/05/30 19:45:34 | MT1-Rear-12 | 4 | -5.477 | -19.55 | 20270 | 0 | 1 |  |  | H2O | 149.8001 | 0.128 | 0.332 | 94 | -0.018 | -0.023 | 28.22 | 844 | 0 | 31.188 | MT1-Rear | 12 | 67 | BSE
73 | P-11794 | 2025/05/30 19:54:23 | MT1-Rear-13 | 1 | -6.429 | -24.089 | 19973 | -1 | 1 |  |  | H2O | 149.8063 | 0.193 | 0.248 | 101 | -0.037 | -0.025 | 27.98 | 844 | 0 | 31 | MT1-Rear | 13 | 67 | BSE
74 | P-11794 | 2025/05/30 20:03:12 | MT1-Rear-13 | 2 | -6.527 | -24.735 | 19921 | -1 | 1 |  |  | H2O | 149.8125 | 0.136 | 0.326 | 96 | -0.017 | 0 | 28.87 | 844 | 0 | 30.813 | MT1-Rear | 13 | 67 | BSE
75 | P-11794 | 2025/05/30 20:12:01 | MT1-Rear-13 | 3 | -6.565 | -24.956 | 19804 | -1 | 1 |  |  | H2O | 149.8186 | 0.115 | 0.256 | 95 | -0.001 | 0.031 | 28.65 | 844 | 0 | 30.938 | MT1-Rear | 13 | 67 | BSE
76 | P-11794 | 2025/05/30 20:20:50 | MT1-Rear-13 | 4 | -6.572 | -24.872 | 20042 | 0 | 1 |  |  | H2O | 149.8246 | 0.148 | 0.348 | 94 | -0.031 | 0 | 28.46 | 844 | 0 | 30.938 | MT1-Rear | 13 | 67 | BSE
77 | P-11795 | 2025/05/30 20:29:40 | MT1-Rear-14 | 1 | -6.706 | -25.167 | 19908 | -1 | 1 |  |  | H2O | 149.8307 | 0.137 | 0.333 | 93 | -0.02 | -0.022 | 28.12 | 844 | 0 | 30.813 | MT1-Rear | 14 | 67 | BSE
78 | P-11795 | 2025/05/30 20:38:29 | MT1-Rear-14 | 2 | -6.728 | -25.43 | 19932 | -1 | 1 |  |  | H2O | 149.837 | 0.072 | 0.468 | 94 | -0.01 | -0.022 | 28.19 | 844 | 0 | 30.875 | MT1-Rear | 14 | 67 | BSE
79 | P-11795 | 2025/05/30 20:47:19 | MT1-Rear-14 | 3 | -6.591 | -25.367 | 20080 | -1 | 1 |  |  | H2O | 149.8431 | 0.086 | 0.242 | 94 | -0.015 | -0.019 | 28.38 | 844 | 0 | 31 | MT1-Rear | 14 | 67 | BSE
80 | P-11795 | 2025/05/30 20:56:09 | MT1-Rear-14 | 4 | -6.74 | -25.309 | 19944 | 0 | 1 |  |  | H2O | 149.8492 | 0.115 | 0.257 | 93 | 0.009 | -0.015 | 27.96 | 844 | 0 | 31 | MT1-Rear | 14 | 67 | BSE
81 | P-11796 | 2025/05/30 21:04:58 | MT1-Rear-15 | 1 | -6.659 | -25.43 | 20208 | -1 | 1 |  |  | H2O | 149.8553 | 0.133 | 0.291 | 90 | -0.009 | 0.002 | 27.19 | 844 | 0 | 31.188 | MT1-Rear | 15 | 67 | BSE
82 | P-11796 | 2025/05/30 21:13:49 | MT1-Rear-15 | 2 | -6.675 | -25.826 | 19908 | -1 | 1 |  |  | H2O | 149.8615 | 0.182 | 0.188 | 96 | -0.036 | 0.028 | 28.85 | 844 | 0 | 31.375 | MT1-Rear | 15 | 67 | BSE
83 | P-11796 | 2025/05/30 21:22:39 | MT1-Rear-15 | 3 | -6.744 | -25.57 | 20209 | -1 | 1 |  |  | H2O | 149.8676 | 0.123 | 0.334 | 95 | -0.012 | 0.011 | 28.47 | 844 | 0 | 31 | MT1-Rear | 15 | 67 | BSE
84 | P-11796 | 2025/05/30 21:31:29 | MT1-Rear-15 | 4 | -6.738 | -25.212 | 20005 | 0 | 1 |  |  | H2O | 149.8737 | 0.085 | 0.247 | 91 | -0.01 | 0.037 | 27.46 | 844 | 0 | 30.938 | MT1-Rear | 15 | 67 | BSE
85 | P-11797 | 2025/05/30 21:40:18 | MT1-Rear-16 | 1 | -6.187 | -23.154 | 19896 | -1 | 1 |  |  | H2O | 149.8799 | 0.129 | 0.312 | 96 | -0.031 | -0.04 | 28.79 | 844 | 0 | 30.563 | MT1-Rear | 16 | 67 | BSE
86 | P-11797 | 2025/05/30 21:49:08 | MT1-Rear-16 | 2 | -6.094 | -23.194 | 19851 | -1 | 1 |  |  | H2O | 149.886 | 0.137 | 0.412 | 94 | -0.031 | -0.082 | 28.47 | 844 | 0 | 31 | MT1-Rear | 16 | 67 | BSE
87 | P-11797 | 2025/05/30 21:57:58 | MT1-Rear-16 | 3 | -6.185 | -23.087 | 20239 | -1 | 1 |  |  | H2O | 149.8921 | 0.072 | 0.323 | 92 | -0.007 | -0.014 | 27.85 | 844 | 0 | 31 | MT1-Rear | 16 | 67 | BSE
88 | P-11797 | 2025/05/30 22:06:48 | MT1-Rear-16 | 4 | -6.257 | -22.822 | 19523 | 0 | 1 |  |  | H2O | 149.8983 | 0.153 | 0.312 | 95 | -0.032 | 0.006 | 28.5 | 844 | 0 | 30.875 | MT1-Rear | 16 | 67 | BSE
89 | P-11798 | 2025/05/30 22:15:38 | MT1-Rear-17 | 1 | -7.923 | -29.619 | 19924 | -1 | 1 |  |  | H2O | 149.9044 | 0.121 | 0.244 | 91 | 0.004 | -0.033 | 27.49 | 844 | 0 | 30.75 | MT1-Rear | 17 | 67 | BSE
90 | P-11798 | 2025/05/30 22:24:28 | MT1-Rear-17 | 2 | -7.988 | -30.603 | 20207 | -1 | 1 |  |  | H2O | 149.9105 | 0.151 | 0.231 | 90 | -0.035 | 0.003 | 27.25 | 844 | 0 | 30.938 | MT1-Rear | 17 | 67 | BSE
91 | P-11798 | 2025/05/30 22:33:19 | MT1-Rear-17 | 3 | -7.976 | -30.88 | 20043 | -1 | 1 |  |  | H2O | 149.9167 | 0.129 | 0.319 | 95 | -0.015 | -0.059 | 28.68 | 844 | 0 | 31.125 | MT1-Rear | 17 | 67 | BSE
92 | P-11798 | 2025/05/30 22:42:09 | MT1-Rear-17 | 4 | -8.041 | -30.61 | 20008 | 0 | 1 |  |  | H2O | 149.9228 | 0.12 | 0.284 | 92 | -0.017 | -0.012 | 27.78 | 844 | 0 | 30.875 | MT1-Rear | 17 | 67 | BSE
93 | P-11799 | 2025/05/30 22:51:00 | MT1-Rear-18 | 1 | -9.091 | -40.714 | 19993 | -1 | 1 |  |  | H2O | 149.9289 | 0.092 | 0.397 | 95 | -0.016 | 0.05 | 28.58 | 844 | 0 | 30.813 | MT1-Rear | 18 | 67 | BSE
94 | P-11799 | 2025/05/30 22:59:50 | MT1-Rear-18 | 2 | -9.173 | -41.839 | 20062 | -1 | 1 |  |  | H2O | 149.9352 | 0.161 | 0.348 | 96 | 0.018 | 0.078 | 28.98 | 844 | 0 | 30.75 | MT1-Rear | 18 | 67 | BSE
95 | P-11799 | 2025/05/30 23:08:39 | MT1-Rear-18 | 3 | -9.158 | -42.242 | 20336 | -1 | 1 |  |  | H2O | 149.9412 | 0.155 | 0.29 | 94 | -0.015 | 0.051 | 28.42 | 844 | 0 | 30.875 | MT1-Rear | 18 | 67 | BSE
96 | P-11799 | 2025/05/30 23:17:30 | MT1-Rear-18 | 4 | -9.133 | -42.125 | 20311 | 0 | 1 |  |  | H2O | 149.9473 | 0.105 | 0.273 | 95 | -0.007 | 0.048 | 28.56 | 844 | 0 | 30.938 | MT1-Rear | 18 | 67 | BSE
97 | P-11800 | 2025/05/30 23:26:19 | MT1-Rear-19 | 1 | -6.414 | -25.188 | 20305 | -1 | 1 |  |  | H2O | 149.9534 | 0.108 | 0.291 | 95 | -0.018 | -0.009 | 28.57 | 844 | 0 | 30.75 | MT1-Rear | 19 | 67 | BSE
98 | P-11800 | 2025/05/30 23:35:10 | MT1-Rear-19 | 2 | -6.22 | -23.487 | 20228 | -1 | 1 |  |  | H2O | 149.9597 | 0.079 | 0.281 | 94 | 0.001 | -0.022 | 28.32 | 844 | 0 | 30.813 | MT1-Rear | 19 | 67 | BSE
99 | P-11800 | 2025/05/30 23:44:02 | MT1-Rear-19 | 3 | -6.12 | -23.314 | 20321 | -1 | 1 |  |  | H2O | 149.9658 | 0.095 | 0.345 | 93 | -0.015 | -0.004 | 27.93 | 844 | 0 | 30.875 | MT1-Rear | 19 | 67 | BSE
100 | P-11800 | 2025/05/30 23:52:52 | MT1-Rear-19 | 4 | -6.131 | -23.289 | 19603 | 0 | 1 |  |  | H2O | 149.972 | 0.157 | 0.356 | 94 | -0.021 | 0.027 | 28.29 | 844 | 0 | 30.875 | MT1-Rear | 19 | 67 | BSE
101 | P-11801 | 2025/05/31 00:01:45 | MT1-Rear-20 | 1 | -6.454 | -24.727 | 19846 | -1 | 1 |  |  | H2O | 149.9781 | 0.104 | 0.359 | 94 | 0.001 | 0.001 | 28.34 | 844 | 0 | 30.625 | MT1-Rear | 20 | 67 | BSE
102 | P-11801 | 2025/05/31 00:10:35 | MT1-Rear-20 | 2 | -6.438 | -24.846 | 19855 | -1 | 1 |  |  | H2O | 149.9842 | 0.08 | 0.295 | 91 | 0.003 | -0.02 | 27.5 | 844 | 0 | 30.938 | MT1-Rear | 20 | 67 | BSE
103 | P-11801 | 2025/05/31 00:19:26 | MT1-Rear-20 | 3 | -6.452 | -24.88 | 20298 | -1 | 1 |  |  | H2O | 149.9904 | 0.105 | 0.289 | 93 | -0.022 | -0.031 | 28.08 | 844 | 0 | 31 | MT1-Rear | 20 | 67 | BSE
104 | P-11801 | 2025/05/31 00:28:17 | MT1-Rear-20 | 4 | -6.424 | -24.705 | 19976 | 0 | 1 |  |  | H2O | 149.9965 | 0.155 | 0.328 | 90 | -0.02 | -0.021 | 27.14 | 844 | 0 | 31 | MT1-Rear | 20 | 67 | BSE
105 | P-11802 | 2025/05/31 00:37:08 | MT1-Rear-21 | 1 | -6.601 | -25.506 | 19970 | -1 | 1 |  |  | H2O | 150.0027 | 0.106 | 0.208 | 93 | -0.016 | 0.024 | 28.02 | 844 | 0 | 30.625 | MT1-Rear | 21 | 67 | BSE
106 | P-11802 | 2025/05/31 00:45:59 | MT1-Rear-21 | 2 | -6.629 | -25.607 | 20233 | -1 | 1 |  |  | H2O | 150.0088 | 0.112 | 0.241 | 95 | 0.004 | -0.021 | 28.77 | 844 | 0 | 30.75 | MT1-Rear | 21 | 67 | BSE
107 | P-11802 | 2025/05/31 00:54:50 | MT1-Rear-21 | 3 | -6.631 | -25.502 | 20378 | -1 | 1 |  |  | H2O | 150.015 | 0.152 | 0.296 | 95 | -0.029 | -0.042 | 28.71 | 844 | 0 | 30.938 | MT1-Rear | 21 | 67 | BSE
108 | P-11802 | 2025/05/31 01:03:40 | MT1-Rear-21 | 4 | -6.601 | -25.599 | 19929 | 0 | 1 |  |  | H2O | 150.0211 | 0.12 | 0.263 | 93 | -0.025 | -0.029 | 28 | 844 | 0 | 30.875 | MT1-Rear | 21 | 67 | BSE
109 | P-11803 | 2025/05/31 01:12:31 | MT1-Rear-22 | 1 | -7.91 | -30.047 | 20223 | -1 | 1 |  |  | H2O | 150.0272 | 0.149 | 0.255 | 93 | -0.007 | 0.01 | 27.91 | 844 | 0 | 30.438 | MT1-Rear | 22 | 67 | BSE
110 | P-11803 | 2025/05/31 01:21:23 | MT1-Rear-22 | 2 | -7.941 | -30.971 | 20354 | -1 | 1 |  |  | H2O | 150.0334 | 0.099 | 0.315 | 94 | -0.019 | 0.021 | 28.37 | 844 | 0 | 30.875 | MT1-Rear | 22 | 67 | BSE
111 | P-11803 | 2025/05/31 01:30:14 | MT1-Rear-22 | 3 | -7.982 | -31.171 | 19705 | -1 | 1 |  |  | H2O | 150.0395 | 0.137 | 0.273 | 91 | -0.014 | 0.011 | 27.48 | 844 | 0 | 30.875 | MT1-Rear | 22 | 67 | BSE
112 | P-11803 | 2025/05/31 01:39:05 | MT1-Rear-22 | 4 | -8.073 | -31.089 | 19928 | 0 | 1 |  |  | H2O | 150.0457 | 0.155 | 0.277 | 92 | 0.016 | 0.048 | 27.7 | 844 | 0 | 31 | MT1-Rear | 22 | 67 | BSE
113 | P-11804 | 2025/05/31 01:47:56 | MT1-Rear-23 | 1 | -6.878 | -26.68 | 20161 | -1 | 1 |  |  | H2O | 150.0518 | 0.08 | 0.301 | 94 | -0.003 | -0.002 | 28.22 | 844 | 0 | 30.625 | MT1-Rear | 23 | 67 | BSE
114 | P-11804 | 2025/05/31 01:56:47 | MT1-Rear-23 | 2 | -6.85 | -26.35 | 20011 | -1 | 1 |  |  | H2O | 150.058 | 0.098 | 0.307 | 100 | -0.017 | -0.027 | 27.67 | 844 | 0 | 30.688 | MT1-Rear | 23 | 67 | BSE
115 | P-11804 | 2025/05/31 02:05:39 | MT1-Rear-23 | 3 | -6.751 | -26.401 | 19872 | -1 | 1 |  |  | H2O | 150.0641 | 0.116 | 0.341 | 93 | -0.013 | 0.041 | 27.93 | 844 | 0 | 30.813 | MT1-Rear | 23 | 67 | BSE
116 | P-11804 | 2025/05/31 02:14:29 | MT1-Rear-23 | 4 | -6.758 | -26.46 | 20200 | 0 | 1 |  |  | H2O | 150.0702 | 0.13 | 0.334 | 94 | -0.003 | -0.014 | 28.46 | 844 | 0 | 30.875 | MT1-Rear | 23 | 67 | BSE
117 | P-11805 | 2025/05/31 02:23:21 | MT1-Rear-24 | 1 | -6.822 | -26.31 | 19747 | -1 | 1 |  |  | H2O | 150.0765 | 0.143 | 0.263 | 94 | -0.028 | 0.003 | 28.16 | 844 | 0 | 30.75 | MT1-Rear | 24 | 67 | BSE
118 | P-11805 | 2025/05/31 02:32:13 | MT1-Rear-24 | 2 | -6.842 | -26.19 | 19824 | -1 | 1 |  |  | H2O | 150.0825 | 0.116 | 0.31 | 93 | -0.014 | -0.036 | 27.96 | 844 | 0 | 30.75 | MT1-Rear | 24 | 67 | BSE
119 | P-11805 | 2025/05/31 02:41:04 | MT1-Rear-24 | 3 | -6.884 | -26.166 | 19858 | -1 | 1 |  |  | H2O | 150.0888 | 0.148 | 0.358 | 96 | -0.025 | 0.021 | 28.87 | 844 | 0 | 30.875 | MT1-Rear | 24 | 67 | BSE
120 | P-11805 | 2025/05/31 02:49:56 | MT1-Rear-24 | 4 | -6.757 | -26.433 | 19858 | 0 | 1 |  |  | H2O | 150.0948 | 0.129 | 0.317 | 93 | -0.027 | 0.063 | 27.99 | 844 | 0 | 30.875 | MT1-Rear | 24 | 67 | BSE
121 | P-11806 | 2025/05/31 02:58:48 | MT1-Rear-25 | 1 | -6.139 | -23.044 | 20347 | -1 | 1 |  |  | H2O | 150.1011 | 0.164 | 0.324 | 93 | -0.033 | -0.063 | 28 | 844 | 0 | 30.813 | MT1-Rear | 25 | 67 | BSE
122 | P-11806 | 2025/05/31 03:07:40 | MT1-Rear-25 | 2 | -6.12 | -23.087 | 19949 | -1 | 1 |  |  | H2O | 150.1072 | 0.142 | 0.257 | 92 | -0.029 | 0 | 27.62 | 844 | 0 | 30.75 | MT1-Rear | 25 | 67 | BSE
123 | P-11806 | 2025/05/31 03:16:32 | MT1-Rear-25 | 3 | -6.198 | -22.667 | 19901 | -1 | 1 |  |  | H2O | 150.1134 | 0.13 | 0.238 | 94 | -0.026 | -0.007 | 28.35 | 844 | 0 | 30.938 | MT1-Rear | 25 | 67 | BSE
124 | P-11806 | 2025/05/31 03:25:24 | MT1-Rear-25 | 4 | -6.079 | -22.815 | 20299 | 0 | 1 |  |  | H2O | 150.1195 | 0.143 | 0.207 | 101 | -0.03 | 0.002 | 27.85 | 844 | 0 | 30.688 | MT1-Rear | 25 | 67 | BSE
125 | P-11807 | 2025/05/31 03:34:16 | MT1-Rear-26 | 1 | -5.171 | -8.378 | 19801 | -1 | 1 |  |  | H2O | 150.1257 | 0.117 | 0.285 | 92 | -0.005 | -0.065 | 27.57 | 844 | 0 | 30.75 | MT1-Rear | 26 | 67 | BSE
126 | P-11807 | 2025/05/31 03:43:08 | MT1-Rear-26 | 2 | -5.085 | -6.746 | 19880 | -1 | 1 |  |  | H2O | 150.1318 | 0.165 | 0.322 | 101 | -0.026 | -0.029 | 27.98 | 844 | 0 | 30.875 | MT1-Rear | 26 | 67 | BSE
127 | P-11807 | 2025/05/31 03:52:01 | MT1-Rear-26 | 3 | -5.202 | -6.227 | 20203 | -1 | 1 |  |  | H2O | 150.138 | 0.041 | 0.429 | 93 | -0.008 | -0.049 | 28.05 | 844 | 0 | 30.938 | MT1-Rear | 26 | 67 | BSE
128 | P-11807 | 2025/05/31 04:00:52 | MT1-Rear-26 | 4 | -5.1 | -6.377 | 20051 | 0 | 1 |  |  | H2O | 150.1442 | 0.134 | 0.272 | 101 | -0.025 | -0.037 | 27.96 | 844 | 0 | 30.75 | MT1-Rear | 26 | 67 | BSE
129 | P-11808 | 2025/05/31 04:09:44 | MT1-Rear-27 | 1 | -8.822 | -37.667 | 19615 | -1 | 1 |  |  | H2O | 150.1503 | 0.103 | 0.306 | 92 | 0.002 | 0.006 | 27.79 | 844 | 0 | 30.688 | MT1-Rear | 27 | 67 | BSE
130 | P-11808 | 2025/05/31 04:18:36 | MT1-Rear-27 | 2 | -9.027 | -41.413 | 19809 | -1 | 1 |  |  | H2O | 150.1565 | 0.199 | 0.321 | 102 | -0.035 | -0.016 | 28.29 | 844 | 0 | 30.75 | MT1-Rear | 27 | 67 | BSE
131 | P-11808 | 2025/05/31 04:27:28 | MT1-Rear-27 | 3 | -9.091 | -41.958 | 19857 | -1 | 1 |  |  | H2O | 150.1626 | 0.119 | 0.437 | 92 | -0.019 | 0.024 | 27.83 | 844 | 0 | 30.938 | MT1-Rear | 27 | 67 | BSE
132 | P-11808 | 2025/05/31 04:36:20 | MT1-Rear-27 | 4 | -9.126 | -42.159 | 20291 | 0 | 1 |  |  | H2O | 150.1688 | 0.116 | 0.328 | 102 | -0.003 | 0.041 | 28.2 | 844 | 0 | 30.75 | MT1-Rear | 27 | 67 | BSE

## Run Summary
Line | Sample | Inj Nr | Sample Name | H2O_Mean | Good | Ignore | Time Code | d(18_16)Mean | d(D_H)Mean
1 | 1 | 1 | Blacksburg | 19818 | 1 | 0 | 2025/05/30 09:21:18 | -9.115 | -49.164 |  | Min | Max
2 | 1 | 2 | Blacksburg | 20160 | 1 | 0 | 2025/05/30 09:30:04 | -8.833 | -44.946 | 0 | 17000 | 23000
3 | 1 | 3 | Blacksburg | 20062 | 1 | 0 | 2025/05/30 09:38:49 | -8.751 | -44.572 | 500 | 17000 | 23000
4 | 1 | 4 | Blacksburg | 19988 | 1 | 0 | 2025/05/30 09:47:36 | -8.79 | -44.16
5 | 1 | 5 | Blacksburg | 19540 | 1 | 0 | 2025/05/30 09:56:23 | -8.846 | -44.057
6 | 1 | 6 | Blacksburg | 19461 | 1 | 0 | 2025/05/30 10:05:09 | -8.816 | -43.958
7 | 1 | 7 | Blacksburg | 19412 | 1 | 0 | 2025/05/30 10:13:56 | -8.789 | -43.858
8 | 1 | 8 | Blacksburg | 19504 | 1 | 0 | 2025/05/30 10:22:42 | -8.729 | -43.923
9 | 1 | 9 | Blacksburg | 19383 | 1 | 0 | 2025/05/30 10:31:29 | -8.719 | -44.118
10 | 1 | 10 | Blacksburg | 19691 | 1 | 0 | 2025/05/30 10:40:15 | -8.719 | -43.772
11 | 2 | 1 | Myrtle | 19676 | 1 | 0 | 2025/05/30 10:49:02 | -3.698 | -9.712
12 | 2 | 2 | Myrtle | 19667 | 1 | 0 | 2025/05/30 10:57:49 | -3.315 | -5.97
13 | 2 | 3 | Myrtle | 19364 | 1 | 0 | 2025/05/30 11:06:35 | -3.27 | -5.228
14 | 2 | 4 | Myrtle | 19657 | 1 | 0 | 2025/05/30 11:15:22 | -3.195 | -4.65
15 | 2 | 5 | Myrtle | 19707 | 1 | 0 | 2025/05/30 11:24:09 | -3.216 | -4.594
16 | 2 | 6 | Myrtle | 19648 | 1 | 0 | 2025/05/30 11:32:56 | -3.226 | -4.396
17 | 2 | 7 | Myrtle | 19633 | 1 | 0 | 2025/05/30 11:41:42 | -3.186 | -4.446 | Concentration of water vapor should be about 20,000 ppmV (roughly +/- 1000; but not outside 17,000 and 23,000 ppmV).
18 | 2 | 8 | Myrtle | 19707 | 1 | 0 | 2025/05/30 11:50:29 | -3.144 | -4.367 | Large scatter may indicate a problem with the syringe or the septa.
19 | 2 | 9 | Myrtle | 19743 | 1 | 0 | 2025/05/30 11:59:15 | -3.2 | -4.179
20 | 2 | 10 | Myrtle | 19723 | 1 | 0 | 2025/05/30 12:08:02 | -3.153 | -4.425
21 | 3 | 1 | Homer | 19785 | 1 | 0 | 2025/05/30 12:16:49 | -14.826 | -94.75
22 | 3 | 2 | Homer | 19525 | 1 | 0 | 2025/05/30 12:25:35 | -15.529 | -105.054
23 | 3 | 3 | Homer | 19482 | 1 | 0 | 2025/05/30 12:34:22 | -15.726 | -106.917
24 | 3 | 4 | Homer | 19482 | 1 | 0 | 2025/05/30 12:43:08 | -15.775 | -107.625
25 | 3 | 5 | Homer | 19509 | 1 | 0 | 2025/05/30 12:51:56 | -15.836 | -108.105
26 | 3 | 6 | Homer | 20068 | 1 | 0 | 2025/05/30 13:00:43 | -15.812 | -108.481
27 | 3 | 7 | Homer | 19447 | 1 | 0 | 2025/05/30 13:09:30 | -15.766 | -108.906
28 | 3 | 8 | Homer | 19900 | 1 | 0 | 2025/05/30 13:18:17 | -15.791 | -109.01
29 | 3 | 9 | Homer | 19452 | 1 | 0 | 2025/05/30 13:27:04 | -15.915 | -109.403
30 | 3 | 10 | Homer | 19993 | 1 | 0 | 2025/05/30 13:35:51 | -15.796 | -108.936
31 | 4 | 1 | Blacksburg | 19820 | 1 | 0 | 2025/05/30 13:44:38 | -9.146 | -52.767
32 | 4 | 2 | Blacksburg | 19553 | 1 | 0 | 2025/05/30 13:53:26 | -8.892 | -46.38
33 | 4 | 3 | Blacksburg | 19788 | 1 | 0 | 2025/05/30 14:02:14 | -8.664 | -45.136
34 | 4 | 4 | Blacksburg | 19766 | 1 | 0 | 2025/05/30 14:11:00 | -8.709 | -44.613
35 | 4 | 5 | Blacksburg | 19755 | 1 | 0 | 2025/05/30 14:19:48 | -8.641 | -44.607
36 | 4 | 6 | Blacksburg | 19501 | 1 | 0 | 2025/05/30 14:28:35 | -8.726 | -44.295
37 | 4 | 7 | Blacksburg | 19814 | 1 | 0 | 2025/05/30 14:37:23 | -8.591 | -44.044
38 | 4 | 8 | Blacksburg | 20180 | 1 | 0 | 2025/05/30 14:46:10 | -8.589 | -43.919
39 | 4 | 9 | Blacksburg | 19479 | 1 | 0 | 2025/05/30 14:54:57 | -8.611 | -44.061
40 | 4 | 10 | Blacksburg | 19479 | 1 | 0 | 2025/05/30 15:03:45 | -8.677 | -44.008
41 | 5 | 1 | Hawaii | 19425 | 1 | 0 | 2025/05/30 15:12:33 | -4.878 | -12.652
42 | 5 | 2 | Hawaii | 19913 | 1 | 0 | 2025/05/30 15:21:20 | -4.647 | -8.963
43 | 5 | 3 | Hawaii | 19945 | 1 | 0 | 2025/05/30 15:30:08 | -4.595 | -8.292
44 | 5 | 4 | Hawaii | 19762 | 1 | 0 | 2025/05/30 15:38:56 | -4.682 | -7.909
45 | 6 | 1 | CC4 17feb25 9m | 20018 | 1 | 0 | 2025/05/30 15:47:44 | -5.99 | -24.127
46 | 6 | 2 | CC4 17feb25 9m | 19780 | 1 | 0 | 2025/05/30 15:56:31 | -6.186 | -25.301
47 | 6 | 3 | CC4 17feb25 9m | 19955 | 1 | 0 | 2025/05/30 16:05:18 | -6.174 | -25.765
48 | 6 | 4 | CC4 17feb25 9m | 19667 | 1 | 0 | 2025/05/30 16:14:07 | -6.211 | -25.554
49 | 7 | 1 | C50 22may24 20m | 19892 | 1 | 0 | 2025/05/30 16:22:55 | -6.426 | -26.612
50 | 7 | 2 | C50 22may24 20m | 19907 | 1 | 0 | 2025/05/30 16:31:42 | -6.429 | -26.738
51 | 7 | 3 | C50 22may24 20m | 19811 | 1 | 0 | 2025/05/30 16:40:30 | -6.45 | -26.686
52 | 7 | 4 | C50 22may24 20m | 19960 | 1 | 0 | 2025/05/30 16:49:17 | -6.545 | -26.607
53 | 8 | 1 | CC2 16apr25 0.1m | 19919 | 1 | 0 | 2025/05/30 16:58:05 | -6.391 | -25.127
54 | 8 | 2 | CC2 16apr25 0.1m | 20051 | 1 | 0 | 2025/05/30 17:06:53 | -6.473 | -24.991
55 | 8 | 3 | CC2 16apr25 0.1m | 19993 | 1 | 0 | 2025/05/30 17:15:42 | -6.488 | -24.864
56 | 8 | 4 | CC2 16apr25 0.1m | 20058 | 1 | 0 | 2025/05/30 17:24:30 | -6.465 | -24.833
57 | 9 | 1 | Blacksburg | 20364 | 1 | 0 | 2025/05/30 17:33:19 | -8.815 | -39.96
58 | 9 | 2 | Blacksburg | 19979 | 1 | 0 | 2025/05/30 17:42:07 | -8.972 | -41.66
59 | 9 | 3 | Blacksburg | 19919 | 1 | 0 | 2025/05/30 17:50:54 | -8.971 | -41.945
60 | 9 | 4 | Blacksburg | 19462 | 1 | 0 | 2025/05/30 17:59:42 | -8.92 | -42.316
61 | 10 | 1 | CC4 16apr25 6m | 19958 | 1 | 0 | 2025/05/30 18:08:31 | -6.609 | -27.453
62 | 10 | 2 | CC4 16apr25 6m | 19940 | 1 | 0 | 2025/05/30 18:17:20 | -6.537 | -25.735
63 | 10 | 3 | CC4 16apr25 6m | 20027 | 1 | 0 | 2025/05/30 18:26:10 | -6.517 | -25.13
64 | 10 | 4 | CC4 16apr25 6m | 19929 | 1 | 0 | 2025/05/30 18:35:00 | -6.506 | -25.001
65 | 11 | 1 | CC3 19jun24 1.5m | 20353 | 1 | 0 | 2025/05/30 18:43:49 | -6.016 | -23.125
66 | 11 | 2 | CC3 19jun24 1.5m | 19942 | 1 | 0 | 2025/05/30 18:52:38 | -5.948 | -22.95
67 | 11 | 3 | CC3 19jun24 1.5m | 20051 | 1 | 0 | 2025/05/30 19:01:27 | -5.851 | -23.03
68 | 11 | 4 | CC3 19jun24 1.5m | 20046 | 1 | 0 | 2025/05/30 19:10:16 | -5.971 | -22.911
69 | 12 | 1 | CC2 28oct24 0.1m | 19966 | 1 | 0 | 2025/05/30 19:19:05 | -5.507 | -19.774
70 | 12 | 2 | CC2 28oct24 0.1m | 19895 | 1 | 0 | 2025/05/30 19:27:55 | -5.482 | -19.472
71 | 12 | 3 | CC2 28oct24 0.1m | 19696 | 1 | 0 | 2025/05/30 19:36:44 | -5.497 | -19.697
72 | 12 | 4 | CC2 28oct24 0.1m | 20270 | 1 | 0 | 2025/05/30 19:45:34 | -5.477 | -19.55
73 | 13 | 1 | C50 16sep24 9m | 19973 | 1 | 0 | 2025/05/30 19:54:23 | -6.429 | -24.089
74 | 13 | 2 | C50 16sep24 9m | 19921 | 1 | 0 | 2025/05/30 20:03:12 | -6.527 | -24.735
75 | 13 | 3 | C50 16sep24 9m | 19804 | 1 | 0 | 2025/05/30 20:12:01 | -6.565 | -24.956
76 | 13 | 4 | C50 16sep24 9m | 20042 | 1 | 0 | 2025/05/30 20:20:50 | -6.572 | -24.872
77 | 14 | 1 | CC4 16apr25 9m | 19908 | 1 | 0 | 2025/05/30 20:29:40 | -6.706 | -25.167
78 | 14 | 2 | CC4 16apr25 9m | 19932 | 1 | 0 | 2025/05/30 20:38:29 | -6.728 | -25.43
79 | 14 | 3 | CC4 16apr25 9m | 20080 | 1 | 0 | 2025/05/30 20:47:19 | -6.591 | -25.367
80 | 14 | 4 | CC4 16apr25 9m | 19944 | 1 | 0 | 2025/05/30 20:56:09 | -6.74 | -25.309
81 | 15 | 1 | C50 16apr25 20m | 20208 | 1 | 0 | 2025/05/30 21:04:58 | -6.659 | -25.43
82 | 15 | 2 | C50 16apr25 20m | 19908 | 1 | 0 | 2025/05/30 21:13:49 | -6.675 | -25.826
83 | 15 | 3 | C50 16apr25 20m | 20209 | 1 | 0 | 2025/05/30 21:22:39 | -6.744 | -25.57
84 | 15 | 4 | C50 16apr25 20m | 20005 | 1 | 0 | 2025/05/30 21:31:29 | -6.738 | -25.212
85 | 16 | 1 | CP2 19jun24 BOT | 19896 | 1 | 0 | 2025/05/30 21:40:18 | -6.187 | -23.154
86 | 16 | 2 | CP2 19jun24 BOT | 19851 | 1 | 0 | 2025/05/30 21:49:08 | -6.094 | -23.194
87 | 16 | 3 | CP2 19jun24 BOT | 20239 | 1 | 0 | 2025/05/30 21:57:58 | -6.185 | -23.087
88 | 16 | 4 | CP2 19jun24 BOT | 19523 | 1 | 0 | 2025/05/30 22:06:48 | -6.257 | -22.822
89 | 17 | 1 | CS2 16apr25 0.1m | 19924 | 1 | 0 | 2025/05/30 22:15:38 | -7.923 | -29.619
90 | 17 | 2 | CS2 16apr25 0.1m | 20207 | 1 | 0 | 2025/05/30 22:24:28 | -7.988 | -30.603
91 | 17 | 3 | CS2 16apr25 0.1m | 20043 | 1 | 0 | 2025/05/30 22:33:19 | -7.976 | -30.88
92 | 17 | 4 | CS2 16apr25 0.1m | 20008 | 1 | 0 | 2025/05/30 22:42:09 | -8.041 | -30.61
93 | 18 | 1 | Blacksburg | 19993 | 1 | 0 | 2025/05/30 22:51:00 | -9.091 | -40.714
94 | 18 | 2 | Blacksburg | 20062 | 1 | 0 | 2025/05/30 22:59:50 | -9.173 | -41.839
95 | 18 | 3 | Blacksburg | 20336 | 1 | 0 | 2025/05/30 23:08:39 | -9.158 | -42.242
96 | 18 | 4 | Blacksburg | 20311 | 1 | 0 | 2025/05/30 23:17:30 | -9.133 | -42.125
97 | 19 | 1 | CC2 19jun24 BOT | 20305 | 1 | 0 | 2025/05/30 23:26:19 | -6.414 | -25.188
98 | 19 | 2 | CC2 19jun24 BOT | 20228 | 1 | 0 | 2025/05/30 23:35:10 | -6.22 | -23.487
99 | 19 | 3 | CC2 19jun24 BOT | 20321 | 1 | 0 | 2025/05/30 23:44:02 | -6.12 | -23.314
100 | 19 | 4 | CC2 19jun24 BOT | 19603 | 1 | 0 | 2025/05/30 23:52:52 | -6.131 | -23.289
101 | 20 | 1 | CC4 22may24 1.5m | 19846 | 1 | 0 | 2025/05/31 00:01:45 | -6.454 | -24.727
102 | 20 | 2 | CC4 22may24 1.5m | 19855 | 1 | 0 | 2025/05/31 00:10:35 | -6.438 | -24.846
103 | 20 | 3 | CC4 22may24 1.5m | 20298 | 1 | 0 | 2025/05/31 00:19:26 | -6.452 | -24.88
104 | 20 | 4 | CC4 22may24 1.5m | 19976 | 1 | 0 | 2025/05/31 00:28:17 | -6.424 | -24.705
105 | 21 | 1 | C50 22may24 6m | 19970 | 1 | 0 | 2025/05/31 00:37:08 | -6.601 | -25.506
106 | 21 | 2 | C50 22may24 6m | 20233 | 1 | 0 | 2025/05/31 00:45:59 | -6.629 | -25.607
107 | 21 | 3 | C50 22may24 6m | 20378 | 1 | 0 | 2025/05/31 00:54:50 | -6.631 | -25.502
108 | 21 | 4 | C50 22may24 6m | 19929 | 1 | 0 | 2025/05/31 01:03:40 | -6.601 | -25.599
109 | 22 | 1 | CS1 16apr25 0.1m | 20223 | 1 | 0 | 2025/05/31 01:12:31 | -7.91 | -30.047
110 | 22 | 2 | CS1 16apr25 0.1m | 20354 | 1 | 0 | 2025/05/31 01:21:23 | -7.941 | -30.971
111 | 22 | 3 | CS1 16apr25 0.1m | 19705 | 1 | 0 | 2025/05/31 01:30:14 | -7.982 | -31.171
112 | 22 | 4 | CS1 16apr25 0.1m | 19928 | 1 | 0 | 2025/05/31 01:39:05 | -8.073 | -31.089
113 | 23 | 1 | CC4 19jun24 BOT | 20161 | 1 | 0 | 2025/05/31 01:47:56 | -6.878 | -26.68
114 | 23 | 2 | CC4 19jun24 BOT | 20011 | 1 | 0 | 2025/05/31 01:56:47 | -6.85 | -26.35
115 | 23 | 3 | CC4 19jun24 BOT | 19872 | 1 | 0 | 2025/05/31 02:05:39 | -6.751 | -26.401
116 | 23 | 4 | CC4 19jun24 BOT | 20200 | 1 | 0 | 2025/05/31 02:14:29 | -6.758 | -26.46
117 | 24 | 1 | C50 19jun24 20m | 19747 | 1 | 0 | 2025/05/31 02:23:21 | -6.822 | -26.31
118 | 24 | 2 | C50 19jun24 20m | 19824 | 1 | 0 | 2025/05/31 02:32:13 | -6.842 | -26.19
119 | 24 | 3 | C50 19jun24 20m | 19858 | 1 | 0 | 2025/05/31 02:41:04 | -6.884 | -26.166
120 | 24 | 4 | C50 19jun24 20m | 19858 | 1 | 0 | 2025/05/31 02:49:56 | -6.757 | -26.433
121 | 25 | 1 | CP2 19jun24 BOT - DUP | 20347 | 1 | 0 | 2025/05/31 02:58:48 | -6.139 | -23.044
122 | 25 | 2 | CP2 19jun24 BOT - DUP | 19949 | 1 | 0 | 2025/05/31 03:07:40 | -6.12 | -23.087
123 | 25 | 3 | CP2 19jun24 BOT - DUP | 19901 | 1 | 0 | 2025/05/31 03:16:32 | -6.198 | -22.667
124 | 25 | 4 | CP2 19jun24 BOT - DUP | 20299 | 1 | 0 | 2025/05/31 03:25:24 | -6.079 | -22.815
125 | 26 | 1 | Hawaii | 19801 | 1 | 0 | 2025/05/31 03:34:16 | -5.171 | -8.378
126 | 26 | 2 | Hawaii | 19880 | 1 | 0 | 2025/05/31 03:43:08 | -5.085 | -6.746
127 | 26 | 3 | Hawaii | 20203 | 1 | 0 | 2025/05/31 03:52:01 | -5.202 | -6.227
128 | 26 | 4 | Hawaii | 20051 | 1 | 0 | 2025/05/31 04:00:52 | -5.1 | -6.377
129 | 27 | 1 | Blacksburg | 19615 | 1 | 0 | 2025/05/31 04:09:44 | -8.822 | -37.667
130 | 27 | 2 | Blacksburg | 19809 | 1 | 0 | 2025/05/31 04:18:36 | -9.027 | -41.413
131 | 27 | 3 | Blacksburg | 19857 | 1 | 0 | 2025/05/31 04:27:28 | -9.091 | -41.958
132 | 27 | 4 | Blacksburg | 20291 | 1 | 0 | 2025/05/31 04:36:20 | -9.126 | -42.159

## 18O
Line | Sample | Injection | Sample Name | Correction Type | Good | Ignore | d(18_16) Picarro Raw | Stdev Raw | d(18_16) Memory Corrected | Stdev Memory Corrected | d(18_16) Memory Corrected Drifts | d(18_16) Drift Corrected | Drift Corrected Drifts | d(18_16) Drift Corrected High, Low, and Drift | d(18_16) Reference Defined | d18O Final Corrected Value | Post run corrections:
1 | 1 | 1 | Blacksburg | Drift | 1 | -1 | -9.115 | [9 empty] | (1) Memory correction |  |  | (1) Graphs for memory correction
2 | 1 | 2 | Blacksburg | Drift | 1 | -1 | -8.833 | [9 empty] | Combined Stdev | 0.074545
3 | 1 | 3 | Blacksburg | Drift | 1 | -1 | -8.751 | [9 empty] | Date:
4 | 1 | 4 | Blacksburg | Drift | 1 | -1 | -8.79 | [9 empty] | injection | 18O memory coef
5 | 1 | 5 | Blacksburg | Drift | 1 | -1 | -8.846 | [9 empty] | 1 | 0.9127
6 | 1 | 6 | Blacksburg | Drift | 1 | -1 | -8.816 | [9 empty] | 2 | 0.97183
7 | 1 | 7 | Blacksburg | Drift | 1 | 0 | -8.789 |  | -8.789 |  | -8.789 | -8.7644 | -8.7644 |  |  | -7.7752 | 3 | 0.99052
8 | 1 | 8 | Blacksburg | Drift | 1 | 0 | -8.729 |  | -8.729 |  | -8.729 | -8.7009 | -8.7009 |  |  | -7.7137 | 4 | 0.99351
9 | 1 | 9 | Blacksburg | Drift | 1 | 0 | -8.719 |  | -8.719 |  | -8.719 | -8.6873 | -8.6873 |  |  | -7.7005 | 5 | 0.99626
10 | 1 | 10 | Blacksburg | Drift | 1 | 0 | -8.719 | 0.033665 | -8.719 | 0.033665 | -8.719 | -8.6838 | -8.6838 |  |  | -7.6971 | 6 | 0.99626
11 | 2 | 1 | Myrtle | High Reference | 1 | 0 | -3.698 |  | -3.2596 |  |  | -3.221 |  | -3.221 | -2.2908 | -2.401 | 7 | 0.99677
12 | 2 | 2 | Myrtle | High Reference | 1 | 0 | -3.315 |  | -3.1627 |  |  | -3.1205 |  | -3.1205 | -2.2908 | -2.3036 | 8 | 0.9991
13 | 2 | 3 | Myrtle | High Reference | 1 | 0 | -3.27 |  | -3.2184 |  |  | -3.1726 |  | -3.1726 | -2.2908 | -2.3541 | 9 | 1
14 | 2 | 4 | Myrtle | High Reference | 1 | 0 | -3.195 |  | -3.1591 |  |  | -3.1099 |  | -3.1099 | -2.2908 | -2.2933 | 10 | 1
15 | 2 | 5 | Myrtle | High Reference | 1 | 0 | -3.216 |  | -3.1954 |  |  | -3.1427 |  | -3.1427 | -2.2908 | -2.3251
16 | 2 | 6 | Myrtle | High Reference | 1 | 0 | -3.226 |  | -3.2055 |  |  | -3.1492 |  | -3.1492 | -2.2908 | -2.3314 | (2) Drift correction
17 | 2 | 7 | Myrtle | High Reference | 1 | 0 | -3.186 |  | -3.1681 |  |  | -3.1083 |  | -3.1083 | -2.2908 | -2.2918 | Slope m: | -0.0035174
18 | 2 | 8 | Myrtle | High Reference | 1 | 0 | -3.144 |  | -3.139 |  |  | -3.0757 |  | -3.0757 | -2.2908 | -2.2601 | Stdev Mem Corr. | 0.17947
19 | 2 | 9 | Myrtle | High Reference | 1 | 0 | -3.2 |  | -3.2 |  |  | -3.1332 |  | -3.1332 | -2.2908 | -2.3159 | Stdev Drift Corr. | 0.073075
20 | 2 | 10 | Myrtle | High Reference | 1 | 0 | -3.153 | 0.16202 | -3.153 | 0.036539 |  | -3.0827 |  | -3.0827 | -2.2908 | -2.2669
21 | 3 | 1 | Homer | Low Reference | 1 | 0 | -14.826 |  | -15.8451 |  |  | -15.7712 |  | -15.7712 | -14.5399 | -14.5683 | (3) Normalization to VSMOW scale
22 | 3 | 2 | Homer | Low Reference | 1 | 0 | -15.529 |  | -15.8777 |  |  | -15.8003 |  | -15.8003 | -14.5399 | -14.5964 | slope m | 0.96948
23 | 3 | 3 | Homer | Low Reference | 1 | 0 | -15.726 |  | -15.8452 |  |  | -15.7643 |  | -15.7643 | -14.5399 | -14.5615 | interception t | 0.72169
24 | 3 | 4 | Homer | Low Reference | 1 | 0 | -15.775 |  | -15.857 |  |  | -15.7725 |  | -15.7725 | -14.5399 | -14.5695
25 | 3 | 5 | Homer | Low Reference | 1 | 0 | -15.836 |  | -15.8834 |  |  | -15.7955 |  | -15.7955 | -14.5399 | -14.5918 | (4) Quality control |  | offset
26 | 3 | 6 | Homer | Low Reference | 1 | 0 | -15.812 |  | -15.8594 |  |  | -15.7679 |  | -15.7679 | -14.5399 | -14.565 | QC Check Defined | -3.6108 | accepted ± 0.15
27 | 3 | 7 | Homer | Low Reference | 1 | 0 | -15.766 |  | -15.8068 |  |  | -15.7118 |  | -15.7118 | -14.5399 | -14.5107 | QC Check Vial 5 | -3.5669 | 0.043948
28 | 3 | 8 | Homer | Low Reference | 1 | 0 | -15.791 |  | -15.8024 |  |  | -15.7039 |  | -15.7039 | -14.5399 | -14.503 | QC Check Vial 26 | -3.8 | -0.18923
29 | 3 | 9 | Homer | Low Reference | 1 | 0 | -15.915 |  | -15.915 |  |  | -15.813 |  | -15.813 | -14.5399 | -14.6087
30 | 3 | 10 | Homer | Low Reference | 1 | 0 | -15.796 | 0.31506 | -15.796 | 0.038482 |  | -15.6905 |  | -15.6905 | -14.5399 | -14.49 | Drift Defined | -7.5615 | accepted ± 0.15
31 | 4 | 1 | Blacksburg | Drift | 1 | 0 | -9.146 |  | -8.5654 |  |  | -8.4564 |  | -8.4564 | -7.5615 | -7.4766 | Drift Average | -7.7128 | -0.15131
32 | 4 | 2 | Blacksburg | Drift | 1 | 0 | -8.892 |  | -8.6975 |  |  | -8.5849 |  | -8.5849 | -7.5615 | -7.6013 | Drift Vial 4 | -7.4938 | 0.067645
33 | 4 | 3 | Blacksburg | Drift | 1 | 0 | -8.664 |  | -8.5964 |  |  | -8.4803 |  | -8.4803 | -7.5615 | -7.4998 | Drift Vial 9 | -7.8026 | -0.24111
34 | 4 | 4 | Blacksburg | Drift | 1 | 0 | -8.709 |  | -8.663 |  |  | -8.5434 |  | -8.5434 | -7.5615 | -7.561 | Drift Vial 18 | -7.8502 | -0.28867
35 | 4 | 5 | Blacksburg | Drift | 1 | 0 | -8.641 |  | -8.6142 |  |  | -8.4911 |  | -8.4911 | -7.5615 | -7.5103 | Drift Vial 27 | -7.6957 | -0.13424
36 | 4 | 6 | Blacksburg | Drift | 1 | 0 | -8.726 |  | -8.6996 |  |  | -8.5729 |  | -8.5729 | -7.5615 | -7.5896
37 | 4 | 7 | Blacksburg | Drift | 1 | 0 | -8.591 |  | -8.5677 |  |  | -8.4376 |  | -8.4376 | -7.5615 | -7.4584 | Reference Defined Values |  | d18O
38 | 4 | 8 | Blacksburg | Drift | 1 | 0 | -8.589 |  | -8.5825 |  |  | -8.4488 |  | -8.4488 | -7.5615 | -7.4693 | High Reference | Myrtle Beach | -2.2908
39 | 4 | 9 | Blacksburg | Drift | 1 | 0 | -8.611 |  | -8.611 |  |  | -8.4738 |  | -8.4738 | -7.5615 | -7.4935 | Low Reference | Homer | -14.5399
40 | 4 | 10 | Blacksburg | Drift | 1 | 0 | -8.677 | 0.17253 | -8.677 | 0.052354 |  | -8.5363 |  | -8.5363 | -7.5615 | -7.5541 | Drift | Blacksburg | -7.5615
41 | 5 | 1 | Hawaii | QC Check | 1 | 0 | -4.878 |  | -4.5463 |  |  | -4.4021 |  |  |  | -3.5461 | QC Check | Hawaii | -3.6108
42 | 5 | 2 | Hawaii | QC Check | 1 | 0 | -4.647 |  | -4.5335 |  |  | -4.3857 |  |  |  | -3.5302 | Calibration Date:
43 | 5 | 3 | Hawaii | QC Check | 1 | 0 | -4.595 |  | -4.5563 |  |  | -4.4051 |  |  |  | -3.5489
44 | 5 | 4 | Hawaii | QC Check | 1 | 0 | -4.682 | 0.12361 | -4.6561 | 0.056131 |  | -4.5013 |  |  |  | -3.6422
45 | 6 | 1 | CC4 17feb25 9m |  | 1 | 0 | -5.99 |  | -6.1042 |  |  | -5.9459 |  |  |  | -5.0428
46 | 6 | 2 | CC4 17feb25 9m |  | 1 | 0 | -6.186 |  | -6.2284 |  |  | -6.0666 |  |  |  | -5.1598 | Denali | -19.9287
47 | 6 | 3 | CC4 17feb25 9m |  | 1 | 0 | -6.174 |  | -6.1881 |  |  | -6.0228 |  |  |  | -5.1173 | Alberta | -18.8392
48 | 6 | 4 | CC4 17feb25 9m |  | 1 | 0 | -6.211 | 0.10135 | -6.2209 | 0.056894 |  | -6.0521 |  |  |  | -5.1457 | MT BHB | -18.431
49 | 7 | 1 | C50 22may24 20m |  | 1 | 0 | -6.426 |  | -6.4448 |  |  | -6.2724 |  |  |  | -5.3593 | Anchorage | -16.4204
50 | 7 | 2 | C50 22may24 20m |  | 1 | 0 | -6.429 |  | -6.4351 |  |  | -6.2593 |  |  |  | -5.3466 | Homer | -14.5399
51 | 7 | 3 | C50 22may24 20m |  | 1 | 0 | -6.45 |  | -6.4523 |  |  | -6.2729 |  |  |  | -5.3598 | MT Creek | -13.3562
52 | 7 | 4 | C50 22may24 20m |  | 1 | 0 | -6.545 | 0.056027 | -6.5472 | 0.052029 |  | -6.3643 |  |  |  | -5.4484 | Nickel Mines | -7.876
53 | 8 | 1 | CC2 16apr25 0.1m |  | 1 | 0 | -6.391 |  | -6.3776 |  |  | -6.1911 |  |  |  | -5.2805 | Blacksburg | -7.5615
54 | 8 | 2 | CC2 16apr25 0.1m |  | 1 | 0 | -6.473 |  | -6.471 |  |  | -6.281 |  |  |  | -5.3677 | Denton's Well | -6.87
55 | 8 | 3 | CC2 16apr25 0.1m |  | 1 | 0 | -6.488 |  | -6.4875 |  |  | -6.294 |  |  |  | -5.3802 | Hawaii | -3.6108
56 | 8 | 4 | CC2 16apr25 0.1m |  | 1 | 0 | -6.465 | 0.043231 | -6.4645 | 0.049332 |  | -6.2675 |  |  |  | -5.3546 | Myrtle Beach | -2.2908
57 | 9 | 1 | Blacksburg | Drift | 1 | 0 | -8.815 |  | -9.0202 |  | -9.0202 | -8.8197 | -8.8197 |  |  | -7.8288 | Heavy Oxygen | 3.1753
58 | 9 | 2 | Blacksburg | Drift | 1 | 0 | -8.972 |  | -9.0426 |  | -9.0426 | -8.8386 | -8.8386 |  |  | -7.8472 | Heavy Hydrogen | 11.1372
59 | 9 | 3 | Blacksburg | Drift | 1 | 0 | -8.971 |  | -8.9948 |  | -8.9948 | -8.7872 | -8.7872 |  |  | -7.7974 | Andres Heavy Water | 17.5664
60 | 9 | 4 | Blacksburg | Drift | 1 | 0 | -8.92 | 0.073777 | -8.9359 | 0.045988 | -8.9359 | -8.7249 | -8.7249 |  |  | -7.737 | Picarro ZERO | 0.3
61 | 10 | 1 | CC4 16apr25 6m |  | 1 | 0 | -6.609 |  | -6.4072 |  |  | -6.1927 |  |  |  | -5.282 | Picarro MID | -20.6
62 | 10 | 2 | CC4 16apr25 6m |  | 1 | 0 | -6.537 |  | -6.4699 |  |  | -6.2518 |  |  |  | -5.3393 | Picarro DEPL | -29.6
63 | 10 | 3 | CC4 16apr25 6m |  | 1 | 0 | -6.517 |  | -6.4942 |  |  | -6.2726 |  |  |  | -5.3595
64 | 10 | 4 | CC4 16apr25 6m |  | 1 | 0 | -6.506 | 0.046313 | -6.4903 | 0.040227 |  | -6.2652 |  |  |  | -5.3523
65 | 11 | 1 | CC3 19jun24 1.5m |  | 1 | 0 | -6.016 |  | -5.9732 |  |  | -5.7446 |  |  |  | -4.8476
66 | 11 | 2 | CC3 19jun24 1.5m |  | 1 | 0 | -5.948 |  | -5.9323 |  |  | -5.7001 |  |  |  | -4.8045
67 | 11 | 3 | CC3 19jun24 1.5m |  | 1 | 0 | -5.851 |  | -5.8448 |  |  | -5.6091 |  |  |  | -4.7163
68 | 11 | 4 | CC3 19jun24 1.5m |  | 1 | 0 | -5.971 | 0.069649 | -5.9675 | 0.059275 |  | -5.7283 |  |  |  | -4.8318
69 | 12 | 1 | CC2 28oct24 0.1m |  | 1 | 0 | -5.507 |  | -5.4665 |  |  | -5.2238 |  |  |  | -4.3427
70 | 12 | 2 | CC2 28oct24 0.1m |  | 1 | 0 | -5.482 |  | -5.4682 |  |  | -5.222 |  |  |  | -4.341
71 | 12 | 3 | CC2 28oct24 0.1m |  | 1 | 0 | -5.497 |  | -5.4925 |  |  | -5.2428 |  |  |  | -4.3611
72 | 12 | 4 | CC2 28oct24 0.1m |  | 1 | 0 | -5.477 | 0.013769 | -5.4738 | 0.011917 |  | -5.2205 |  |  |  | -4.3395
73 | 13 | 1 | C50 16sep24 9m |  | 1 | 0 | -6.429 |  | -6.5121 |  |  | -6.2553 |  |  |  | -5.3428
74 | 13 | 2 | C50 16sep24 9m |  | 1 | 0 | -6.527 |  | -6.5566 |  |  | -6.2963 |  |  |  | -5.3825
75 | 13 | 3 | C50 16sep24 9m |  | 1 | 0 | -6.565 |  | -6.5753 |  |  | -6.3115 |  |  |  | -5.3972
76 | 13 | 4 | C50 16sep24 9m |  | 1 | 0 | -6.572 | 0.06587 | -6.5791 | 0.030731 |  | -6.3118 |  |  |  | -5.3975
77 | 14 | 1 | CC4 16apr25 9m |  | 1 | 0 | -6.706 |  | -6.7177 |  |  | -6.4469 |  |  |  | -5.5284
78 | 14 | 2 | CC4 16apr25 9m |  | 1 | 0 | -6.728 |  | -6.7324 |  |  | -6.458 |  |  |  | -5.5393
79 | 14 | 3 | CC4 16apr25 9m |  | 1 | 0 | -6.591 |  | -6.5912 |  |  | -6.3133 |  |  |  | -5.399
80 | 14 | 4 | CC4 16apr25 9m |  | 1 | 0 | -6.74 | 0.0683 | -6.7411 | 0.070274 |  | -6.4597 |  |  |  | -5.5409
81 | 15 | 1 | C50 16apr25 20m |  | 1 | 0 | -6.659 |  | -6.6519 |  |  | -6.367 |  |  |  | -5.451
82 | 15 | 2 | C50 16apr25 20m |  | 1 | 0 | -6.675 |  | -6.6732 |  |  | -6.3847 |  |  |  | -5.4682
83 | 15 | 3 | C50 16apr25 20m |  | 1 | 0 | -6.744 |  | -6.744 |  |  | -6.4521 |  |  |  | -5.5335
84 | 15 | 4 | C50 16apr25 20m |  | 1 | 0 | -6.738 | 0.04329 | -6.738 | 0.04619 |  | -6.4425 |  |  |  | -5.5242
85 | 16 | 1 | CP2 19jun24 BOT |  | 1 | 0 | -6.187 |  | -6.1389 |  |  | -5.8399 |  |  |  | -4.94
86 | 16 | 2 | CP2 19jun24 BOT |  | 1 | 0 | -6.094 |  | -6.0759 |  |  | -5.7734 |  |  |  | -4.8755
87 | 16 | 3 | CP2 19jun24 BOT |  | 1 | 0 | -6.185 |  | -6.1798 |  |  | -5.8737 |  |  |  | -4.9728
88 | 16 | 4 | CP2 19jun24 BOT |  | 1 | 0 | -6.257 | 0.066825 | -6.2539 | 0.074635 |  | -5.9443 |  |  |  | -5.0413
89 | 17 | 1 | CS2 16apr25 0.1m |  | 1 | 0 | -7.923 |  | -8.0684 |  |  | -7.7554 |  |  |  | -6.797
90 | 17 | 2 | CS2 16apr25 0.1m |  | 1 | 0 | -7.988 |  | -8.0368 |  |  | -7.7202 |  |  |  | -6.7629
91 | 17 | 3 | CS2 16apr25 0.1m |  | 1 | 0 | -7.976 |  | -7.9923 |  |  | -7.6722 |  |  |  | -6.7164
92 | 17 | 4 | CS2 16apr25 0.1m |  | 1 | 0 | -8.041 | 0.048422 | -8.0526 | 0.03281 |  | -7.729 |  |  |  | -6.7714
93 | 18 | 1 | Blacksburg | Drift | 1 | 0 | -9.091 |  | -9.1827 |  | -9.1827 | -8.8556 | -8.8556 |  |  | -7.8636
94 | 18 | 2 | Blacksburg | Drift | 1 | 0 | -9.173 |  | -9.2049 |  | -9.2049 | -8.8743 | -8.8743 |  |  | -7.8818
95 | 18 | 3 | Blacksburg | Drift | 1 | 0 | -9.158 |  | -9.1686 |  | -9.1686 | -8.8344 | -8.8344 |  |  | -7.8432
96 | 18 | 4 | Blacksburg | Drift | 1 | 0 | -9.133 | 0.035855 | -9.1401 | 0.027133 | -9.1401 | -8.8024 | -8.8024 |  |  | -7.8121
97 | 19 | 1 | CC2 19jun24 BOT |  | 1 | 0 | -6.414 |  | -6.1766 |  |  | -5.8354 |  |  |  | -4.9357
98 | 19 | 2 | CC2 19jun24 BOT |  | 1 | 0 | -6.22 |  | -6.1379 |  |  | -5.7932 |  |  |  | -4.8948
99 | 19 | 3 | CC2 19jun24 BOT |  | 1 | 0 | -6.12 |  | -6.0914 |  |  | -5.7432 |  |  |  | -4.8463
100 | 19 | 4 | CC2 19jun24 BOT |  | 1 | 0 | -6.131 | 0.13608 | -6.1115 | 0.036803 |  | -5.7598 |  |  |  | -4.8623
101 | 20 | 1 | CC4 22may24 1.5m |  | 1 | 0 | -6.454 |  | -6.4822 |  |  | -6.1269 |  |  |  | -5.2183
102 | 20 | 2 | CC4 22may24 1.5m |  | 1 | 0 | -6.438 |  | -6.4466 |  |  | -6.0879 |  |  |  | -5.1804
103 | 20 | 3 | CC4 22may24 1.5m |  | 1 | 0 | -6.452 |  | -6.455 |  |  | -6.0928 |  |  |  | -5.1851
104 | 20 | 4 | CC4 22may24 1.5m |  | 1 | 0 | -6.424 | 0.013952 | -6.4259 | 0.023311 |  | -6.0601 |  |  |  | -5.1535
105 | 21 | 1 | C50 22may24 6m |  | 1 | 0 | -6.601 |  | -6.6165 |  |  | -6.2471 |  |  |  | -5.3348
106 | 21 | 2 | C50 22may24 6m |  | 1 | 0 | -6.629 |  | -6.6348 |  |  | -6.2619 |  |  |  | -5.3492
107 | 21 | 3 | C50 22may24 6m |  | 1 | 0 | -6.631 |  | -6.633 |  |  | -6.2566 |  |  |  | -5.344
108 | 21 | 4 | C50 22may24 6m |  | 1 | 0 | -6.601 | 0.016763 | -6.6021 | 0.015357 |  | -6.2223 |  |  |  | -5.3107
109 | 22 | 1 | CS1 16apr25 0.1m |  | 1 | 0 | -7.91 |  | -8.0243 |  |  | -7.6409 |  |  |  | -6.686
110 | 22 | 2 | CS1 16apr25 0.1m |  | 1 | 0 | -7.941 |  | -7.9788 |  |  | -7.5918 |  |  |  | -6.6385
111 | 22 | 3 | CS1 16apr25 0.1m |  | 1 | 0 | -7.982 |  | -7.9951 |  |  | -7.6047 |  |  |  | -6.6509
112 | 22 | 4 | CS1 16apr25 0.1m |  | 1 | 0 | -8.073 | 0.07077 | -8.0826 | 0.045657 |  | -7.6886 |  |  |  | -6.7323
113 | 23 | 1 | CC4 19jun24 BOT |  | 1 | 0 | -6.878 |  | -6.7737 |  |  | -6.3762 |  |  |  | -5.4599
114 | 23 | 2 | CC4 19jun24 BOT |  | 1 | 0 | -6.85 |  | -6.8155 |  |  | -6.4146 |  |  |  | -5.4971
115 | 23 | 3 | CC4 19jun24 BOT |  | 1 | 0 | -6.751 |  | -6.7385 |  |  | -6.334 |  |  |  | -5.419
116 | 23 | 4 | CC4 19jun24 BOT |  | 1 | 0 | -6.758 | 0.064309 | -6.7495 | 0.034165 |  | -6.3414 |  |  |  | -5.4262
117 | 24 | 1 | C50 19jun24 20m |  | 1 | 0 | -6.822 |  | -6.8276 |  |  | -6.4161 |  |  |  | -5.4986
118 | 24 | 2 | C50 19jun24 20m |  | 1 | 0 | -6.842 |  | -6.8444 |  |  | -6.4293 |  |  |  | -5.5114
119 | 24 | 3 | C50 19jun24 20m |  | 1 | 0 | -6.884 |  | -6.8852 |  |  | -6.4666 |  |  |  | -5.5476
120 | 24 | 4 | C50 19jun24 20m |  | 1 | 0 | -6.757 | 0.052905 | -6.757 | 0.053479 |  | -6.3349 |  |  |  | -5.4199
121 | 25 | 1 | CP2 19jun24 BOT - DUP |  | 1 | 0 | -6.139 |  | -6.085 |  |  | -5.6594 |  |  |  | -4.765
122 | 25 | 2 | CP2 19jun24 BOT - DUP |  | 1 | 0 | -6.12 |  | -6.1021 |  |  | -5.6729 |  |  |  | -4.7781
123 | 25 | 3 | CP2 19jun24 BOT - DUP |  | 1 | 0 | -6.198 |  | -6.1927 |  |  | -5.7601 |  |  |  | -4.8626
124 | 25 | 4 | CP2 19jun24 BOT - DUP |  | 1 | 0 | -6.079 | 0.049471 | -6.0746 | 0.053935 |  | -5.6384 |  |  |  | -4.7447
125 | 26 | 1 | Hawaii |  | 1 | 0 | -5.171 |  | -5.0917 |  |  | -4.6521 |  |  |  | -3.7884
126 | 26 | 2 | Hawaii |  | 1 | 0 | -5.085 |  | -5.057 |  |  | -4.6138 |  |  |  | -3.7513
127 | 26 | 3 | Hawaii |  | 1 | 0 | -5.202 |  | -5.1937 |  |  | -4.747 |  |  |  | -3.8804
128 | 26 | 4 | Hawaii |  | 1 | 0 | -5.1 | 0.056062 | -5.0936 | 0.058908 |  | -4.6434 |  |  |  | -3.78
129 | 27 | 1 | Blacksburg | Drift | 1 | 0 | -8.822 |  | -9.1469 |  | -9.1469 | -8.6932 | -8.6932 |  |  | -7.7062
130 | 27 | 2 | Blacksburg | Drift | 1 | 0 | -9.027 |  | -9.1376 |  | -9.1376 | -8.6804 | -8.6804 |  |  | -7.6938
131 | 27 | 3 | Blacksburg | Drift | 1 | 0 | -9.091 |  | -9.1288 |  | -9.1288 | -8.6681 | -8.6681 |  |  | -7.6818
132 | 27 | 4 | Blacksburg | Drift | 1 | 0 | -9.126 | 0.13599 | -9.1521 | 0.010302 | -9.1521 | -8.6878 | -8.6878 |  |  | -7.701
Note on G2: Manually change first six injections to ignore "-1".
Note on G3: Manually change first six injections to ignore "-1".
Note on G4: Manually change first six injections to ignore "-1".
Note on G5: Manually change first six injections to ignore "-1".
Note on G6: Manually change first six injections to ignore "-1".
Note on G7: Manually change first six injections to ignore "-1".

## 2H
Line | Sample | Injection | Sample Name | Correction Type | Good | Ignore | dD_H) Picarro Raw | Stdev Raw | d(D_H) Memory Corrected | Stdev Memory Corrected | d(D_H) Memory Corrected Drift | d(D_H) Drift Corrected | Drift Corrected Drift | d(D_H) Drift Corrected High, Low, Drift | d(D_H) Reference Defined | dD Final Corrected Value | Post run corrections:
1 | 1 | 1 | Blacksburg | Drift | 1 | -1 | -49.164 | [9 empty] | (1) Memory correction |  |  | (1) Graphs for memory correction
2 | 1 | 2 | Blacksburg | Drift | 1 | -1 | -44.946 | [9 empty] | Combined Stdev | 0.19285
3 | 1 | 3 | Blacksburg | Drift | 1 | -1 | -44.572 | [9 empty] | Date:
4 | 1 | 4 | Blacksburg | Drift | 1 | -1 | -44.16 | [9 empty] | injection | 18O memory coef
5 | 1 | 5 | Blacksburg | Drift | 1 | -1 | -44.057 | [9 empty] | 1 | 0.84149
6 | 1 | 6 | Blacksburg | Drift | 1 | -1 | -43.958 | [9 empty] | 2 | 0.95967
7 | 1 | 7 | Blacksburg | Drift | 1 | 0 | -43.858 |  | -43.858 |  | -43.858 | -43.9285 | -43.9285 |  |  | -47.7668 | 3 | 0.97903
8 | 1 | 8 | Blacksburg | Drift | 1 | 0 | -43.923 |  | -43.923 |  | -43.923 | -44.0035 | -44.0035 |  |  | -47.8385 | 4 | 0.98712
9 | 1 | 9 | Blacksburg | Drift | 1 | 0 | -44.118 |  | -44.118 |  | -44.118 | -44.2086 | -44.2086 |  |  | -48.0344 | 5 | 0.99035
10 | 1 | 10 | Blacksburg | Drift | 1 | 0 | -43.772 | 0.14713 | -43.772 | 0.14713 | -43.772 | -43.8727 | -43.8727 |  |  | -47.7135 | 6 | 0.99444
11 | 2 | 1 | Myrtle | High Reference | 1 | 0 | -9.712 |  | -4.3133 |  |  | -4.424 |  | -4.424 | -10.453 | -10.0214 | 7 | 0.99798
12 | 2 | 2 | Myrtle | High Reference | 1 | 0 | -5.97 |  | -4.4454 |  |  | -4.5662 |  | -4.5662 | -10.453 | -10.1572 | 8 | 0.9993
13 | 2 | 3 | Myrtle | High Reference | 1 | 0 | -5.228 |  | -4.4198 |  |  | -4.5506 |  | -4.5506 | -10.453 | -10.1423 | 9 | 1
14 | 2 | 4 | Myrtle | High Reference | 1 | 0 | -4.65 |  | -4.1461 |  |  | -4.2871 |  | -4.2871 | -10.453 | -9.8905 | 10 | 1
15 | 2 | 5 | Myrtle | High Reference | 1 | 0 | -4.594 |  | -4.2158 |  |  | -4.3667 |  | -4.3667 | -10.453 | -9.9666
16 | 2 | 6 | Myrtle | High Reference | 1 | 0 | -4.396 |  | -4.1769 |  |  | -4.338 |  | -4.338 | -10.453 | -9.9392 | (2) Drift correction
17 | 2 | 7 | Myrtle | High Reference | 1 | 0 | -4.446 |  | -4.3664 |  |  | -4.5376 |  | -4.5376 | -10.453 | -10.1298 | Slope m: | 0.010066
18 | 2 | 8 | Myrtle | High Reference | 1 | 0 | -4.367 |  | -4.3395 |  |  | -4.5207 |  | -4.5207 | -10.453 | -10.1137 | Stdev Mem Corr. | 0.66864
19 | 2 | 9 | Myrtle | High Reference | 1 | 0 | -4.179 |  | -4.179 |  |  | -4.3702 |  | -4.3702 | -10.453 | -9.97 | Stdev Drift Corr. | 0.47648
20 | 2 | 10 | Myrtle | High Reference | 1 | 0 | -4.425 | 1.6724 | -4.425 | 0.11438 |  | -4.6263 |  | -4.6263 | -10.453 | -10.2147
21 | 3 | 1 | Homer | Low Reference | 1 | 0 | -94.75 |  | -109.067 |  |  | -109.2784 |  | -109.2784 | -110.5169 | -110.2069 | (3) Normalization to VSMOW scale
22 | 3 | 2 | Homer | Low Reference | 1 | 0 | -105.054 |  | -109.1124 |  |  | -109.3339 |  | -109.3339 | -110.5169 | -110.2599 | slope m | 0.95547
23 | 3 | 3 | Homer | Low Reference | 1 | 0 | -106.917 |  | -109.0661 |  |  | -109.2976 |  | -109.2976 | -110.5169 | -110.2253 | interception t | -5.7943
24 | 3 | 4 | Homer | Low Reference | 1 | 0 | -107.625 |  | -108.9542 |  |  | -109.1957 |  | -109.1957 | -110.5169 | -110.1279
25 | 3 | 5 | Homer | Low Reference | 1 | 0 | -108.105 |  | -109.106 |  |  | -109.3576 |  | -109.3576 | -110.5169 | -110.2826 | (4) Quality control |  | offset
26 | 3 | 6 | Homer | Low Reference | 1 | 0 | -108.481 |  | -109.0599 |  |  | -109.3216 |  | -109.3216 | -110.5169 | -110.2481 | QC Check Defined | -13.2294 | accepted ±1.5
27 | 3 | 7 | Homer | Low Reference | 1 | 0 | -108.906 |  | -109.1174 |  |  | -109.3892 |  | -109.3892 | -110.5169 | -110.3127 | QC Check Vial 5 | -13.4214 | -0.19197
28 | 3 | 8 | Homer | Low Reference | 1 | 0 | -109.01 |  | -109.083 |  |  | -109.3648 |  | -109.3648 | -110.5169 | -110.2894 | QC Check Vial 26 | -12.7992 | 0.43017
29 | 3 | 9 | Homer | Low Reference | 1 | 0 | -109.403 |  | -109.403 |  |  | -109.6949 |  | -109.6949 | -110.5169 | -110.6048
30 | 3 | 10 | Homer | Low Reference | 1 | 0 | -108.936 | 4.3978 | -108.936 | 0.12636 |  | -109.238 |  | -109.238 | -110.5169 | -110.1683 | Drift Defined | -47.446 | accepted ±1.5
31 | 4 | 1 | Blacksburg | Drift | 1 | 0 | -52.767 |  | -43.8639 |  |  | -44.1759 |  | -44.1759 | -47.446 | -48.0032 | Drift Average | -47.5732 | -0.12722
32 | 4 | 2 | Blacksburg | Drift | 1 | 0 | -46.38 |  | -43.8571 |  |  | -44.1792 |  | -44.1792 | -47.446 | -48.0063 | Drift Vial 4 | -48.1709 | -0.72485
33 | 4 | 3 | Blacksburg | Drift | 1 | 0 | -45.136 |  | -43.7982 |  |  | -44.1304 |  | -44.1304 | -47.446 | -47.9597 | Drift Vial 9 | -46.855 | 0.59102
34 | 4 | 4 | Blacksburg | Drift | 1 | 0 | -44.613 |  | -43.7846 |  |  | -44.1268 |  | -44.1268 | -47.446 | -47.9563 | Drift Vial 18 | -47.1595 | 0.28652
35 | 4 | 5 | Blacksburg | Drift | 1 | 0 | -44.607 |  | -43.9859 |  |  | -44.3382 |  | -44.3382 | -47.446 | -48.1583 | Drift Vial 27 | -47.8425 | -0.39651
36 | 4 | 6 | Blacksburg | Drift | 1 | 0 | -44.295 |  | -43.9354 |  |  | -44.2978 |  | -44.2978 | -47.446 | -48.1196
37 | 4 | 7 | Blacksburg | Drift | 1 | 0 | -44.044 |  | -43.9127 |  |  | -44.2852 |  | -44.2852 | -47.446 | -48.1076 | Reference Defined Values |  | d2H
38 | 4 | 8 | Blacksburg | Drift | 1 | 0 | -43.919 |  | -43.8736 |  |  | -44.2561 |  | -44.2561 | -47.446 | -48.0799 | High Reference | North Myrtle Beach, SC | -10.453
39 | 4 | 9 | Blacksburg | Drift | 1 | 0 | -44.061 |  | -44.061 |  |  | -44.4536 |  | -44.4536 | -47.446 | -48.2685 | Low Reference | Homer, AK | -110.5169
40 | 4 | 10 | Blacksburg | Drift | 1 | 0 | -44.008 | 2.6984 | -44.008 | 0.090229 |  | -44.4106 |  | -44.4106 | -47.446 | -48.2275 | Drift | Blacksburg | -47.446
41 | 5 | 1 | Hawaii | QC Check | 1 | 0 | -12.652 |  | -7.6819 |  |  | -8.0946 |  |  |  | -13.5285 | QC Check | Hawaii | -13.2294
42 | 5 | 2 | Hawaii | QC Check | 1 | 0 | -8.963 |  | -7.5496 |  |  | -7.9724 |  |  |  | -13.4117 | Calibration Date:
43 | 5 | 3 | Hawaii | QC Check | 1 | 0 | -8.292 |  | -7.5431 |  |  | -7.9759 |  |  |  | -13.4151
44 | 5 | 4 | Hawaii | QC Check | 1 | 0 | -7.909 | 2.176 | -7.4441 | 0.097604 |  | -7.887 |  |  |  | -13.3301
45 | 6 | 1 | CC4 17feb25 9m |  | 1 | 0 | -24.127 |  | -26.6976 |  |  | -27.1506 |  |  |  | -31.736
46 | 6 | 2 | CC4 17feb25 9m |  | 1 | 0 | -25.301 |  | -26.0024 |  |  | -26.4655 |  |  |  | -31.0814 | Denali National Park | -153.7172
47 | 6 | 3 | CC4 17feb25 9m |  | 1 | 0 | -25.765 |  | -26.1394 |  |  | -26.6125 |  |  |  | -31.2219 | Edmonton, Alberta | -146.8245
48 | 6 | 4 | CC4 17feb25 9m |  | 1 | 0 | -25.554 | 0.73152 | -25.7813 | 0.39058 |  | -26.2644 |  |  |  | -30.8893 | MT BHB | -140.5778
49 | 7 | 1 | C50 22may24 20m |  | 1 | 0 | -26.612 |  | -26.7797 |  |  | -27.2729 |  |  |  | -31.8529 | Anchorage, AK | -129.9028
50 | 7 | 2 | C50 22may24 20m |  | 1 | 0 | -26.738 |  | -26.7858 |  |  | -27.289 |  |  |  | -31.8683 | Homer, AK | -110.5169
51 | 7 | 3 | C50 22may24 20m |  | 1 | 0 | -26.686 |  | -26.7097 |  |  | -27.2231 |  |  |  | -31.8053 | MT Creek | -97.7428
52 | 7 | 4 | C50 22may24 20m |  | 1 | 0 | -26.607 | 0.062904 | -26.6206 | 0.077068 |  | -27.144 |  |  |  | -31.7297 | Nickel Mines, PA | -48.7951
53 | 8 | 1 | CC2 16apr25 0.1m |  | 1 | 0 | -25.127 |  | -24.8924 |  |  | -25.4259 |  |  |  | -30.0881 | Blacksburg | -47.446
54 | 8 | 2 | CC2 16apr25 0.1m |  | 1 | 0 | -24.991 |  | -24.9258 |  |  | -25.4694 |  |  |  | -30.1296 | Denton's Well | -41.771
55 | 8 | 3 | CC2 16apr25 0.1m |  | 1 | 0 | -24.864 |  | -24.8275 |  |  | -25.3811 |  |  |  | -30.0453 | Hawaii | -13.2294
56 | 8 | 4 | CC2 16apr25 0.1m |  | 1 | 0 | -24.833 | 0.13421 | -24.8102 | 0.05436 |  | -25.3738 |  |  |  | -30.0383 | North Myrtle Beach, SC | -10.453
57 | 9 | 1 | Blacksburg | Drift | 1 | 0 | -39.96 |  | -42.3577 |  | -42.3577 | -42.9315 | -42.9315 |  |  | -46.8142 | WQL Heavy Oxygen | -15.7881
58 | 9 | 2 | Blacksburg | Drift | 1 | 0 | -41.66 |  | -42.3386 |  | -42.3386 | -42.9225 | -42.9225 |  |  | -46.8056 | WQL Heavy Hydrogen | 4.0332
59 | 9 | 3 | Blacksburg | Drift | 1 | 0 | -41.945 |  | -42.3038 |  | -42.3038 | -42.8977 | -42.8977 |  |  | -46.7819 | Andres Heavy Water | 20.5382
60 | 9 | 4 | Blacksburg | Drift | 1 | 0 | -42.316 | 1.042 | -42.5412 | 0.10626 | -42.5412 | -43.1451 | -43.1451 |  |  | -47.0183
61 | 10 | 1 | CC4 16apr25 6m |  | 1 | 0 | -27.453 |  | -25.0971 |  |  | -25.7111 |  |  |  | -30.3606
62 | 10 | 2 | CC4 16apr25 6m |  | 1 | 0 | -25.735 |  | -25.0663 |  |  | -25.6904 |  |  |  | -30.3408
63 | 10 | 3 | CC4 16apr25 6m |  | 1 | 0 | -25.13 |  | -24.7696 |  |  | -25.4038 |  |  |  | -30.0669
64 | 10 | 4 | CC4 16apr25 6m |  | 1 | 0 | -25.001 | 1.1285 | -24.778 | 0.17824 |  | -25.4222 |  |  |  | -30.0846
65 | 11 | 1 | CC3 19jun24 1.5m |  | 1 | 0 | -23.125 |  | -22.8276 |  |  | -23.4819 |  |  |  | -28.2307
66 | 11 | 2 | CC3 19jun24 1.5m |  | 1 | 0 | -22.95 |  | -22.8673 |  |  | -23.5316 |  |  |  | -28.2782
67 | 11 | 3 | CC3 19jun24 1.5m |  | 1 | 0 | -23.03 |  | -22.9887 |  |  | -23.6631 |  |  |  | -28.4038
68 | 11 | 4 | CC3 19jun24 1.5m |  | 1 | 0 | -22.911 | 0.094661 | -22.8841 | 0.068704 |  | -23.5686 |  |  |  | -28.3134
69 | 12 | 1 | CC2 28oct24 0.1m |  | 1 | 0 | -19.774 |  | -19.2768 |  |  | -19.9713 |  |  |  | -24.8764
70 | 12 | 2 | CC2 28oct24 0.1m |  | 1 | 0 | -19.472 |  | -19.3333 |  |  | -20.0379 |  |  |  | -24.94
71 | 12 | 3 | CC2 28oct24 0.1m |  | 1 | 0 | -19.697 |  | -19.6296 |  |  | -20.3443 |  |  |  | -25.2327
72 | 12 | 4 | CC2 28oct24 0.1m |  | 1 | 0 | -19.55 | 0.13712 | -19.5067 | 0.16164 |  | -20.2314 |  |  |  | -25.1249
73 | 13 | 1 | C50 16sep24 9m |  | 1 | 0 | -24.089 |  | -24.8085 |  |  | -25.5433 |  |  |  | -30.2002
74 | 13 | 2 | C50 16sep24 9m |  | 1 | 0 | -24.735 |  | -24.9441 |  |  | -25.689 |  |  |  | -30.3395
75 | 13 | 3 | C50 16sep24 9m |  | 1 | 0 | -24.956 |  | -25.0694 |  |  | -25.8243 |  |  |  | -30.4687
76 | 13 | 4 | C50 16sep24 9m |  | 1 | 0 | -24.872 | 0.39336 | -24.9405 | 0.10654 |  | -25.7055 |  |  |  | -30.3553
77 | 14 | 1 | CC4 16apr25 9m |  | 1 | 0 | -25.167 |  | -25.2138 |  |  | -25.9888 |  |  |  | -30.6259
78 | 14 | 2 | CC4 16apr25 9m |  | 1 | 0 | -25.43 |  | -25.4525 |  |  | -26.2376 |  |  |  | -30.8637
79 | 14 | 3 | CC4 16apr25 9m |  | 1 | 0 | -25.367 |  | -25.3774 |  |  | -26.1726 |  |  |  | -30.8015
80 | 14 | 4 | CC4 16apr25 9m |  | 1 | 0 | -25.309 | 0.11229 | -25.3146 | 0.10105 |  | -26.1199 |  |  |  | -30.7512
81 | 15 | 1 | C50 16apr25 20m |  | 1 | 0 | -25.43 |  | -25.4492 |  |  | -26.2645 |  |  |  | -30.8894
82 | 15 | 2 | C50 16apr25 20m |  | 1 | 0 | -25.826 |  | -25.8469 |  |  | -26.6722 |  |  |  | -31.2789
83 | 15 | 3 | C50 16apr25 20m |  | 1 | 0 | -25.57 |  | -25.5755 |  |  | -26.4109 |  |  |  | -31.0293
84 | 15 | 4 | C50 16apr25 20m |  | 1 | 0 | -25.212 | 0.25733 | -25.2108 | 0.26493 |  | -26.0563 |  |  |  | -30.6904
85 | 16 | 1 | CP2 19jun24 BOT |  | 1 | 0 | -23.154 |  | -22.8278 |  |  | -23.6834 |  |  |  | -28.4232
86 | 16 | 2 | CP2 19jun24 BOT |  | 1 | 0 | -23.194 |  | -23.1126 |  |  | -23.9783 |  |  |  | -28.7049
87 | 16 | 3 | CP2 19jun24 BOT |  | 1 | 0 | -23.087 |  | -23.0424 |  |  | -23.9182 |  |  |  | -28.6475
88 | 16 | 4 | CP2 19jun24 BOT |  | 1 | 0 | -22.822 | 0.16742 | -22.7912 | 0.15808 |  | -23.677 |  |  |  | -28.4171
89 | 17 | 1 | CS2 16apr25 0.1m |  | 1 | 0 | -29.619 |  | -30.6964 |  |  | -31.5922 |  |  |  | -35.9798
90 | 17 | 2 | CS2 16apr25 0.1m |  | 1 | 0 | -30.603 |  | -30.9168 |  |  | -31.8227 |  |  |  | -36.2001
91 | 17 | 3 | CS2 16apr25 0.1m |  | 1 | 0 | -30.88 |  | -31.049 |  |  | -31.965 |  |  |  | -36.336
92 | 17 | 4 | CS2 16apr25 0.1m |  | 1 | 0 | -30.61 | 0.55454 | -30.7103 | 0.17027 |  | -31.6364 |  |  |  | -36.022
93 | 18 | 1 | Blacksburg | Drift | 1 | 0 | -40.714 |  | -42.3155 |  | -42.3155 | -43.2517 | -43.2517 |  |  | -47.1201
94 | 18 | 2 | Blacksburg | Drift | 1 | 0 | -41.839 |  | -42.2919 |  | -42.2919 | -43.2381 | -43.2381 |  |  | -47.1071
95 | 18 | 3 | Blacksburg | Drift | 1 | 0 | -42.242 |  | -42.4859 |  | -42.4859 | -43.4422 | -43.4422 |  |  | -47.3021
96 | 18 | 4 | Blacksburg | Drift | 1 | 0 | -42.125 | 0.69817 | -42.2733 | 0.097709 | -42.2733 | -43.2396 | -43.2396 |  |  | -47.1086
97 | 19 | 1 | CC2 19jun24 BOT |  | 1 | 0 | -25.188 |  | -22.5034 |  |  | -23.4798 |  |  |  | -28.2286
98 | 19 | 2 | CC2 19jun24 BOT |  | 1 | 0 | -23.487 |  | -22.7353 |  |  | -23.7218 |  |  |  | -28.4598
99 | 19 | 3 | CC2 19jun24 BOT |  | 1 | 0 | -23.314 |  | -22.9196 |  |  | -23.9161 |  |  |  | -28.6455
100 | 19 | 4 | CC2 19jun24 BOT |  | 1 | 0 | -23.289 | 0.91657 | -23.0464 | 0.23605 |  | -24.053 |  |  |  | -28.7763
101 | 20 | 1 | CC4 22may24 1.5m |  | 1 | 0 | -24.727 |  | -24.9549 |  |  | -25.9716 |  |  |  | -30.6095
102 | 20 | 2 | CC4 22may24 1.5m |  | 1 | 0 | -24.846 |  | -24.9088 |  |  | -25.9355 |  |  |  | -30.575
103 | 20 | 3 | CC4 22may24 1.5m |  | 1 | 0 | -24.88 |  | -24.9134 |  |  | -25.9501 |  |  |  | -30.589
104 | 20 | 4 | CC4 22may24 1.5m |  | 1 | 0 | -24.705 | 0.086466 | -24.7232 | 0.10334 |  | -25.7701 |  |  |  | -30.4169
105 | 21 | 1 | C50 22may24 6m |  | 1 | 0 | -25.506 |  | -25.633 |  |  | -26.6899 |  |  |  | -31.2958
106 | 21 | 2 | C50 22may24 6m |  | 1 | 0 | -25.607 |  | -25.6434 |  |  | -26.7104 |  |  |  | -31.3153
107 | 21 | 3 | C50 22may24 6m |  | 1 | 0 | -25.502 |  | -25.5187 |  |  | -26.5958 |  |  |  | -31.2059
108 | 21 | 4 | C50 22may24 6m |  | 1 | 0 | -25.599 | 0.057274 | -25.6105 | 0.0568 |  | -26.6976 |  |  |  | -31.3032
109 | 22 | 1 | CS1 16apr25 0.1m |  | 1 | 0 | -30.047 |  | -30.752 |  |  | -31.8492 |  |  |  | -36.2254
110 | 22 | 2 | CS1 16apr25 0.1m |  | 1 | 0 | -30.971 |  | -31.1877 |  |  | -32.2949 |  |  |  | -36.6512
111 | 22 | 3 | CS1 16apr25 0.1m |  | 1 | 0 | -31.171 |  | -31.2878 |  |  | -32.4051 |  |  |  | -36.7566
112 | 22 | 4 | CS1 16apr25 0.1m |  | 1 | 0 | -31.089 | 0.5215 | -31.1597 | 0.23634 |  | -32.2871 |  |  |  | -36.6438
113 | 23 | 1 | CC4 19jun24 BOT |  | 1 | 0 | -26.68 |  | -25.9811 |  |  | -27.1186 |  |  |  | -31.7054
114 | 23 | 2 | CC4 19jun24 BOT |  | 1 | 0 | -26.35 |  | -26.1589 |  |  | -27.3064 |  |  |  | -31.8848
115 | 23 | 3 | CC4 19jun24 BOT |  | 1 | 0 | -26.401 |  | -26.3027 |  |  | -27.4603 |  |  |  | -32.0319
116 | 23 | 4 | CC4 19jun24 BOT |  | 1 | 0 | -26.46 | 0.14529 | -26.4004 | 0.18241 |  | -27.568 |  |  |  | -32.1348
117 | 24 | 1 | C50 19jun24 20m |  | 1 | 0 | -26.31 |  | -26.2862 |  |  | -27.4639 |  |  |  | -32.0354
118 | 24 | 2 | C50 19jun24 20m |  | 1 | 0 | -26.19 |  | -26.1791 |  |  | -27.3669 |  |  |  | -31.9426
119 | 24 | 3 | C50 19jun24 20m |  | 1 | 0 | -26.166 |  | -26.1598 |  |  | -27.3577 |  |  |  | -31.9338
120 | 24 | 4 | C50 19jun24 20m |  | 1 | 0 | -26.433 | 0.12287 | -26.4327 | 0.12516 |  | -27.6405 |  |  |  | -32.2041
121 | 25 | 1 | CP2 19jun24 BOT - DUP |  | 1 | 0 | -23.044 |  | -22.5068 |  |  | -23.7248 |  |  |  | -28.4627
122 | 25 | 2 | CP2 19jun24 BOT - DUP |  | 1 | 0 | -23.087 |  | -22.9521 |  |  | -24.1801 |  |  |  | -28.8977
123 | 25 | 3 | CP2 19jun24 BOT - DUP |  | 1 | 0 | -22.667 |  | -22.588 |  |  | -23.8261 |  |  |  | -28.5595
124 | 25 | 4 | CP2 19jun24 BOT - DUP |  | 1 | 0 | -22.815 | 0.19763 | -22.7684 | 0.19833 |  | -24.0166 |  |  |  | -28.7415
125 | 26 | 1 | Hawaii |  | 1 | 0 | -8.378 |  | -6.0897 |  |  | -7.3479 |  |  |  | -12.815
126 | 26 | 2 | Hawaii |  | 1 | 0 | -6.746 |  | -6.0979 |  |  | -7.3662 |  |  |  | -12.8326
127 | 26 | 3 | Hawaii |  | 1 | 0 | -6.227 |  | -5.8792 |  |  | -7.1575 |  |  |  | -12.6332
128 | 26 | 4 | Hawaii |  | 1 | 0 | -6.377 | 0.98836 | -6.1653 | 0.12395 |  | -7.4537 |  |  |  | -12.9161
129 | 27 | 1 | Blacksburg | Drift | 1 | 0 | -37.667 |  | -42.6266 |  | -42.6266 | -43.9251 | -43.9251 |  |  | -47.7636
130 | 27 | 2 | Blacksburg | Drift | 1 | 0 | -41.413 |  | -42.826 |  | -42.826 | -44.1346 | -44.1346 |  |  | -47.9637
131 | 27 | 3 | Blacksburg | Drift | 1 | 0 | -41.958 |  | -42.7041 |  | -42.7041 | -44.0227 | -44.0227 |  |  | -47.8568
132 | 27 | 4 | Blacksburg | Drift | 1 | 0 | -42.159 | 2.1118 | -42.6199 | 0.095857 | -42.6199 | -43.9485 | -43.9485 |  |  | -47.786
Note on G2: Manually change first six injections to ignore "-1".
Note on G3: Manually change first six injections to ignore "-1".
Note on G4: Manually change first six injections to ignore "-1".
Note on G5: Manually change first six injections to ignore "-1".
Note on G6: Manually change first six injections to ignore "-1".
Note on G7: Manually change first six injections to ignore "-1".

## RESULTS
| Analyst: | DWH |  |  |  | Complete Sequence |  |  | Run Summary | 18O | 18O | 2H | 2H
| Analysis Date: | 2025-05-30 |  |  |  | Line | Inj Number | Identifier | Ignore | Ignore | d18O Final Value | d2H Final Value | Ignore
| Project: | CCR |  |  |  | 1 | 1 | Blacksburg | 0 | -1 |  |  | -1
[6 empty] | 2 | 2 | Blacksburg | 0 | -1 |  |  | -1
[6 empty] | 3 | 3 | Blacksburg | 0 | -1 |  |  | -1
Vial | Sample Name | d18O | 1 s.d. | d2H | 1 s.d. | 4 | 4 | Blacksburg | 0 | -1 |  |  | -1
|  | ‰ vs. VSMOW |  | ‰ vs. VSMOW |  | 5 | 5 | Blacksburg | 0 | -1 |  |  | -1
1 | Blacksburg | -7.7216 | 0.036429 | -47.8383 | 0.14044 | 6 | 6 | Blacksburg | 0 | -1 |  |  | -1
2 | Myrtle | -2.3143 | 0.041838 | -10.0545 | 0.11018 | 7 | 7 | Blacksburg | 0 | 0 | -7.7752 | -47.7668 | 0
3 | Homer | -14.5565 | 0.041323 | -110.2726 | 0.12975 | 8 | 8 | Blacksburg | 0 | 0 | -7.7137 | -47.8385 | 0
4 | Blacksburg | -7.5214 | 0.051337 | -48.0887 | 0.10857 | 9 | 9 | Blacksburg | 0 | 0 | -7.7005 | -48.0344 | 0
5 | Hawaii | -3.5669 | 0.05092 | -13.4214 | 0.081522 | 10 | 10 | Blacksburg | 0 | 0 | -7.6971 | -47.7135 | 0
6 | CC4 17feb25 9m | -5.1164 | 0.052155 | -31.2321 | 0.36253 | 11 | 1 | Myrtle | 0 | 0 | -2.401 | -10.0214 | 0
7 | C50 22may24 20m | -5.3785 | 0.046971 | -31.814 | 0.0623 | 12 | 2 | Myrtle | 0 | 0 | -2.3036 | -10.1572 | 0
8 | CC2 16apr25 0.1m | -5.3457 | 0.044735 | -30.0753 | 0.042363 | 13 | 3 | Myrtle | 0 | 0 | -2.3541 | -10.1423 | 0
9 | Blacksburg | -7.8026 | 0.048357 | -46.855 | 0.10973 | 14 | 4 | Myrtle | 0 | 0 | -2.2933 | -9.8905 | 0
10 | CC4 16apr25 6m | -5.3333 | 0.035197 | -30.2132 | 0.15911 | 15 | 5 | Myrtle | 0 | 0 | -2.3251 | -9.9666 | 0
11 | CC3 19jun24 1.5m | -4.8001 | 0.058626 | -28.3065 | 0.07317 | 16 | 6 | Myrtle | 0 | 0 | -2.3314 | -9.9392 | 0
12 | CC2 28oct24 0.1m | -4.3461 | 0.010097 | -25.0435 | 0.1644 | 17 | 7 | Myrtle | 0 | 0 | -2.2918 | -10.1298 | 0
13 | C50 16sep24 9m | -5.38 | 0.025785 | -30.3409 | 0.11006 | 18 | 8 | Myrtle | 0 | 0 | -2.2601 | -10.1137 | 0
14 | CC4 16apr25 9m | -5.5019 | 0.068842 | -30.7606 | 0.10086 | 19 | 9 | Myrtle | 0 | 0 | -2.3159 | -9.97 | 0
15 | C50 16apr25 20m | -5.4942 | 0.040768 | -30.972 | 0.2474 | 20 | 10 | Myrtle | 0 | 0 | -2.2669 | -10.2147 | 0
16 | CP2 19jun24 BOT | -4.9574 | 0.068996 | -28.5482 | 0.14972 | 21 | 1 | Homer | 0 | 0 | -14.5683 | -110.2069 | 0
17 | CS2 16apr25 0.1m | -6.762 | 0.033655 | -36.1345 | 0.16479 | 22 | 2 | Homer | 0 | 0 | -14.5964 | -110.2599 | 0
18 | Blacksburg | -7.8502 | 0.029868 | -47.1595 | 0.095269 | 23 | 3 | Homer | 0 | 0 | -14.5615 | -110.2253 | 0
19 | CC2 19jun24 BOT | -4.8848 | 0.039483 | -28.5276 | 0.23786 | 24 | 4 | Homer | 0 | 0 | -14.5695 | -110.1279 | 0
20 | CC4 22may24 1.5m | -5.1843 | 0.026589 | -30.5476 | 0.088247 | 25 | 5 | Homer | 0 | 0 | -14.5918 | -110.2826 | 0
21 | C50 22may24 6m | -5.3347 | 0.017039 | -31.28 | 0.050113 | 26 | 6 | Homer | 0 | 0 | -14.565 | -110.2481 | 0
22 | CS1 16apr25 0.1m | -6.6769 | 0.042047 | -36.5692 | 0.23495 | 27 | 7 | Homer | 0 | 0 | -14.5107 | -110.3127 | 0
23 | CC4 19jun24 BOT | -5.4506 | 0.035797 | -31.9392 | 0.18661 | 28 | 8 | Homer | 0 | 0 | -14.503 | -110.2894 | 0
24 | C50 19jun24 20m | -5.4944 | 0.053813 | -32.029 | 0.12546 | 29 | 9 | Homer | 0 | 0 | -14.6087 | -110.6048 | 0
25 | CP2 19jun24 BOT - DUP | -4.7876 | 0.051848 | -28.6654 | 0.19327 | 30 | 10 | Homer | 0 | 0 | -14.49 | -110.1683 | 0
26 | Hawaii | -3.8 | 0.055893 | -12.7992 | 0.11918 | 31 | 1 | Blacksburg | 0 | 0 | -7.4766 | -48.0032 | 0
27 | Blacksburg | -7.6957 | 0.010567 | -47.8425 | 0.090038 | 32 | 2 | Blacksburg | 0 | 0 | -7.6013 | -48.0063 | 0
[6 empty] | 33 | 3 | Blacksburg | 0 | 0 | -7.4998 | -47.9597 | 0
[6 empty] | 34 | 4 | Blacksburg | 0 | 0 | -7.561 | -47.9563 | 0
[6 empty] | 35 | 5 | Blacksburg | 0 | 0 | -7.5103 | -48.1583 | 0
| Quality Assurance Verification |  |  |  |  | 36 | 6 | Blacksburg | 0 | 0 | -7.5896 | -48.1196 | 0
| Metric |  | Reviewer 1 | Reviewer 2 |  | 37 | 7 | Blacksburg | 0 | 0 | -7.4584 | -48.1076 | 0
| d18O Cal Curve R^2 | 0.999+ | pass |  |  | 38 | 8 | Blacksburg | 0 | 0 | -7.4693 | -48.0799 | 0
| d18O QC checks | < 0.15 | pass |  |  | 39 | 9 | Blacksburg | 0 | 0 | -7.4935 | -48.2685 | 0
| dD Cal Curve R^2 | 0.999+ | pass |  |  | 40 | 10 | Blacksburg | 0 | 0 | -7.5541 | -48.2275 | 0
| dD QC checks | < 1.5 | pass |  |  | 41 | 1 | Hawaii | 0 | 0 | -3.5461 | -13.5285 | 0
| Initials |  | DWH |  |  | 42 | 2 | Hawaii | 0 | 0 | -3.5302 | -13.4117 | 0
| Date |  | 2025-06-03 |  |  | 43 | 3 | Hawaii | 0 | 0 | -3.5489 | -13.4151 | 0
| Comments: |  |  |  |  | 44 | 4 | Hawaii | 0 | 0 | -3.6422 | -13.3301 | 0
[6 empty] | 45 | 1 | CC4 17feb25 9m | 0 | 0 | -5.0428 | -31.736 | 0
[6 empty] | 46 | 2 | CC4 17feb25 9m | 0 | 0 | -5.1598 | -31.0814 | 0
[6 empty] | 47 | 3 | CC4 17feb25 9m | 0 | 0 | -5.1173 | -31.2219 | 0
[6 empty] | 48 | 4 | CC4 17feb25 9m | 0 | 0 | -5.1457 | -30.8893 | 0
[6 empty] | 49 | 1 | C50 22may24 20m | 0 | 0 | -5.3593 | -31.8529 | 0
[6 empty] | 50 | 2 | C50 22may24 20m | 0 | 0 | -5.3466 | -31.8683 | 0
[6 empty] | 51 | 3 | C50 22may24 20m | 0 | 0 | -5.3598 | -31.8053 | 0
[6 empty] | 52 | 4 | C50 22may24 20m | 0 | 0 | -5.4484 | -31.7297 | 0
[6 empty] | 53 | 1 | CC2 16apr25 0.1m | 0 | 0 | -5.2805 | -30.0881 | 0
[6 empty] | 54 | 2 | CC2 16apr25 0.1m | 0 | 0 | -5.3677 | -30.1296 | 0
[6 empty] | 55 | 3 | CC2 16apr25 0.1m | 0 | 0 | -5.3802 | -30.0453 | 0
[6 empty] | 56 | 4 | CC2 16apr25 0.1m | 0 | 0 | -5.3546 | -30.0383 | 0
[6 empty] | 57 | 1 | Blacksburg | 0 | 0 | -7.8288 | -46.8142 | 0
[6 empty] | 58 | 2 | Blacksburg | 0 | 0 | -7.8472 | -46.8056 | 0
[6 empty] | 59 | 3 | Blacksburg | 0 | 0 | -7.7974 | -46.7819 | 0
[6 empty] | 60 | 4 | Blacksburg | 0 | 0 | -7.737 | -47.0183 | 0
[6 empty] | 61 | 1 | CC4 16apr25 6m | 0 | 0 | -5.282 | -30.3606 | 0
[6 empty] | 62 | 2 | CC4 16apr25 6m | 0 | 0 | -5.3393 | -30.3408 | 0
[6 empty] | 63 | 3 | CC4 16apr25 6m | 0 | 0 | -5.3595 | -30.0669 | 0
[6 empty] | 64 | 4 | CC4 16apr25 6m | 0 | 0 | -5.3523 | -30.0846 | 0
[6 empty] | 65 | 1 | CC3 19jun24 1.5m | 0 | 0 | -4.8476 | -28.2307 | 0
[6 empty] | 66 | 2 | CC3 19jun24 1.5m | 0 | 0 | -4.8045 | -28.2782 | 0
[6 empty] | 67 | 3 | CC3 19jun24 1.5m | 0 | 0 | -4.7163 | -28.4038 | 0
[6 empty] | 68 | 4 | CC3 19jun24 1.5m | 0 | 0 | -4.8318 | -28.3134 | 0
[6 empty] | 69 | 1 | CC2 28oct24 0.1m | 0 | 0 | -4.3427 | -24.8764 | 0
[6 empty] | 70 | 2 | CC2 28oct24 0.1m | 0 | 0 | -4.341 | -24.94 | 0
[6 empty] | 71 | 3 | CC2 28oct24 0.1m | 0 | 0 | -4.3611 | -25.2327 | 0
[6 empty] | 72 | 4 | CC2 28oct24 0.1m | 0 | 0 | -4.3395 | -25.1249 | 0
[6 empty] | 73 | 1 | C50 16sep24 9m | 0 | 0 | -5.3428 | -30.2002 | 0
[6 empty] | 74 | 2 | C50 16sep24 9m | 0 | 0 | -5.3825 | -30.3395 | 0
[6 empty] | 75 | 3 | C50 16sep24 9m | 0 | 0 | -5.3972 | -30.4687 | 0
[6 empty] | 76 | 4 | C50 16sep24 9m | 0 | 0 | -5.3975 | -30.3553 | 0
[6 empty] | 77 | 1 | CC4 16apr25 9m | 0 | 0 | -5.5284 | -30.6259 | 0
[6 empty] | 78 | 2 | CC4 16apr25 9m | 0 | 0 | -5.5393 | -30.8637 | 0
[6 empty] | 79 | 3 | CC4 16apr25 9m | 0 | 0 | -5.399 | -30.8015 | 0
[6 empty] | 80 | 4 | CC4 16apr25 9m | 0 | 0 | -5.5409 | -30.7512 | 0
[6 empty] | 81 | 1 | C50 16apr25 20m | 0 | 0 | -5.451 | -30.8894 | 0
[6 empty] | 82 | 2 | C50 16apr25 20m | 0 | 0 | -5.4682 | -31.2789 | 0
[6 empty] | 83 | 3 | C50 16apr25 20m | 0 | 0 | -5.5335 | -31.0293 | 0
[6 empty] | 84 | 4 | C50 16apr25 20m | 0 | 0 | -5.5242 | -30.6904 | 0
[6 empty] | 85 | 1 | CP2 19jun24 BOT | 0 | 0 | -4.94 | -28.4232 | 0
[6 empty] | 86 | 2 | CP2 19jun24 BOT | 0 | 0 | -4.8755 | -28.7049 | 0
[6 empty] | 87 | 3 | CP2 19jun24 BOT | 0 | 0 | -4.9728 | -28.6475 | 0
[6 empty] | 88 | 4 | CP2 19jun24 BOT | 0 | 0 | -5.0413 | -28.4171 | 0
[6 empty] | 89 | 1 | CS2 16apr25 0.1m | 0 | 0 | -6.797 | -35.9798 | 0
[6 empty] | 90 | 2 | CS2 16apr25 0.1m | 0 | 0 | -6.7629 | -36.2001 | 0
[6 empty] | 91 | 3 | CS2 16apr25 0.1m | 0 | 0 | -6.7164 | -36.336 | 0
[6 empty] | 92 | 4 | CS2 16apr25 0.1m | 0 | 0 | -6.7714 | -36.022 | 0
[6 empty] | 93 | 1 | Blacksburg | 0 | 0 | -7.8636 | -47.1201 | 0
[6 empty] | 94 | 2 | Blacksburg | 0 | 0 | -7.8818 | -47.1071 | 0
[6 empty] | 95 | 3 | Blacksburg | 0 | 0 | -7.8432 | -47.3021 | 0
[6 empty] | 96 | 4 | Blacksburg | 0 | 0 | -7.8121 | -47.1086 | 0
[6 empty] | 97 | 1 | CC2 19jun24 BOT | 0 | 0 | -4.9357 | -28.2286 | 0
[6 empty] | 98 | 2 | CC2 19jun24 BOT | 0 | 0 | -4.8948 | -28.4598 | 0
[6 empty] | 99 | 3 | CC2 19jun24 BOT | 0 | 0 | -4.8463 | -28.6455 | 0
[6 empty] | 100 | 4 | CC2 19jun24 BOT | 0 | 0 | -4.8623 | -28.7763 | 0
[6 empty] | 101 | 1 | CC4 22may24 1.5m | 0 | 0 | -5.2183 | -30.6095 | 0
[6 empty] | 102 | 2 | CC4 22may24 1.5m | 0 | 0 | -5.1804 | -30.575 | 0
[6 empty] | 103 | 3 | CC4 22may24 1.5m | 0 | 0 | -5.1851 | -30.589 | 0
[6 empty] | 104 | 4 | CC4 22may24 1.5m | 0 | 0 | -5.1535 | -30.4169 | 0
[6 empty] | 105 | 1 | C50 22may24 6m | 0 | 0 | -5.3348 | -31.2958 | 0
[6 empty] | 106 | 2 | C50 22may24 6m | 0 | 0 | -5.3492 | -31.3153 | 0
[6 empty] | 107 | 3 | C50 22may24 6m | 0 | 0 | -5.344 | -31.2059 | 0
[6 empty] | 108 | 4 | C50 22may24 6m | 0 | 0 | -5.3107 | -31.3032 | 0
[6 empty] | 109 | 1 | CS1 16apr25 0.1m | 0 | 0 | -6.686 | -36.2254 | 0
[6 empty] | 110 | 2 | CS1 16apr25 0.1m | 0 | 0 | -6.6385 | -36.6512 | 0
[6 empty] | 111 | 3 | CS1 16apr25 0.1m | 0 | 0 | -6.6509 | -36.7566 | 0
[6 empty] | 112 | 4 | CS1 16apr25 0.1m | 0 | 0 | -6.7323 | -36.6438 | 0
[6 empty] | 113 | 1 | CC4 19jun24 BOT | 0 | 0 | -5.4599 | -31.7054 | 0
[6 empty] | 114 | 2 | CC4 19jun24 BOT | 0 | 0 | -5.4971 | -31.8848 | 0
[6 empty] | 115 | 3 | CC4 19jun24 BOT | 0 | 0 | -5.419 | -32.0319 | 0
[6 empty] | 116 | 4 | CC4 19jun24 BOT | 0 | 0 | -5.4262 | -32.1348 | 0
[6 empty] | 117 | 1 | C50 19jun24 20m | 0 | 0 | -5.4986 | -32.0354 | 0
[6 empty] | 118 | 2 | C50 19jun24 20m | 0 | 0 | -5.5114 | -31.9426 | 0
[6 empty] | 119 | 3 | C50 19jun24 20m | 0 | 0 | -5.5476 | -31.9338 | 0
[6 empty] | 120 | 4 | C50 19jun24 20m | 0 | 0 | -5.4199 | -32.2041 | 0
[6 empty] | 121 | 1 | CP2 19jun24 BOT - DUP | 0 | 0 | -4.765 | -28.4627 | 0
[6 empty] | 122 | 2 | CP2 19jun24 BOT - DUP | 0 | 0 | -4.7781 | -28.8977 | 0
[6 empty] | 123 | 3 | CP2 19jun24 BOT - DUP | 0 | 0 | -4.8626 | -28.5595 | 0
[6 empty] | 124 | 4 | CP2 19jun24 BOT - DUP | 0 | 0 | -4.7447 | -28.7415 | 0
[6 empty] | 125 | 1 | Hawaii | 0 | 0 | -3.7884 | -12.815 | 0
[6 empty] | 126 | 2 | Hawaii | 0 | 0 | -3.7513 | -12.8326 | 0
[6 empty] | 127 | 3 | Hawaii | 0 | 0 | -3.8804 | -12.6332 | 0
[6 empty] | 128 | 4 | Hawaii | 0 | 0 | -3.78 | -12.9161 | 0
[6 empty] | 129 | 1 | Blacksburg | 0 | 0 | -7.7062 | -47.7636 | 0
[6 empty] | 130 | 2 | Blacksburg | 0 | 0 | -7.6938 | -47.9637 | 0
[6 empty] | 131 | 3 | Blacksburg | 0 | 0 | -7.6818 | -47.8568 | 0
[6 empty] | 132 | 4 | Blacksburg | 0 | 0 | -7.701 | -47.786 | 0