## Instructions
Picarro Isotope Analyzer
Data Processor 3.0 (KAP)
How to Use this Spreadsheet:
Step 1: Retrieve raw result file from Piccaro CPU (C:Piccaro\IsotopeData). Add initials and project name to end
of file name. Copy file into your folder on Water Server.
Step 2: On syncing computer, move raw file into Raw Data folder on Google Drive:
Instruments and Analyses --> Picarro Water Isotope Analyzer --> Raw Data
Step 3: Save this spreadsheet in the Processed Data folder on Google Drive:
Instruments and Analyses --> Picarro Water Isotope Analyzer --> Processed Data
Use file name: YYYYMMDD_Initials_(anything else to identify ex. project, run number, etc.)
Step 4: Enter run information and sample positions in the "Tray Configuration" tab. (orange cells)
(It is important to use the same names for the standards throughout this spreadsheet)
Step 5: Copy text from raw result file and paste into "Picarro Output" tab.
Step 6: In Run Summary tab, check that H2O Mean is within "good" bounds. (between 17,000 - 23,000)
Step 7: In 18O tab, verify that the reference defined values match the references used in run (cells U39 - 42). Run Solver and
verify stdev memory corrected are reduced (column K values less than corresponding column I values). Verify stdev of
drift corrected reference is lower than memory corrected. Verify normalization R2 is 0.99+. Verify QC checks are within
0.15 tolerance.
Step 8: In dD tab, verify that the reference defined values match the references used in run (cells U39 - 42. Run Solver and
verify stdev memory corrected are reduced (column K values less than corresponding column I values). Verify stdev of
drift corrected reference is lower than memory corrected. Verify normalization R2 is 0.99+. Verify QC checks are within
1.5 tolerance.
Step 9: Fill in Quality Assurance Verification chart on RESULTS tab. Dr. Scott or Kelly Peeler will fill in as reviewer 2.
Step 10: Add reference results to Water Isotope Analyzer run log.

## Tray Configuration
| Analyst |  | DWH
| PI |  | Durelle Scott
| Project |  | CCR
| Sampling Date |  | 2024 - 2025
| Sampling Location |  | Carvins Cove Reservoir
| Analysis Date |  | 2025-06-05
Notes
Rear Rack
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
1 | Blacksburg | Myrtle | Homer | Blacksburg | Hawaii | C50 28oct24 1.5m | C50 30sep24 1.5m | CC3 11jul24 1.5m | Blacksburg | 9
10 | CS1 17dec24 0.1m | CC2 17dec24 0.1m | C50 17dec24 6m | CC4 17dec24 9m | CC4 17dec24 0.1m | CC4 30sep24 9m | CC2 11jul24 0.1m | CP2 28oct24 0.1m | Blacksburg | 18
19 | CP1 17dec24 0.1m | CS2 17dec24 0.1m | CC3 28oct24 0.1m | C50 17dec24 0.1m | CC4 28oct24 BOT | CC3 28oct24 BOT | CC2 11jul24 0.1m - DUP | Hawaii | Blacksburg | 27
28 | [9 empty] | 36
37 | [9 empty] | 45
46 | [9 empty] | 54
Rear Rack
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
1 | Drift | High reference | Low Reference | Drift | QC Check | Sample | Sample | Sample | Drift | 9
10 | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Drift | 18
19 | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Sample | Drift | 27
28 | [9 empty] | 36
37 | [9 empty] | 45
46 | [9 empty] | 54
Standards
|  | Standard ID | Actual d(18_16) | Actual d(D_H)
| #1 | Denali | -19.9287 | -153.7172 | [6 empty] | -3.37
| #2 | Alberta | -18.8392 | -146.8245 | [6 empty] | -20.92
| #3 | MT BHB | -18.431 | -140.5778 | [6 empty] | -17.35
| #4 | Anchorage | -16.4204 | -129.9028 | [6 empty] | -7.39
| #5 | Homer | -14.5399 | -110.5169
| #6 | MT Creek | -13.3562 | -97.7428
| #7 | Nickel Mines | -7.876 | -48.7951
| #8 | Blacksburg | -7.5615 | -47.446
| #9 | Denton's Well | -6.9583 | -42.6011
| #10 | Hawaii | -3.6108 | -13.2294
| #11 | Myrtle Beach | -2.2908 | -10.453
| #12 | Heavy Oxygen | 3.1753 | -15.7881
| #13 | Heavy Hydrogen | 11.1372 | 4.0332
| #14 | Andres Heavy Water | 17.5664 | 20.5382
| #15 | Picarro ZERO | 0.3 | 1.8
| #16 | Picarro MID | -20.6 | -159
| #17 | Picarro DEPL | -29.6 | -235
#18
#19
#20
|  | DI | -6 | -36
Note on C34: Nate: The Standard ID must match the Standard ID used above in the trey protocol. For example, if you are using the Boulder Colorado Standard, and you labeled it "B" above, it must be labled "B" in this collum.

## Timing
| Enter start time and date --> | 2019-01-10 09:30
Tray Position | Sample ID | Start Time | # Injections
1 | Blacksburg | 2019-01-10 10:58 | 10
2 | Myrtle | 2019-01-10 12:27 | 10
3 | Homer | 2019-01-10 13:55 | 10
4 | Blacksburg | 2019-01-10 15:23 | 10
5 | Hawaii | 2019-01-10 15:59 | 4
6 | C50 28oct24 1.5m | 2019-01-10 16:34 | 4
7 | C50 30sep24 1.5m | 2019-01-10 17:09 | 4
8 | CC3 11jul24 1.5m | 2019-01-10 17:45 | 4
9 | Blacksburg | 2019-01-10 18:20 | 4
10 | CS1 17dec24 0.1m | 2019-01-10 18:55 | 4
11 | CC2 17dec24 0.1m | 2019-01-10 19:31 | 4
12 | C50 17dec24 6m | 2019-01-10 20:06 | 4
13 | CC4 17dec24 9m | 2019-01-10 20:41 | 4
14 | CC4 17dec24 0.1m | 2019-01-10 21:17 | 4
15 | CC4 30sep24 9m | 2019-01-10 21:52 | 4
16 | CC2 11jul24 0.1m | 2019-01-10 22:27 | 4
17 | CP2 28oct24 0.1m | 2019-01-10 23:03 | 4
18 | Blacksburg | 2019-01-10 23:38 | 4
19 | CP1 17dec24 0.1m | 2019-01-11 00:13 | 4
20 | CS2 17dec24 0.1m | 2019-01-11 00:49 | 4
21 | CC3 28oct24 0.1m | 2019-01-11 01:24 | 4
22 | C50 17dec24 0.1m | 2019-01-11 01:59 | 4
23 | CC4 28oct24 BOT | 2019-01-11 02:35 | 4
24 | CC3 28oct24 BOT | 2019-01-11 03:10 | 4
25 | CC2 11jul24 0.1m - DUP | 2019-01-11 03:45 | 4
26 | Hawaii | 2019-01-11 04:21 | 4
27 | Blacksburg | 2019-01-11 04:56 | 4
Formulas:
Start Time = [Start Time, row above]+"0:08:50"*10

## Picarro Output
Line | Analysis | Time Code | Port | Inj Nr | d(18_16)Mean | d(D_H)Mean | H2O_Mean | Ignore | Good | Identifier 1 | Identifier 2 | Gas Configuration | days_Mean | d(18_16)_SD | d(D_H)_SD | H2O_SD | d(18_16)_Sl | d(D_H)_Sl | H2O_Sl | BASELINE_Mean | BASELINE_SD | DAS Temp | Tray | Sample | Job | Method
1 | P-11890 | 2025/06/05 09:44:57 | MT1-Rear-01 | 1 | -8.803 | -45.843 | 23699 | -1 | 0 |  |  | H2O | 155.3836 | 0.109 | 0.326 | 100 | 0.009 | -0.007 | 30.13 | 845 | 0 | 30.813 | MT1-Rear | 1 | 66 | BSE
2 | P-11890 | 2025/06/05 09:53:44 | MT1-Rear-01 | 2 | -8.992 | -44.361 | 21705 | -1 | 1 |  |  | H2O | 155.3897 | 0.14 | 0.342 | 93 | -0.007 | -0.043 | 28.1 | 845 | 0 | 30.688 | MT1-Rear | 1 | 66 | BSE
3 | P-11890 | 2025/06/05 10:02:30 | MT1-Rear-01 | 3 | -8.956 | -44.276 | 19898 | -1 | 1 |  |  | H2O | 155.3958 | 0.127 | 0.385 | 93 | -0.029 | -0.043 | 27.95 | 845 | 0 | 30.813 | MT1-Rear | 1 | 66 | BSE
4 | P-11890 | 2025/06/05 10:11:17 | MT1-Rear-01 | 4 | -9.005 | -44.314 | 19793 | 0 | 1 |  |  | H2O | 155.4019 | 0.079 | 0.403 | 93 | -0.017 | 0.07 | 27.96 | 845 | 0 | 30.813 | MT1-Rear | 1 | 66 | BSE
5 | P-11890 | 2025/06/05 10:20:03 | MT1-Rear-01 | 5 | -9 | -43.827 | 19890 | 0 | 1 |  |  | H2O | 155.408 | 0.133 | 0.446 | 100 | -0.022 | -0.016 | 27.75 | 844 | 0 | 30.938 | MT1-Rear | 1 | 66 | BSE
6 | P-11890 | 2025/06/05 10:28:49 | MT1-Rear-01 | 6 | -9.017 | -43.933 | 19904 | 0 | 1 |  |  | H2O | 155.4141 | 0.111 | 0.287 | 101 | -0.024 | -0.015 | 28.04 | 844 | 0 | 31 | MT1-Rear | 1 | 66 | BSE
7 | P-11890 | 2025/06/05 10:37:34 | MT1-Rear-01 | 7 | -8.94 | -43.394 | 19850 | 0 | 1 |  |  | H2O | 155.4202 | 0.155 | 0.306 | 92 | -0.027 | -0.032 | 27.65 | 844 | 0 | 30.688 | MT1-Rear | 1 | 66 | BSE
8 | P-11890 | 2025/06/05 10:46:21 | MT1-Rear-01 | 8 | -8.993 | -43.998 | 19658 | 0 | 1 |  |  | H2O | 155.4263 | 0.145 | 0.233 | 93 | -0.011 | -0.003 | 27.91 | 844 | 0 | 30.813 | MT1-Rear | 1 | 66 | BSE
9 | P-11890 | 2025/06/05 10:55:07 | MT1-Rear-01 | 9 | -9.018 | -43.676 | 19849 | 0 | 1 |  |  | H2O | 155.4323 | 0.099 | 0.281 | 95 | -0.016 | 0.005 | 28.75 | 844 | 0 | 31 | MT1-Rear | 1 | 66 | BSE
10 | P-11890 | 2025/06/05 11:03:54 | MT1-Rear-01 | 10 | -8.902 | -44.029 | 19863 | 0 | 1 |  |  | H2O | 155.4385 | 0.084 | 0.24 | 99 | -0.012 | -0.031 | 27.37 | 844 | 0 | 31.063 | MT1-Rear | 1 | 66 | BSE
11 | P-11891 | 2025/06/05 11:12:40 | MT1-Rear-02 | 1 | -3.761 | -9.111 | 19800 | -1 | 1 |  |  | H2O | 155.4446 | 0.138 | 0.269 | 100 | -0.023 | -0.038 | 27.69 | 844 | 0 | 31 | MT1-Rear | 2 | 66 | BSE
12 | P-11891 | 2025/06/05 11:21:27 | MT1-Rear-02 | 2 | -3.451 | -5.914 | 19827 | -1 | 1 |  |  | H2O | 155.4507 | 0.055 | 0.387 | 92 | -0.007 | -0.044 | 27.64 | 845 | 0 | 30.875 | MT1-Rear | 2 | 66 | BSE
13 | P-11891 | 2025/06/05 11:30:14 | MT1-Rear-02 | 3 | -3.545 | -4.767 | 19777 | -1 | 1 |  |  | H2O | 155.4568 | 0.117 | 0.488 | 94 | -0.02 | -0.026 | 28.37 | 844 | 0 | 30.688 | MT1-Rear | 2 | 66 | BSE
14 | P-11891 | 2025/06/05 11:39:01 | MT1-Rear-02 | 4 | -3.488 | -4.409 | 19637 | 0 | 1 |  |  | H2O | 155.4628 | 0.089 | 0.291 | 92 | -0.007 | 0.007 | 27.8 | 844 | 0 | 31.125 | MT1-Rear | 2 | 66 | BSE
15 | P-11891 | 2025/06/05 11:47:46 | MT1-Rear-02 | 5 | -3.364 | -4.163 | 19745 | 0 | 1 |  |  | H2O | 155.4689 | 0.155 | 0.229 | 92 | -0.022 | -0.035 | 27.63 | 844 | 0 | 30.625 | MT1-Rear | 2 | 66 | BSE
16 | P-11891 | 2025/06/05 11:56:33 | MT1-Rear-02 | 6 | -3.313 | -4.12 | 19632 | 0 | 1 |  |  | H2O | 155.475 | 0.095 | 0.247 | 100 | -0.003 | 0.043 | 27.65 | 844 | 0 | 31.063 | MT1-Rear | 2 | 66 | BSE
17 | P-11891 | 2025/06/05 12:05:19 | MT1-Rear-02 | 7 | -3.381 | -4.281 | 19762 | 0 | 1 |  |  | H2O | 155.4811 | 0.131 | 0.4 | 100 | -0.023 | -0.012 | 27.79 | 844 | 0 | 30.813 | MT1-Rear | 2 | 66 | BSE
18 | P-11891 | 2025/06/05 12:14:05 | MT1-Rear-02 | 8 | -3.376 | -4.096 | 19710 | 0 | 1 |  |  | H2O | 155.4872 | 0.167 | 0.332 | 93 | -0.044 | -0.042 | 27.86 | 844 | 0 | 31 | MT1-Rear | 2 | 66 | BSE
19 | P-11891 | 2025/06/05 12:22:53 | MT1-Rear-02 | 9 | -3.455 | -4.093 | 19580 | 0 | 1 |  |  | H2O | 155.4933 | 0.071 | 0.232 | 91 | -0.001 | -0.014 | 27.39 | 844 | 0 | 30.813 | MT1-Rear | 2 | 66 | BSE
20 | P-11891 | 2025/06/05 12:31:40 | MT1-Rear-02 | 10 | -3.335 | -4.029 | 19518 | 0 | 1 |  |  | H2O | 155.4994 | 0.123 | 0.382 | 90 | -0.015 | 0.09 | 27.2 | 844 | 0 | 31 | MT1-Rear | 2 | 66 | BSE
21 | P-11892 | 2025/06/05 12:40:27 | MT1-Rear-03 | 1 | -15.242 | -95.259 | 19780 | -1 | 1 |  |  | H2O | 155.5055 | 0.085 | 0.566 | 98 | -0.011 | 0.101 | 27.08 | 844 | 0 | 30.75 | MT1-Rear | 3 | 66 | BSE
22 | P-11892 | 2025/06/05 12:49:14 | MT1-Rear-03 | 2 | -15.925 | -105.092 | 19756 | -1 | 1 |  |  | H2O | 155.5116 | 0.115 | 0.631 | 99 | 0.019 | 0.155 | 27.47 | 844 | 0 | 30.875 | MT1-Rear | 3 | 66 | BSE
23 | P-11892 | 2025/06/05 12:58:00 | MT1-Rear-03 | 3 | -16.039 | -107.017 | 19807 | -1 | 1 |  |  | H2O | 155.5177 | 0.12 | 0.498 | 91 | 0 | 0.111 | 27.43 | 844 | 0 | 31.063 | MT1-Rear | 3 | 66 | BSE
24 | P-11892 | 2025/06/05 13:06:47 | MT1-Rear-03 | 4 | -16.065 | -107.939 | 19808 | 0 | 1 |  |  | H2O | 155.5238 | 0.159 | 0.391 | 93 | 0.016 | 0.091 | 28.14 | 844 | 0 | 30.813 | MT1-Rear | 3 | 66 | BSE
25 | P-11892 | 2025/06/05 13:15:35 | MT1-Rear-03 | 5 | -16.05 | -108.058 | 19887 | 0 | 1 |  |  | H2O | 155.5299 | 0.092 | 0.404 | 91 | -0.016 | 0.046 | 27.39 | 844 | 0 | 30.75 | MT1-Rear | 3 | 66 | BSE
26 | P-11892 | 2025/06/05 13:24:22 | MT1-Rear-03 | 6 | -16.122 | -108.404 | 19784 | 0 | 1 |  |  | H2O | 155.5361 | 0.113 | 0.361 | 90 | -0.004 | 0.092 | 27.08 | 844 | 0 | 30.75 | MT1-Rear | 3 | 66 | BSE
27 | P-11892 | 2025/06/05 13:33:09 | MT1-Rear-03 | 7 | -16.182 | -108.426 | 19880 | 0 | 1 |  |  | H2O | 155.5421 | 0.168 | 0.454 | 91 | -0.009 | 0.099 | 27.35 | 844 | 0 | 30.875 | MT1-Rear | 3 | 66 | BSE
28 | P-11892 | 2025/06/05 13:41:56 | MT1-Rear-03 | 8 | -16.188 | -108.751 | 19796 | 0 | 1 |  |  | H2O | 155.5482 | 0.116 | 0.55 | 93 | 0.01 | 0.126 | 27.95 | 844 | 0 | 31 | MT1-Rear | 3 | 66 | BSE
29 | P-11892 | 2025/06/05 13:50:43 | MT1-Rear-03 | 9 | -16.168 | -108.941 | 19786 | 0 | 1 |  |  | H2O | 155.5543 | 0.094 | 0.253 | 91 | -0.003 | 0.015 | 27.48 | 844 | 0 | 30.5 | MT1-Rear | 3 | 66 | BSE
30 | P-11892 | 2025/06/05 13:59:31 | MT1-Rear-03 | 10 | -16.265 | -108.882 | 19820 | 0 | 1 |  |  | H2O | 155.5604 | 0.136 | 0.51 | 102 | -0.03 | 0.09 | 28.21 | 844 | 0 | 30.875 | MT1-Rear | 3 | 66 | BSE
31 | P-11893 | 2025/06/05 14:08:19 | MT1-Rear-04 | 1 | -9.49 | -51.762 | 19841 | -1 | 1 |  |  | H2O | 155.5666 | 0.098 | 0.327 | 93 | 0.012 | -0.068 | 28.1 | 844 | 0 | 30.813 | MT1-Rear | 4 | 66 | BSE
32 | P-11893 | 2025/06/05 14:17:07 | MT1-Rear-04 | 2 | -9.082 | -45.924 | 19719 | -1 | 1 |  |  | H2O | 155.5726 | 0.148 | 0.32 | 91 | -0.033 | 0.052 | 27.38 | 844 | 0 | 30.938 | MT1-Rear | 4 | 66 | BSE
33 | P-11893 | 2025/06/05 14:25:54 | MT1-Rear-04 | 3 | -9.084 | -44.616 | 19893 | -1 | 1 |  |  | H2O | 155.5788 | 0.138 | 0.233 | 102 | -0.017 | -0.043 | 28.18 | 844 | 0 | 30.938 | MT1-Rear | 4 | 66 | BSE
34 | P-11893 | 2025/06/05 14:34:41 | MT1-Rear-04 | 4 | -8.968 | -44.29 | 19902 | 0 | 1 |  |  | H2O | 155.5848 | 0.084 | 0.28 | 99 | -0.01 | -0.011 | 27.55 | 844 | 0 | 31 | MT1-Rear | 4 | 66 | BSE
35 | P-11893 | 2025/06/05 14:43:29 | MT1-Rear-04 | 5 | -9.045 | -43.863 | 19734 | 0 | 1 |  |  | H2O | 155.591 | 0.106 | 0.333 | 93 | -0.004 | 0.02 | 28.12 | 844 | 0 | 31.25 | MT1-Rear | 4 | 66 | BSE
36 | P-11893 | 2025/06/05 14:52:17 | MT1-Rear-04 | 6 | -8.954 | -43.727 | 19833 | 0 | 1 |  |  | H2O | 155.597 | 0.121 | 0.384 | 94 | -0.022 | 0.069 | 28.22 | 844 | 0 | 30.313 | MT1-Rear | 4 | 66 | BSE
37 | P-11893 | 2025/06/05 15:01:05 | MT1-Rear-04 | 7 | -8.991 | -43.7 | 19773 | 0 | 1 |  |  | H2O | 155.6032 | 0.121 | 0.255 | 99 | -0.012 | 0.019 | 27.36 | 844 | 0 | 30.563 | MT1-Rear | 4 | 66 | BSE
38 | P-11893 | 2025/06/05 15:09:53 | MT1-Rear-04 | 8 | -8.984 | -44.084 | 19783 | 0 | 1 |  |  | H2O | 155.6093 | 0.114 | 0.293 | 93 | 0.002 | 0.025 | 27.88 | 844 | 0 | 30.75 | MT1-Rear | 4 | 66 | BSE
39 | P-11893 | 2025/06/05 15:18:41 | MT1-Rear-04 | 9 | -8.981 | -43.898 | 19748 | 0 | 1 |  |  | H2O | 155.6154 | 0.123 | 0.338 | 91 | -0.012 | -0.034 | 27.4 | 844 | 0 | 31 | MT1-Rear | 4 | 66 | BSE
40 | P-11893 | 2025/06/05 15:27:28 | MT1-Rear-04 | 10 | -8.958 | -43.769 | 19754 | 0 | 1 |  |  | H2O | 155.6215 | 0.12 | 0.352 | 92 | -0.016 | -0.003 | 27.76 | 844 | 0 | 30.938 | MT1-Rear | 4 | 66 | BSE
41 | P-11894 | 2025/06/05 15:36:17 | MT1-Rear-05 | 1 | -5.19 | -11.789 | 19861 | -1 | 1 |  |  | H2O | 155.6277 | 0.106 | 0.42 | 95 | -0.001 | 0.012 | 28.52 | 844 | 0 | 30.875 | MT1-Rear | 5 | 67 | BSE
42 | P-11894 | 2025/06/05 15:45:06 | MT1-Rear-05 | 2 | -4.904 | -8.366 | 19904 | -1 | 1 |  |  | H2O | 155.6337 | 0.072 | 0.279 | 90 | -0.006 | -0.035 | 27.22 | 844 | 0 | 31 | MT1-Rear | 5 | 67 | BSE
43 | P-11894 | 2025/06/05 15:53:54 | MT1-Rear-05 | 3 | -4.899 | -7.992 | 19819 | -1 | 1 |  |  | H2O | 155.6398 | 0.155 | 0.295 | 90 | 0.007 | 0.016 | 27.01 | 844 | 0 | 30.563 | MT1-Rear | 5 | 67 | BSE
44 | P-11894 | 2025/06/05 16:02:42 | MT1-Rear-05 | 4 | -4.904 | -7.759 | 19754 | 0 | 1 |  |  | H2O | 155.646 | 0.077 | 0.195 | 96 | 0 | 0.024 | 29 | 844 | 0 | 31.063 | MT1-Rear | 5 | 67 | BSE
45 | P-11895 | 2025/06/05 16:11:30 | MT1-Rear-06 | 1 | -5.478 | -19.234 | 19856 | -1 | 1 |  |  | H2O | 155.6521 | 0.059 | 0.243 | 90 | -0.005 | 0.027 | 27.19 | 844 | 0 | 31.5 | MT1-Rear | 6 | 67 | BSE
46 | P-11895 | 2025/06/05 16:20:18 | MT1-Rear-06 | 2 | -5.405 | -20.629 | 19869 | -1 | 1 |  |  | H2O | 155.6582 | 0.113 | 0.263 | 90 | -0.021 | 0.01 | 27.22 | 844 | 0 | 31.188 | MT1-Rear | 6 | 67 | BSE
47 | P-11895 | 2025/06/05 16:29:06 | MT1-Rear-06 | 3 | -5.489 | -20.759 | 19699 | -1 | 1 |  |  | H2O | 155.6643 | 0.104 | 0.347 | 100 | -0.02 | 0.052 | 27.82 | 844 | 0 | 31.375 | MT1-Rear | 6 | 67 | BSE
48 | P-11895 | 2025/06/05 16:37:55 | MT1-Rear-06 | 4 | -5.508 | -20.341 | 19937 | 0 | 1 |  |  | H2O | 155.6704 | 0.16 | 0.34 | 94 | -0.025 | 0.015 | 28.27 | 844 | 0 | 30.938 | MT1-Rear | 6 | 67 | BSE
49 | P-11896 | 2025/06/05 16:46:43 | MT1-Rear-07 | 1 | -5.063 | -19.41 | 19849 | -1 | 1 |  |  | H2O | 155.6765 | 0.084 | 0.465 | 92 | -0.01 | 0.07 | 27.83 | 844 | 0 | 30.625 | MT1-Rear | 7 | 67 | BSE
50 | P-11896 | 2025/06/05 16:55:32 | MT1-Rear-07 | 2 | -5.041 | -19.369 | 19959 | -1 | 1 |  |  | H2O | 155.6827 | 0.104 | 0.325 | 98 | -0.014 | 0.012 | 27.21 | 844 | 0 | 30.938 | MT1-Rear | 7 | 67 | BSE
51 | P-11896 | 2025/06/05 17:04:20 | MT1-Rear-07 | 3 | -5.109 | -19.418 | 19951 | -1 | 1 |  |  | H2O | 155.6888 | 0.143 | 0.288 | 93 | -0.029 | 0.009 | 27.92 | 844 | 0 | 31.063 | MT1-Rear | 7 | 67 | BSE
52 | P-11896 | 2025/06/05 17:13:09 | MT1-Rear-07 | 4 | -5.027 | -19.499 | 19676 | 0 | 1 |  |  | H2O | 155.6949 | 0.121 | 0.406 | 93 | -0.013 | -0.009 | 27.88 | 844 | 0 | 30.875 | MT1-Rear | 7 | 67 | BSE
53 | P-11897 | 2025/06/05 17:21:57 | MT1-Rear-08 | 1 | -5.52 | -21.289 | 19993 | -1 | 1 |  |  | H2O | 155.701 | 0.198 | 0.37 | 101 | -0.034 | -0.037 | 28.09 | 844 | 0 | 31.125 | MT1-Rear | 8 | 67 | BSE
54 | P-11897 | 2025/06/05 17:30:44 | MT1-Rear-08 | 2 | -5.562 | -21.666 | 19862 | -1 | 1 |  |  | H2O | 155.7072 | 0.112 | 0.362 | 96 | 0.008 | 0.069 | 28.78 | 844 | 0 | 30.75 | MT1-Rear | 8 | 67 | BSE
55 | P-11897 | 2025/06/05 17:39:33 | MT1-Rear-08 | 3 | -5.582 | -21.572 | 19900 | -1 | 1 |  |  | H2O | 155.7132 | 0.109 | 0.443 | 93 | -0.014 | 0.08 | 27.94 | 844 | 0 | 31.063 | MT1-Rear | 8 | 67 | BSE
56 | P-11897 | 2025/06/05 17:48:20 | MT1-Rear-08 | 4 | -5.494 | -21.874 | 19864 | 0 | 1 |  |  | H2O | 155.7194 | 0.166 | 0.433 | 100 | -0.005 | 0.098 | 27.66 | 844 | 0 | 30.875 | MT1-Rear | 8 | 67 | BSE
57 | P-11898 | 2025/06/05 17:57:08 | MT1-Rear-09 | 1 | -8.747 | -40.169 | 19846 | -1 | 1 |  |  | H2O | 155.7255 | 0.123 | 0.322 | 94 | -0.019 | 0.022 | 28.43 | 844 | 0 | 31 | MT1-Rear | 9 | 67 | BSE
58 | P-11898 | 2025/06/05 18:05:56 | MT1-Rear-09 | 2 | -8.87 | -42.844 | 19830 | -1 | 1 |  |  | H2O | 155.7316 | 0.141 | 0.405 | 94 | 0.002 | 0.078 | 28.29 | 844 | 0 | 31 | MT1-Rear | 9 | 67 | BSE
59 | P-11898 | 2025/06/05 18:14:45 | MT1-Rear-09 | 3 | -8.999 | -42.625 | 19823 | -1 | 1 |  |  | H2O | 155.7377 | 0.155 | 0.444 | 103 | -0.018 | 0.065 | 28.68 | 844 | 0 | 31.125 | MT1-Rear | 9 | 67 | BSE
60 | P-11898 | 2025/06/05 18:23:34 | MT1-Rear-09 | 4 | -8.908 | -43.041 | 19792 | 0 | 1 |  |  | H2O | 155.7438 | 0.146 | 0.375 | 91 | -0.013 | 0.049 | 27.43 | 844 | 0 | 31 | MT1-Rear | 9 | 67 | BSE
61 | P-11899 | 2025/06/05 18:32:23 | MT1-Rear-10 | 1 | -7.978 | -33.069 | 19808 | -1 | 1 |  |  | H2O | 155.7499 | 0.092 | 0.306 | 92 | -0.012 | 0.034 | 27.85 | 844 | 0 | 30.938 | MT1-Rear | 10 | 67 | BSE
62 | P-11899 | 2025/06/05 18:41:12 | MT1-Rear-10 | 2 | -7.921 | -32.042 | 19683 | -1 | 1 |  |  | H2O | 155.756 | 0.154 | 0.316 | 103 | -0.019 | 0.07 | 28.51 | 844 | 0 | 30.688 | MT1-Rear | 10 | 67 | BSE
63 | P-11899 | 2025/06/05 18:50:01 | MT1-Rear-10 | 3 | -7.92 | -31.588 | 19857 | -1 | 1 |  |  | H2O | 155.7622 | 0.095 | 0.287 | 91 | 0.009 | 0.003 | 27.53 | 844 | 0 | 30.938 | MT1-Rear | 10 | 67 | BSE
64 | P-11899 | 2025/06/05 18:58:51 | MT1-Rear-10 | 4 | -7.878 | -31.614 | 19773 | 0 | 1 |  |  | H2O | 155.7683 | 0.121 | 0.413 | 94 | -0.021 | 0.035 | 28.17 | 844 | 0 | 30.938 | MT1-Rear | 10 | 67 | BSE
65 | P-11900 | 2025/06/05 19:07:40 | MT1-Rear-11 | 1 | -7.771 | -31.371 | 19742 | -1 | 1 |  |  | H2O | 155.7745 | 0.161 | 0.324 | 93 | -0.024 | 0.047 | 27.93 | 844 | 0 | 30.938 | MT1-Rear | 11 | 67 | BSE
66 | P-11900 | 2025/06/05 19:16:28 | MT1-Rear-11 | 2 | -7.806 | -31.118 | 19856 | -1 | 1 |  |  | H2O | 155.7805 | 0.084 | 0.314 | 93 | -0.008 | 0.045 | 27.98 | 844 | 0 | 32.063 | MT1-Rear | 11 | 67 | BSE
67 | P-11900 | 2025/06/05 19:25:18 | MT1-Rear-11 | 3 | -7.712 | -30.863 | 19972 | -1 | 1 |  |  | H2O | 155.7867 | 0.118 | 0.311 | 103 | -0.019 | -0.04 | 28.67 | 844 | 0 | 30.813 | MT1-Rear | 11 | 67 | BSE
68 | P-11900 | 2025/06/05 19:34:08 | MT1-Rear-11 | 4 | -7.927 | -31.306 | 19781 | 0 | 1 |  |  | H2O | 155.7928 | 0.085 | 0.374 | 91 | -0.018 | 0.017 | 27.34 | 844 | 0 | 30.688 | MT1-Rear | 11 | 67 | BSE
69 | P-11901 | 2025/06/05 19:42:57 | MT1-Rear-12 | 1 | -5.665 | -22.684 | 19782 | -1 | 1 |  |  | H2O | 155.799 | 0.096 | 0.397 | 100 | -0.012 | 0.037 | 27.77 | 844 | 0 | 30.938 | MT1-Rear | 12 | 67 | BSE
70 | P-11901 | 2025/06/05 19:51:46 | MT1-Rear-12 | 2 | -5.505 | -21.758 | 19820 | -1 | 1 |  |  | H2O | 155.8051 | 0.12 | 0.334 | 94 | -0.021 | -0.003 | 28.31 | 844 | 0 | 30.875 | MT1-Rear | 12 | 67 | BSE
71 | P-11901 | 2025/06/05 20:00:35 | MT1-Rear-12 | 3 | -5.506 | -21.693 | 19887 | -1 | 1 |  |  | H2O | 155.8112 | 0.134 | 0.314 | 92 | 0.002 | -0.02 | 27.79 | 844 | 0 | 31.25 | MT1-Rear | 12 | 67 | BSE
72 | P-11901 | 2025/06/05 20:09:25 | MT1-Rear-12 | 4 | -5.619 | -21.425 | 19736 | 0 | 1 |  |  | H2O | 155.8173 | 0.105 | 0.415 | 101 | -0.013 | -0.029 | 28.04 | 844 | 0 | 30.813 | MT1-Rear | 12 | 67 | BSE
73 | P-11902 | 2025/06/05 20:18:14 | MT1-Rear-13 | 1 | -5.47 | -21.198 | 19784 | -1 | 1 |  |  | H2O | 155.8235 | 0.133 | 0.271 | 93 | -0.013 | 0.024 | 27.94 | 844 | 0 | 30.813 | MT1-Rear | 13 | 67 | BSE
74 | P-11902 | 2025/06/05 20:27:03 | MT1-Rear-13 | 2 | -5.417 | -21.048 | 19889 | -1 | 1 |  |  | H2O | 155.8295 | 0.123 | 0.31 | 93 | -0.012 | 0.056 | 27.88 | 844 | 0 | 31.375 | MT1-Rear | 13 | 67 | BSE
75 | P-11902 | 2025/06/05 20:35:53 | MT1-Rear-13 | 3 | -5.525 | -21.292 | 19726 | -1 | 1 |  |  | H2O | 155.8357 | 0.158 | 0.321 | 94 | -0.025 | -0.014 | 28.42 | 844 | 0 | 31.125 | MT1-Rear | 13 | 67 | BSE
76 | P-11902 | 2025/06/05 20:44:42 | MT1-Rear-13 | 4 | -5.468 | -21.209 | 19813 | 0 | 1 |  |  | H2O | 155.8418 | 0.169 | 0.288 | 93 | -0.013 | 0.04 | 28.07 | 844 | 0 | 30.875 | MT1-Rear | 13 | 67 | BSE
77 | P-11903 | 2025/06/05 20:53:32 | MT1-Rear-14 | 1 | -5.401 | -21.12 | 19942 | -1 | 1 |  |  | H2O | 155.848 | 0.107 | 0.281 | 97 | 0.004 | 0.042 | 29.12 | 844 | 0 | 30.5 | MT1-Rear | 14 | 67 | BSE
78 | P-11903 | 2025/06/05 21:02:21 | MT1-Rear-14 | 2 | -5.47 | -21.254 | 19851 | -1 | 1 |  |  | H2O | 155.8541 | 0.131 | 0.257 | 94 | -0.022 | 0.037 | 28.19 | 844 | 0 | 30.25 | MT1-Rear | 14 | 67 | BSE
79 | P-11903 | 2025/06/05 21:11:11 | MT1-Rear-14 | 3 | -5.323 | -21.368 | 19764 | -1 | 1 |  |  | H2O | 155.8602 | 0.09 | 0.429 | 101 | -0.009 | 0.078 | 28.12 | 844 | 0 | 30.75 | MT1-Rear | 14 | 67 | BSE
80 | P-11903 | 2025/06/05 21:20:00 | MT1-Rear-14 | 4 | -5.42 | -21.062 | 19810 | 0 | 1 |  |  | H2O | 155.8664 | 0.109 | 0.413 | 90 | -0.005 | -0.088 | 27.12 | 844 | 0 | 30.813 | MT1-Rear | 14 | 67 | BSE
81 | P-11904 | 2025/06/05 21:28:50 | MT1-Rear-15 | 1 | -5.989 | -24.077 | 19790 | -1 | 1 |  |  | H2O | 155.8725 | 0.103 | 0.405 | 99 | -0.015 | -0.087 | 27.39 | 844 | 0 | 31.125 | MT1-Rear | 15 | 67 | BSE
82 | P-11904 | 2025/06/05 21:37:39 | MT1-Rear-15 | 2 | -6.06 | -24.496 | 19785 | -1 | 1 |  |  | H2O | 155.8786 | 0.147 | 0.287 | 93 | -0.026 | 0.014 | 27.92 | 844 | 0 | 30.563 | MT1-Rear | 15 | 67 | BSE
83 | P-11904 | 2025/06/05 21:46:28 | MT1-Rear-15 | 3 | -6.057 | -24.438 | 19828 | -1 | 1 |  |  | H2O | 155.8847 | 0.123 | 0.332 | 93 | -0.012 | 0.088 | 28.1 | 844 | 0 | 30.75 | MT1-Rear | 15 | 67 | BSE
84 | P-11904 | 2025/06/05 21:55:18 | MT1-Rear-15 | 4 | -6.078 | -24.12 | 19860 | 0 | 1 |  |  | H2O | 155.8909 | 0.103 | 0.349 | 100 | -0.005 | 0.002 | 27.59 | 844 | 0 | 30.563 | MT1-Rear | 15 | 67 | BSE
85 | P-11905 | 2025/06/05 22:04:08 | MT1-Rear-16 | 1 | -5.578 | -22.046 | 19831 | -1 | 1 |  |  | H2O | 155.897 | 0.061 | 0.383 | 93 | 0.006 | 0.032 | 27.9 | 844 | 0 | 30.625 | MT1-Rear | 16 | 67 | BSE
86 | P-11905 | 2025/06/05 22:12:57 | MT1-Rear-16 | 2 | -5.595 | -21.703 | 19910 | -1 | 1 |  |  | H2O | 155.9031 | 0.139 | 0.299 | 92 | -0.019 | 0.065 | 27.68 | 844 | 0 | 31.188 | MT1-Rear | 16 | 67 | BSE
87 | P-11905 | 2025/06/05 22:21:46 | MT1-Rear-16 | 3 | -5.481 | -21.67 | 19892 | -1 | 1 |  |  | H2O | 155.9093 | 0.117 | 0.412 | 95 | -0.013 | 0.092 | 28.7 | 844 | 0 | 30.75 | MT1-Rear | 16 | 67 | BSE
88 | P-11905 | 2025/06/05 22:30:36 | MT1-Rear-16 | 4 | -5.564 | -21.501 | 19848 | 0 | 1 |  |  | H2O | 155.9153 | 0.133 | 0.327 | 94 | -0.022 | -0.031 | 28.29 | 844 | 0 | 31 | MT1-Rear | 16 | 67 | BSE
89 | P-11906 | 2025/06/05 22:39:25 | MT1-Rear-17 | 1 | -5.357 | -20.724 | 19735 | -1 | 1 |  |  | H2O | 155.9215 | 0.142 | 0.36 | 101 | 0.011 | 0.045 | 27.86 | 844 | 0 | 31.125 | MT1-Rear | 17 | 67 | BSE
90 | P-11906 | 2025/06/05 22:48:16 | MT1-Rear-17 | 2 | -5.309 | -20.255 | 19843 | -1 | 1 |  |  | H2O | 155.9276 | 0.104 | 0.297 | 96 | -0.022 | 0.057 | 29.01 | 844 | 0 | 31.188 | MT1-Rear | 17 | 67 | BSE
91 | P-11906 | 2025/06/05 22:57:06 | MT1-Rear-17 | 3 | -5.275 | -20.628 | 19792 | -1 | 1 |  |  | H2O | 155.9338 | 0.136 | 0.258 | 97 | -0.009 | -0.052 | 26.98 | 844 | 0 | 30.375 | MT1-Rear | 17 | 67 | BSE
92 | P-11906 | 2025/06/05 23:05:57 | MT1-Rear-17 | 4 | -5.297 | -20.181 | 19932 | 0 | 1 |  |  | H2O | 155.9399 | 0.115 | 0.313 | 90 | -0.031 | -0.016 | 27.24 | 844 | 0 | 30.75 | MT1-Rear | 17 | 67 | BSE
93 | P-11907 | 2025/06/05 23:14:47 | MT1-Rear-18 | 1 | -8.649 | -40.747 | 19743 | -1 | 1 |  |  | H2O | 155.9461 | 0.088 | 0.383 | 92 | 0.005 | 0.055 | 27.74 | 844 | 0 | 30.938 | MT1-Rear | 18 | 67 | BSE
94 | P-11907 | 2025/06/05 23:23:37 | MT1-Rear-18 | 2 | -8.822 | -42.457 | 20089 | -1 | 1 |  |  | H2O | 155.9522 | 0.091 | 0.362 | 94 | -0.014 | 0.046 | 28.23 | 844 | 0 | 30.813 | MT1-Rear | 18 | 67 | BSE
95 | P-11907 | 2025/06/05 23:32:28 | MT1-Rear-18 | 3 | -8.874 | -43.124 | 19801 | -1 | 1 |  |  | H2O | 155.9583 | 0.115 | 0.358 | 101 | -0.011 | 0.076 | 27.93 | 844 | 0 | 30.813 | MT1-Rear | 18 | 67 | BSE
96 | P-11907 | 2025/06/05 23:41:18 | MT1-Rear-18 | 4 | -8.856 | -42.996 | 19807 | 0 | 1 |  |  | H2O | 155.9645 | 0.117 | 0.2 | 92 | -0.01 | 0.031 | 27.66 | 844 | 0 | 30.375 | MT1-Rear | 18 | 67 | BSE
97 | P-11908 | 2025/06/05 23:50:08 | MT1-Rear-19 | 1 | -7.892 | -32.569 | 19809 | -1 | 1 |  |  | H2O | 155.9706 | 0.112 | 0.286 | 98 | 0.008 | -0.054 | 27.04 | 844 | 0 | 30.188 | MT1-Rear | 19 | 67 | BSE
98 | P-11908 | 2025/06/05 23:58:59 | MT1-Rear-19 | 2 | -7.834 | -31.595 | 20077 | -1 | 1 |  |  | H2O | 155.9768 | 0.192 | 0.392 | 90 | -0.039 | 0.024 | 27.05 | 844 | 0 | 30.5 | MT1-Rear | 19 | 67 | BSE
99 | P-11908 | 2025/06/06 00:07:50 | MT1-Rear-19 | 3 | -7.815 | -31.851 | 19956 | -1 | 1 |  |  | H2O | 155.9829 | 0.116 | 0.346 | 98 | -0.004 | -0.034 | 27.05 | 844 | 0 | 30.938 | MT1-Rear | 19 | 67 | BSE
100 | P-11908 | 2025/06/06 00:16:40 | MT1-Rear-19 | 4 | -7.749 | -31.223 | 20017 | 0 | 1 |  |  | H2O | 155.989 | 0.122 | 0.257 | 92 | -0.022 | -0.048 | 27.84 | 844 | 0 | 30.625 | MT1-Rear | 19 | 67 | BSE
101 | P-11909 | 2025/06/06 00:25:32 | MT1-Rear-20 | 1 | -7.819 | -31.413 | 19961 | -1 | 1 |  |  | H2O | 155.9952 | 0.169 | 0.275 | 90 | -0.025 | 0.045 | 27.18 | 844 | 0 | 30.875 | MT1-Rear | 20 | 67 | BSE
102 | P-11909 | 2025/06/06 00:34:23 | MT1-Rear-20 | 2 | -7.814 | -31.423 | 19852 | -1 | 1 |  |  | H2O | 156.0013 | 0.11 | 0.374 | 89 | -0.02 | 0.02 | 26.82 | 844 | 0 | 30.688 | MT1-Rear | 20 | 67 | BSE
103 | P-11909 | 2025/06/06 00:43:14 | MT1-Rear-20 | 3 | -7.818 | -31.532 | 19849 | -1 | 1 |  |  | H2O | 156.0075 | 0.117 | 0.299 | 91 | -0.006 | -0.04 | 27.26 | 844 | 0 | 30.813 | MT1-Rear | 20 | 67 | BSE
104 | P-11909 | 2025/06/06 00:52:05 | MT1-Rear-20 | 4 | -7.872 | -31.328 | 19997 | 0 | 1 |  |  | H2O | 156.0137 | 0.157 | 0.198 | 101 | -0.005 | 0.007 | 28.04 | 844 | 0 | 31 | MT1-Rear | 20 | 67 | BSE
105 | P-11910 | 2025/06/06 01:00:57 | MT1-Rear-21 | 1 | -5.484 | -21.545 | 19810 | -1 | 1 |  |  | H2O | 156.0198 | 0.148 | 0.356 | 91 | -0.025 | 0 | 27.41 | 844 | 0 | 30.813 | MT1-Rear | 21 | 67 | BSE
106 | P-11910 | 2025/06/06 01:09:48 | MT1-Rear-21 | 2 | -5.287 | -20.95 | 20122 | -1 | 1 |  |  | H2O | 156.0259 | 0.149 | 0.343 | 97 | -0.021 | -0.019 | 27.01 | 844 | 0 | 30.938 | MT1-Rear | 21 | 67 | BSE
107 | P-11910 | 2025/06/06 01:18:39 | MT1-Rear-21 | 3 | -5.408 | -20.663 | 19952 | -1 | 1 |  |  | H2O | 156.0321 | 0.169 | 0.44 | 91 | -0.035 | -0.021 | 27.33 | 844 | 0 | 30.938 | MT1-Rear | 21 | 67 | BSE
108 | P-11910 | 2025/06/06 01:27:30 | MT1-Rear-21 | 4 | -5.354 | -20.555 | 19950 | 0 | 1 |  |  | H2O | 156.0382 | 0.168 | 0.215 | 98 | -0.038 | -0.008 | 27.27 | 844 | 0 | 30.75 | MT1-Rear | 21 | 67 | BSE
109 | P-11911 | 2025/06/06 01:36:22 | MT1-Rear-22 | 1 | -5.453 | -21.281 | 20031 | -1 | 1 |  |  | H2O | 156.0444 | 0.116 | 0.257 | 90 | -0.024 | -0.044 | 27.24 | 844 | 0 | 31.063 | MT1-Rear | 22 | 67 | BSE
110 | P-11911 | 2025/06/06 01:45:13 | MT1-Rear-22 | 2 | -5.546 | -21.245 | 19905 | -1 | 1 |  |  | H2O | 156.0506 | 0.102 | 0.457 | 98 | -0.006 | 0.099 | 29.57 | 844 | 0 | 30.875 | MT1-Rear | 22 | 67 | BSE
111 | P-11911 | 2025/06/06 01:54:04 | MT1-Rear-22 | 3 | -5.412 | -21.155 | 19924 | -1 | 1 |  |  | H2O | 156.0567 | 0.145 | 0.322 | 100 | -0.024 | -0.047 | 27.65 | 844 | 0 | 30.688 | MT1-Rear | 22 | 67 | BSE
112 | P-11911 | 2025/06/06 02:02:57 | MT1-Rear-22 | 4 | -5.441 | -21.401 | 19962 | 0 | 1 |  |  | H2O | 156.0628 | 0.15 | 0.217 | 91 | -0.029 | 0.002 | 27.45 | 844 | 0 | 31.188 | MT1-Rear | 22 | 67 | BSE
113 | P-11912 | 2025/06/06 02:11:48 | MT1-Rear-23 | 1 | -5.778 | -21.928 | 19863 | -1 | 1 |  |  | H2O | 156.069 | 0.126 | 0.289 | 103 | -0.025 | 0.06 | 28.43 | 844 | 0 | 30.813 | MT1-Rear | 23 | 67 | BSE
114 | P-11912 | 2025/06/06 02:20:39 | MT1-Rear-23 | 2 | -5.705 | -22.179 | 19795 | -1 | 1 |  |  | H2O | 156.0751 | 0.095 | 0.247 | 90 | -0.013 | -0.015 | 27.15 | 844 | 0 | 30.563 | MT1-Rear | 23 | 67 | BSE
115 | P-11912 | 2025/06/06 02:29:30 | MT1-Rear-23 | 3 | -5.729 | -21.948 | 19883 | -1 | 1 |  |  | H2O | 156.0813 | 0.174 | 0.384 | 93 | -0.034 | 0.036 | 27.98 | 844 | 0 | 30.313 | MT1-Rear | 23 | 67 | BSE
116 | P-11912 | 2025/06/06 02:38:21 | MT1-Rear-23 | 4 | -5.655 | -22.396 | 19829 | 0 | 1 |  |  | H2O | 156.0874 | 0.106 | 0.232 | 90 | -0.007 | 0.025 | 26.98 | 844 | 0 | 30.625 | MT1-Rear | 23 | 67 | BSE
117 | P-11913 | 2025/06/06 02:47:13 | MT1-Rear-24 | 1 | -5.438 | -20.438 | 19957 | -1 | 1 |  |  | H2O | 156.0936 | 0.166 | 0.328 | 91 | -0.041 | -0.06 | 27.38 | 844 | 0 | 30.188 | MT1-Rear | 24 | 67 | BSE
118 | P-11913 | 2025/06/06 02:56:03 | MT1-Rear-24 | 2 | -5.318 | -20.291 | 19947 | -1 | 1 |  |  | H2O | 156.0998 | 0.116 | 0.282 | 104 | -0.019 | -0.009 | 28.73 | 844 | 0 | 30.5 | MT1-Rear | 24 | 67 | BSE
119 | P-11913 | 2025/06/06 03:04:55 | MT1-Rear-24 | 3 | -5.302 | -20.468 | 19904 | -1 | 1 |  |  | H2O | 156.1059 | 0.066 | 0.191 | 95 | 0.007 | 0.007 | 28.49 | 844 | 0 | 30.438 | MT1-Rear | 24 | 67 | BSE
120 | P-11913 | 2025/06/06 03:13:46 | MT1-Rear-24 | 4 | -5.293 | -20.459 | 19946 | 0 | 1 |  |  | H2O | 156.112 | 0.109 | 0.239 | 100 | -0.003 | 0.019 | 27.62 | 844 | 0 | 30.5 | MT1-Rear | 24 | 67 | BSE
121 | P-11914 | 2025/06/06 03:22:38 | MT1-Rear-25 | 1 | -5.541 | -21.671 | 19845 | -1 | 1 |  |  | H2O | 156.1182 | 0.116 | 0.296 | 90 | -0.012 | -0.055 | 27.25 | 844 | 0 | 30.313 | MT1-Rear | 25 | 67 | BSE
122 | P-11914 | 2025/06/06 03:31:29 | MT1-Rear-25 | 2 | -5.573 | -21.51 | 19920 | -1 | 1 |  |  | H2O | 156.1244 | 0.149 | 0.509 | 92 | -0.023 | -0.018 | 27.67 | 844 | 0 | 30.875 | MT1-Rear | 25 | 67 | BSE
123 | P-11914 | 2025/06/06 03:40:21 | MT1-Rear-25 | 3 | -5.634 | -21.338 | 19726 | -1 | 1 |  |  | H2O | 156.1305 | 0.16 | 0.304 | 91 | -0.015 | 0.013 | 27.36 | 844 | 0 | 31 | MT1-Rear | 25 | 67 | BSE
124 | P-11914 | 2025/06/06 03:49:13 | MT1-Rear-25 | 4 | -5.556 | -21.578 | 19906 | 0 | 1 |  |  | H2O | 156.1367 | 0.099 | 0.184 | 93 | -0.006 | -0.022 | 27.96 | 844 | 0 | 31 | MT1-Rear | 25 | 67 | BSE
125 | P-11915 | 2025/06/06 03:58:05 | MT1-Rear-26 | 1 | -4.891 | -8.458 | 19948 | -1 | 1 |  |  | H2O | 156.1428 | 0.144 | 0.356 | 98 | -0.02 | -0.083 | 27.1 | 844 | 0 | 31 | MT1-Rear | 26 | 67 | BSE
126 | P-11915 | 2025/06/06 04:06:56 | MT1-Rear-26 | 2 | -4.844 | -6.978 | 19945 | -1 | 1 |  |  | H2O | 156.149 | 0.156 | 0.278 | 91 | -0.036 | 0.037 | 27.55 | 844 | 0 | 30.625 | MT1-Rear | 26 | 67 | BSE
127 | P-11915 | 2025/06/06 04:15:48 | MT1-Rear-26 | 3 | -4.778 | -6.872 | 19779 | -1 | 1 |  |  | H2O | 156.1551 | 0.147 | 0.268 | 98 | -0.02 | -0.017 | 27.16 | 844 | 0 | 30.438 | MT1-Rear | 26 | 67 | BSE
128 | P-11915 | 2025/06/06 04:24:40 | MT1-Rear-26 | 4 | -4.938 | -6.582 | 19833 | 0 | 1 |  |  | H2O | 156.1612 | 0.082 | 0.304 | 92 | -0.017 | 0.008 | 27.78 | 844 | 0 | 29.813 | MT1-Rear | 26 | 67 | BSE
129 | P-11916 | 2025/06/06 04:33:31 | MT1-Rear-27 | 1 | -8.644 | -38.379 | 19849 | -1 | 1 |  |  | H2O | 156.1675 | 0.125 | 0.433 | 93 | -0.017 | 0.058 | 27.99 | 844 | 0 | 30.688 | MT1-Rear | 27 | 67 | BSE
130 | P-11916 | 2025/06/06 04:42:23 | MT1-Rear-27 | 2 | -8.896 | -41.522 | 19855 | -1 | 1 |  |  | H2O | 156.1736 | 0.111 | 0.365 | 97 | -0.003 | 0.04 | 26.8 | 844 | 0 | 30.875 | MT1-Rear | 27 | 67 | BSE
131 | P-11916 | 2025/06/06 04:51:14 | MT1-Rear-27 | 3 | -8.888 | -42.456 | 19818 | -1 | 1 |  |  | H2O | 156.1797 | 0.114 | 0.414 | 93 | -0.012 | 0.081 | 28.01 | 844 | 0 | 31.125 | MT1-Rear | 27 | 67 | BSE
132 | P-11916 | 2025/06/06 05:00:07 | MT1-Rear-27 | 4 | -8.961 | -42.463 | 19937 | 0 | 1 |  |  | H2O | 156.1859 | 0.095 | 0.163 | 92 | 0.004 | 0.017 | 27.75 | 844 | 0 | 30.938 | MT1-Rear | 27 | 67 | BSE

## Run Summary
Line | Sample | Inj Nr | Sample Name | H2O_Mean | Good | Ignore | Time Code | d(18_16)Mean | d(D_H)Mean
1 | 1 | 1 | Blacksburg | 23699 | 0 | -1 | 2025/06/05 09:44:57 | -8.803 | -45.843 |  | Min | Max
2 | 1 | 2 | Blacksburg | 21705 | 1 | 0 | 2025/06/05 09:53:44 | -8.992 | -44.361 | 0 | 17000 | 23000
3 | 1 | 3 | Blacksburg | 19898 | 1 | 0 | 2025/06/05 10:02:30 | -8.956 | -44.276 | 500 | 17000 | 23000
4 | 1 | 4 | Blacksburg | 19793 | 1 | 0 | 2025/06/05 10:11:17 | -9.005 | -44.314
5 | 1 | 5 | Blacksburg | 19890 | 1 | 0 | 2025/06/05 10:20:03 | -9 | -43.827
6 | 1 | 6 | Blacksburg | 19904 | 1 | 0 | 2025/06/05 10:28:49 | -9.017 | -43.933
7 | 1 | 7 | Blacksburg | 19850 | 1 | 0 | 2025/06/05 10:37:34 | -8.94 | -43.394
8 | 1 | 8 | Blacksburg | 19658 | 1 | 0 | 2025/06/05 10:46:21 | -8.993 | -43.998
9 | 1 | 9 | Blacksburg | 19849 | 1 | 0 | 2025/06/05 10:55:07 | -9.018 | -43.676
10 | 1 | 10 | Blacksburg | 19863 | 1 | 0 | 2025/06/05 11:03:54 | -8.902 | -44.029
11 | 2 | 1 | Myrtle | 19800 | 1 | 0 | 2025/06/05 11:12:40 | -3.761 | -9.111
12 | 2 | 2 | Myrtle | 19827 | 1 | 0 | 2025/06/05 11:21:27 | -3.451 | -5.914
13 | 2 | 3 | Myrtle | 19777 | 1 | 0 | 2025/06/05 11:30:14 | -3.545 | -4.767
14 | 2 | 4 | Myrtle | 19637 | 1 | 0 | 2025/06/05 11:39:01 | -3.488 | -4.409
15 | 2 | 5 | Myrtle | 19745 | 1 | 0 | 2025/06/05 11:47:46 | -3.364 | -4.163
16 | 2 | 6 | Myrtle | 19632 | 1 | 0 | 2025/06/05 11:56:33 | -3.313 | -4.12
17 | 2 | 7 | Myrtle | 19762 | 1 | 0 | 2025/06/05 12:05:19 | -3.381 | -4.281 | Concentration of water vapor should be about 20,000 ppmV (roughly +/- 1000; but not outside 17,000 and 23,000 ppmV).
18 | 2 | 8 | Myrtle | 19710 | 1 | 0 | 2025/06/05 12:14:05 | -3.376 | -4.096 | Large scatter may indicate a problem with the syringe or the septa.
19 | 2 | 9 | Myrtle | 19580 | 1 | 0 | 2025/06/05 12:22:53 | -3.455 | -4.093
20 | 2 | 10 | Myrtle | 19518 | 1 | 0 | 2025/06/05 12:31:40 | -3.335 | -4.029
21 | 3 | 1 | Homer | 19780 | 1 | 0 | 2025/06/05 12:40:27 | -15.242 | -95.259
22 | 3 | 2 | Homer | 19756 | 1 | 0 | 2025/06/05 12:49:14 | -15.925 | -105.092
23 | 3 | 3 | Homer | 19807 | 1 | 0 | 2025/06/05 12:58:00 | -16.039 | -107.017
24 | 3 | 4 | Homer | 19808 | 1 | 0 | 2025/06/05 13:06:47 | -16.065 | -107.939
25 | 3 | 5 | Homer | 19887 | 1 | 0 | 2025/06/05 13:15:35 | -16.05 | -108.058
26 | 3 | 6 | Homer | 19784 | 1 | 0 | 2025/06/05 13:24:22 | -16.122 | -108.404
27 | 3 | 7 | Homer | 19880 | 1 | 0 | 2025/06/05 13:33:09 | -16.182 | -108.426
28 | 3 | 8 | Homer | 19796 | 1 | 0 | 2025/06/05 13:41:56 | -16.188 | -108.751
29 | 3 | 9 | Homer | 19786 | 1 | 0 | 2025/06/05 13:50:43 | -16.168 | -108.941
30 | 3 | 10 | Homer | 19820 | 1 | 0 | 2025/06/05 13:59:31 | -16.265 | -108.882
31 | 4 | 1 | Blacksburg | 19841 | 1 | 0 | 2025/06/05 14:08:19 | -9.49 | -51.762
32 | 4 | 2 | Blacksburg | 19719 | 1 | 0 | 2025/06/05 14:17:07 | -9.082 | -45.924
33 | 4 | 3 | Blacksburg | 19893 | 1 | 0 | 2025/06/05 14:25:54 | -9.084 | -44.616
34 | 4 | 4 | Blacksburg | 19902 | 1 | 0 | 2025/06/05 14:34:41 | -8.968 | -44.29
35 | 4 | 5 | Blacksburg | 19734 | 1 | 0 | 2025/06/05 14:43:29 | -9.045 | -43.863
36 | 4 | 6 | Blacksburg | 19833 | 1 | 0 | 2025/06/05 14:52:17 | -8.954 | -43.727
37 | 4 | 7 | Blacksburg | 19773 | 1 | 0 | 2025/06/05 15:01:05 | -8.991 | -43.7
38 | 4 | 8 | Blacksburg | 19783 | 1 | 0 | 2025/06/05 15:09:53 | -8.984 | -44.084
39 | 4 | 9 | Blacksburg | 19748 | 1 | 0 | 2025/06/05 15:18:41 | -8.981 | -43.898
40 | 4 | 10 | Blacksburg | 19754 | 1 | 0 | 2025/06/05 15:27:28 | -8.958 | -43.769
41 | 5 | 1 | Hawaii | 19861 | 1 | 0 | 2025/06/05 15:36:17 | -5.19 | -11.789
42 | 5 | 2 | Hawaii | 19904 | 1 | 0 | 2025/06/05 15:45:06 | -4.904 | -8.366
43 | 5 | 3 | Hawaii | 19819 | 1 | 0 | 2025/06/05 15:53:54 | -4.899 | -7.992
44 | 5 | 4 | Hawaii | 19754 | 1 | 0 | 2025/06/05 16:02:42 | -4.904 | -7.759
45 | 6 | 1 | C50 28oct24 1.5m | 19856 | 1 | 0 | 2025/06/05 16:11:30 | -5.478 | -19.234
46 | 6 | 2 | C50 28oct24 1.5m | 19869 | 1 | 0 | 2025/06/05 16:20:18 | -5.405 | -20.629
47 | 6 | 3 | C50 28oct24 1.5m | 19699 | 1 | 0 | 2025/06/05 16:29:06 | -5.489 | -20.759
48 | 6 | 4 | C50 28oct24 1.5m | 19937 | 1 | 0 | 2025/06/05 16:37:55 | -5.508 | -20.341
49 | 7 | 1 | C50 30sep24 1.5m | 19849 | 1 | 0 | 2025/06/05 16:46:43 | -5.063 | -19.41
50 | 7 | 2 | C50 30sep24 1.5m | 19959 | 1 | 0 | 2025/06/05 16:55:32 | -5.041 | -19.369
51 | 7 | 3 | C50 30sep24 1.5m | 19951 | 1 | 0 | 2025/06/05 17:04:20 | -5.109 | -19.418
52 | 7 | 4 | C50 30sep24 1.5m | 19676 | 1 | 0 | 2025/06/05 17:13:09 | -5.027 | -19.499
53 | 8 | 1 | CC3 11jul24 1.5m | 19993 | 1 | 0 | 2025/06/05 17:21:57 | -5.52 | -21.289
54 | 8 | 2 | CC3 11jul24 1.5m | 19862 | 1 | 0 | 2025/06/05 17:30:44 | -5.562 | -21.666
55 | 8 | 3 | CC3 11jul24 1.5m | 19900 | 1 | 0 | 2025/06/05 17:39:33 | -5.582 | -21.572
56 | 8 | 4 | CC3 11jul24 1.5m | 19864 | 1 | 0 | 2025/06/05 17:48:20 | -5.494 | -21.874
57 | 9 | 1 | Blacksburg | 19846 | 1 | 0 | 2025/06/05 17:57:08 | -8.747 | -40.169
58 | 9 | 2 | Blacksburg | 19830 | 1 | 0 | 2025/06/05 18:05:56 | -8.87 | -42.844
59 | 9 | 3 | Blacksburg | 19823 | 1 | 0 | 2025/06/05 18:14:45 | -8.999 | -42.625
60 | 9 | 4 | Blacksburg | 19792 | 1 | 0 | 2025/06/05 18:23:34 | -8.908 | -43.041
61 | 10 | 1 | CS1 17dec24 0.1m | 19808 | 1 | 0 | 2025/06/05 18:32:23 | -7.978 | -33.069
62 | 10 | 2 | CS1 17dec24 0.1m | 19683 | 1 | 0 | 2025/06/05 18:41:12 | -7.921 | -32.042
63 | 10 | 3 | CS1 17dec24 0.1m | 19857 | 1 | 0 | 2025/06/05 18:50:01 | -7.92 | -31.588
64 | 10 | 4 | CS1 17dec24 0.1m | 19773 | 1 | 0 | 2025/06/05 18:58:51 | -7.878 | -31.614
65 | 11 | 1 | CC2 17dec24 0.1m | 19742 | 1 | 0 | 2025/06/05 19:07:40 | -7.771 | -31.371
66 | 11 | 2 | CC2 17dec24 0.1m | 19856 | 1 | 0 | 2025/06/05 19:16:28 | -7.806 | -31.118
67 | 11 | 3 | CC2 17dec24 0.1m | 19972 | 1 | 0 | 2025/06/05 19:25:18 | -7.712 | -30.863
68 | 11 | 4 | CC2 17dec24 0.1m | 19781 | 1 | 0 | 2025/06/05 19:34:08 | -7.927 | -31.306
69 | 12 | 1 | C50 17dec24 6m | 19782 | 1 | 0 | 2025/06/05 19:42:57 | -5.665 | -22.684
70 | 12 | 2 | C50 17dec24 6m | 19820 | 1 | 0 | 2025/06/05 19:51:46 | -5.505 | -21.758
71 | 12 | 3 | C50 17dec24 6m | 19887 | 1 | 0 | 2025/06/05 20:00:35 | -5.506 | -21.693
72 | 12 | 4 | C50 17dec24 6m | 19736 | 1 | 0 | 2025/06/05 20:09:25 | -5.619 | -21.425
73 | 13 | 1 | CC4 17dec24 9m | 19784 | 1 | 0 | 2025/06/05 20:18:14 | -5.47 | -21.198
74 | 13 | 2 | CC4 17dec24 9m | 19889 | 1 | 0 | 2025/06/05 20:27:03 | -5.417 | -21.048
75 | 13 | 3 | CC4 17dec24 9m | 19726 | 1 | 0 | 2025/06/05 20:35:53 | -5.525 | -21.292
76 | 13 | 4 | CC4 17dec24 9m | 19813 | 1 | 0 | 2025/06/05 20:44:42 | -5.468 | -21.209
77 | 14 | 1 | CC4 17dec24 0.1m | 19942 | 1 | 0 | 2025/06/05 20:53:32 | -5.401 | -21.12
78 | 14 | 2 | CC4 17dec24 0.1m | 19851 | 1 | 0 | 2025/06/05 21:02:21 | -5.47 | -21.254
79 | 14 | 3 | CC4 17dec24 0.1m | 19764 | 1 | 0 | 2025/06/05 21:11:11 | -5.323 | -21.368
80 | 14 | 4 | CC4 17dec24 0.1m | 19810 | 1 | 0 | 2025/06/05 21:20:00 | -5.42 | -21.062
81 | 15 | 1 | CC4 30sep24 9m | 19790 | 1 | 0 | 2025/06/05 21:28:50 | -5.989 | -24.077
82 | 15 | 2 | CC4 30sep24 9m | 19785 | 1 | 0 | 2025/06/05 21:37:39 | -6.06 | -24.496
83 | 15 | 3 | CC4 30sep24 9m | 19828 | 1 | 0 | 2025/06/05 21:46:28 | -6.057 | -24.438
84 | 15 | 4 | CC4 30sep24 9m | 19860 | 1 | 0 | 2025/06/05 21:55:18 | -6.078 | -24.12
85 | 16 | 1 | CC2 11jul24 0.1m | 19831 | 1 | 0 | 2025/06/05 22:04:08 | -5.578 | -22.046
86 | 16 | 2 | CC2 11jul24 0.1m | 19910 | 1 | 0 | 2025/06/05 22:12:57 | -5.595 | -21.703
87 | 16 | 3 | CC2 11jul24 0.1m | 19892 | 1 | 0 | 2025/06/05 22:21:46 | -5.481 | -21.67
88 | 16 | 4 | CC2 11jul24 0.1m | 19848 | 1 | 0 | 2025/06/05 22:30:36 | -5.564 | -21.501
89 | 17 | 1 | CP2 28oct24 0.1m | 19735 | 1 | 0 | 2025/06/05 22:39:25 | -5.357 | -20.724
90 | 17 | 2 | CP2 28oct24 0.1m | 19843 | 1 | 0 | 2025/06/05 22:48:16 | -5.309 | -20.255
91 | 17 | 3 | CP2 28oct24 0.1m | 19792 | 1 | 0 | 2025/06/05 22:57:06 | -5.275 | -20.628
92 | 17 | 4 | CP2 28oct24 0.1m | 19932 | 1 | 0 | 2025/06/05 23:05:57 | -5.297 | -20.181
93 | 18 | 1 | Blacksburg | 19743 | 1 | 0 | 2025/06/05 23:14:47 | -8.649 | -40.747
94 | 18 | 2 | Blacksburg | 20089 | 1 | 0 | 2025/06/05 23:23:37 | -8.822 | -42.457
95 | 18 | 3 | Blacksburg | 19801 | 1 | 0 | 2025/06/05 23:32:28 | -8.874 | -43.124
96 | 18 | 4 | Blacksburg | 19807 | 1 | 0 | 2025/06/05 23:41:18 | -8.856 | -42.996
97 | 19 | 1 | CP1 17dec24 0.1m | 19809 | 1 | 0 | 2025/06/05 23:50:08 | -7.892 | -32.569
98 | 19 | 2 | CP1 17dec24 0.1m | 20077 | 1 | 0 | 2025/06/05 23:58:59 | -7.834 | -31.595
99 | 19 | 3 | CP1 17dec24 0.1m | 19956 | 1 | 0 | 2025/06/06 00:07:50 | -7.815 | -31.851
100 | 19 | 4 | CP1 17dec24 0.1m | 20017 | 1 | 0 | 2025/06/06 00:16:40 | -7.749 | -31.223
101 | 20 | 1 | CS2 17dec24 0.1m | 19961 | 1 | 0 | 2025/06/06 00:25:32 | -7.819 | -31.413
102 | 20 | 2 | CS2 17dec24 0.1m | 19852 | 1 | 0 | 2025/06/06 00:34:23 | -7.814 | -31.423
103 | 20 | 3 | CS2 17dec24 0.1m | 19849 | 1 | 0 | 2025/06/06 00:43:14 | -7.818 | -31.532
104 | 20 | 4 | CS2 17dec24 0.1m | 19997 | 1 | 0 | 2025/06/06 00:52:05 | -7.872 | -31.328
105 | 21 | 1 | CC3 28oct24 0.1m | 19810 | 1 | 0 | 2025/06/06 01:00:57 | -5.484 | -21.545
106 | 21 | 2 | CC3 28oct24 0.1m | 20122 | 1 | 0 | 2025/06/06 01:09:48 | -5.287 | -20.95
107 | 21 | 3 | CC3 28oct24 0.1m | 19952 | 1 | 0 | 2025/06/06 01:18:39 | -5.408 | -20.663
108 | 21 | 4 | CC3 28oct24 0.1m | 19950 | 1 | 0 | 2025/06/06 01:27:30 | -5.354 | -20.555
109 | 22 | 1 | C50 17dec24 0.1m | 20031 | 1 | 0 | 2025/06/06 01:36:22 | -5.453 | -21.281
110 | 22 | 2 | C50 17dec24 0.1m | 19905 | 1 | 0 | 2025/06/06 01:45:13 | -5.546 | -21.245
111 | 22 | 3 | C50 17dec24 0.1m | 19924 | 1 | 0 | 2025/06/06 01:54:04 | -5.412 | -21.155
112 | 22 | 4 | C50 17dec24 0.1m | 19962 | 1 | 0 | 2025/06/06 02:02:57 | -5.441 | -21.401
113 | 23 | 1 | CC4 28oct24 BOT | 19863 | 1 | 0 | 2025/06/06 02:11:48 | -5.778 | -21.928
114 | 23 | 2 | CC4 28oct24 BOT | 19795 | 1 | 0 | 2025/06/06 02:20:39 | -5.705 | -22.179
115 | 23 | 3 | CC4 28oct24 BOT | 19883 | 1 | 0 | 2025/06/06 02:29:30 | -5.729 | -21.948
116 | 23 | 4 | CC4 28oct24 BOT | 19829 | 1 | 0 | 2025/06/06 02:38:21 | -5.655 | -22.396
117 | 24 | 1 | CC3 28oct24 BOT | 19957 | 1 | 0 | 2025/06/06 02:47:13 | -5.438 | -20.438
118 | 24 | 2 | CC3 28oct24 BOT | 19947 | 1 | 0 | 2025/06/06 02:56:03 | -5.318 | -20.291
119 | 24 | 3 | CC3 28oct24 BOT | 19904 | 1 | 0 | 2025/06/06 03:04:55 | -5.302 | -20.468
120 | 24 | 4 | CC3 28oct24 BOT | 19946 | 1 | 0 | 2025/06/06 03:13:46 | -5.293 | -20.459
121 | 25 | 1 | CC2 11jul24 0.1m - DUP | 19845 | 1 | 0 | 2025/06/06 03:22:38 | -5.541 | -21.671
122 | 25 | 2 | CC2 11jul24 0.1m - DUP | 19920 | 1 | 0 | 2025/06/06 03:31:29 | -5.573 | -21.51
123 | 25 | 3 | CC2 11jul24 0.1m - DUP | 19726 | 1 | 0 | 2025/06/06 03:40:21 | -5.634 | -21.338
124 | 25 | 4 | CC2 11jul24 0.1m - DUP | 19906 | 1 | 0 | 2025/06/06 03:49:13 | -5.556 | -21.578
125 | 26 | 1 | Hawaii | 19948 | 1 | 0 | 2025/06/06 03:58:05 | -4.891 | -8.458
126 | 26 | 2 | Hawaii | 19945 | 1 | 0 | 2025/06/06 04:06:56 | -4.844 | -6.978
127 | 26 | 3 | Hawaii | 19779 | 1 | 0 | 2025/06/06 04:15:48 | -4.778 | -6.872
128 | 26 | 4 | Hawaii | 19833 | 1 | 0 | 2025/06/06 04:24:40 | -4.938 | -6.582
129 | 27 | 1 | Blacksburg | 19849 | 1 | 0 | 2025/06/06 04:33:31 | -8.644 | -38.379
130 | 27 | 2 | Blacksburg | 19855 | 1 | 0 | 2025/06/06 04:42:23 | -8.896 | -41.522
131 | 27 | 3 | Blacksburg | 19818 | 1 | 0 | 2025/06/06 04:51:14 | -8.888 | -42.456
132 | 27 | 4 | Blacksburg | 19937 | 1 | 0 | 2025/06/06 05:00:07 | -8.961 | -42.463

## 18O
Line | Sample | Injection | Sample Name | Correction Type | Good | Ignore | d(18_16) Picarro Raw | Stdev Raw | d(18_16) Memory Corrected | Stdev Memory Corrected | d(18_16) Memory Corrected Drifts | d(18_16) Drift Corrected | Drift Corrected Drifts | d(18_16) Drift Corrected High, Low, and Drift | d(18_16) Reference Defined | d18O Final Corrected Value | Post run corrections:
1 | 1 | 1 | Blacksburg | Drift | 0 | -1 | -8.803 | [9 empty] | (1) Memory correction |  |  | (1) Graphs for memory correction
2 | 1 | 2 | Blacksburg | Drift | 1 | -1 | -8.992 | [9 empty] | Combined Stdev | 0.066447
3 | 1 | 3 | Blacksburg | Drift | 1 | -1 | -8.956 | [9 empty] | Date:
4 | 1 | 4 | Blacksburg | Drift | 1 | -1 | -9.005 | [9 empty] | injection | 18O memory coef
5 | 1 | 5 | Blacksburg | Drift | 1 | -1 | -9 | [9 empty] | 1 | 0.91621
6 | 1 | 6 | Blacksburg | Drift | 1 | -1 | -9.017 | [9 empty] | 2 | 0.97595
7 | 1 | 7 | Blacksburg | Drift | 1 | 0 | -8.94 |  | -8.94 |  | -8.94 | -8.9402 | -8.9402 |  |  | -7.5903 | 3 | 0.97987
8 | 1 | 8 | Blacksburg | Drift | 1 | 0 | -8.993 |  | -8.993 |  | -8.993 | -8.9933 | -8.9933 |  |  | -7.6407 | 4 | 0.9859
9 | 1 | 9 | Blacksburg | Drift | 1 | 0 | -9.018 |  | -9.018 |  | -9.018 | -9.0183 | -9.0183 |  |  | -7.6644 | 5 | 0.9859
10 | 1 | 10 | Blacksburg | Drift | 1 | 0 | -8.902 | 0.052201 | -8.902 | 0.052201 | -8.902 | -8.9024 | -8.9024 |  |  | -7.5543 | 6 | 0.9934
11 | 2 | 1 | Myrtle | High Reference | 1 | 0 | -3.761 |  | -3.3302 |  |  | -3.3306 |  | -3.3306 | -2.2908 | -2.2632 | 7 | 0.9934
12 | 2 | 2 | Myrtle | High Reference | 1 | 0 | -3.451 |  | -3.3199 |  |  | -3.3203 |  | -3.3203 | -2.2908 | -2.2535 | 8 | 0.9934
13 | 2 | 3 | Myrtle | High Reference | 1 | 0 | -3.545 |  | -3.4372 |  |  | -3.4376 |  | -3.4376 | -2.2908 | -2.3648 | 9 | 0.9934
14 | 2 | 4 | Myrtle | High Reference | 1 | 0 | -3.488 |  | -3.4117 |  |  | -3.4122 |  | -3.4122 | -2.2908 | -2.3406 | 10 | 1
15 | 2 | 5 | Myrtle | High Reference | 1 | 0 | -3.364 |  | -3.2859 |  |  | -3.2864 |  | -3.2864 | -2.2908 | -2.2213
16 | 2 | 6 | Myrtle | High Reference | 1 | 0 | -3.313 |  | -3.2761 |  |  | -3.2767 |  | -3.2767 | -2.2908 | -2.212 | (2) Drift correction
17 | 2 | 7 | Myrtle | High Reference | 1 | 0 | -3.381 |  | -3.3445 |  |  | -3.3451 |  | -3.3451 | -2.2908 | -2.277 | Slope m: | 3.5086e-05
18 | 2 | 8 | Myrtle | High Reference | 1 | 0 | -3.376 |  | -3.3395 |  |  | -3.3401 |  | -3.3401 | -2.2908 | -2.2723 | Stdev Mem Corr. | 0.047846
19 | 2 | 9 | Myrtle | High Reference | 1 | 0 | -3.455 |  | -3.419 |  |  | -3.4197 |  | -3.4197 | -2.2908 | -2.3478 | Stdev Drift Corr. | 0.047818
20 | 2 | 10 | Myrtle | High Reference | 1 | 0 | -3.335 | 0.13197 | -3.335 | 0.055174 |  | -3.3357 |  | -3.3357 | -2.2908 | -2.268
21 | 3 | 1 | Homer | Low Reference | 1 | 0 | -15.242 |  | -16.2397 |  |  | -16.2404 |  | -16.2404 | -14.5399 | -14.5228 | (3) Normalization to VSMOW scale
22 | 3 | 2 | Homer | Low Reference | 1 | 0 | -15.925 |  | -16.2277 |  |  | -16.2285 |  | -16.2285 | -14.5399 | -14.5115 | slope m | 0.94963
23 | 3 | 3 | Homer | Low Reference | 1 | 0 | -16.039 |  | -16.2947 |  |  | -16.2955 |  | -16.2955 | -14.5399 | -14.5751 | interception t | 0.89965
24 | 3 | 4 | Homer | Low Reference | 1 | 0 | -16.065 |  | -16.2445 |  |  | -16.2453 |  | -16.2453 | -14.5399 | -14.5274
25 | 3 | 5 | Homer | Low Reference | 1 | 0 | -16.05 |  | -16.2293 |  |  | -16.2302 |  | -16.2302 | -14.5399 | -14.513 | (4) Quality control |  | offset
26 | 3 | 6 | Homer | Low Reference | 1 | 0 | -16.122 |  | -16.2064 |  |  | -16.2073 |  | -16.2073 | -14.5399 | -14.4914 | QC Check Defined | -3.6108 | accepted ± 0.15
27 | 3 | 7 | Homer | Low Reference | 1 | 0 | -16.182 |  | -16.2668 |  |  | -16.2678 |  | -16.2678 | -14.5399 | -14.5487 | QC Check Vial 5 | -3.6944 | -0.083593
28 | 3 | 8 | Homer | Low Reference | 1 | 0 | -16.188 |  | -16.2729 |  |  | -16.2739 |  | -16.2739 | -14.5399 | -14.5545 | QC Check Vial 26 | -3.6993 | -0.088491
29 | 3 | 9 | Homer | Low Reference | 1 | 0 | -16.168 |  | -16.2527 |  |  | -16.2538 |  | -16.2538 | -14.5399 | -14.5354
30 | 3 | 10 | Homer | Low Reference | 1 | 0 | -16.265 | 0.29139 | -16.265 | 0.025808 |  | -16.2661 |  | -16.2661 | -14.5399 | -14.5471 | Drift Defined | -7.5615 | accepted ± 0.15
31 | 4 | 1 | Blacksburg | Drift | 1 | 0 | -9.49 |  | -8.9223 |  |  | -8.9234 |  | -8.9234 | -7.5615 | -7.5743 | Drift Average | -7.6131 | -0.051617
32 | 4 | 2 | Blacksburg | Drift | 1 | 0 | -9.082 |  | -8.9093 |  |  | -8.9104 |  | -8.9104 | -7.5615 | -7.562 | Drift Vial 4 | -7.5937 | -0.032181
33 | 4 | 3 | Blacksburg | Drift | 1 | 0 | -9.084 |  | -8.9395 |  |  | -8.9406 |  | -8.9406 | -7.5615 | -7.5907 | Drift Vial 9 | -7.6481 | -0.086652
34 | 4 | 4 | Blacksburg | Drift | 1 | 0 | -8.968 |  | -8.8651 |  |  | -8.8663 |  | -8.8663 | -7.5615 | -7.5201 | Drift Vial 18 | -7.5763 | -0.014818
35 | 4 | 5 | Blacksburg | Drift | 1 | 0 | -9.045 |  | -8.9432 |  |  | -8.9444 |  | -8.9444 | -7.5615 | -7.5943 | Drift Vial 27 | -7.635 | -0.073499
36 | 4 | 6 | Blacksburg | Drift | 1 | 0 | -8.954 |  | -8.9057 |  |  | -8.907 |  | -8.907 | -7.5615 | -7.5587
37 | 4 | 7 | Blacksburg | Drift | 1 | 0 | -8.991 |  | -8.943 |  |  | -8.9443 |  | -8.9443 | -7.5615 | -7.5941 | Reference Defined Values |  | d18O
38 | 4 | 8 | Blacksburg | Drift | 1 | 0 | -8.984 |  | -8.9359 |  |  | -8.9373 |  | -8.9373 | -7.5615 | -7.5875 | High Reference | Myrtle Beach | -2.2908
39 | 4 | 9 | Blacksburg | Drift | 1 | 0 | -8.981 |  | -8.9329 |  |  | -8.9343 |  | -8.9343 | -7.5615 | -7.5846 | Low Reference | Homer | -14.5399
40 | 4 | 10 | Blacksburg | Drift | 1 | 0 | -8.958 | 0.16079 | -8.958 | 0.026551 |  | -8.9594 |  | -8.9594 | -7.5615 | -7.6085 | Drift | Blacksburg | -7.5615
41 | 5 | 1 | Hawaii | QC Check | 1 | 0 | -5.19 |  | -4.8743 |  |  | -4.8757 |  |  |  | -3.7305 | QC Check | Hawaii | -3.6108
42 | 5 | 2 | Hawaii | QC Check | 1 | 0 | -4.904 |  | -4.8065 |  |  | -4.808 |  |  |  | -3.6662 | Calibration Date:
43 | 5 | 3 | Hawaii | QC Check | 1 | 0 | -4.899 |  | -4.8173 |  |  | -4.8188 |  |  |  | -3.6765
44 | 5 | 4 | Hawaii | QC Check | 1 | 0 | -4.904 | 0.14385 | -4.8468 | 0.030558 |  | -4.8484 |  |  |  | -3.7045
45 | 6 | 1 | C50 28oct24 1.5m |  | 1 | 0 | -5.478 |  | -5.5261 |  |  | -5.5277 |  |  |  | -4.3496
46 | 6 | 2 | C50 28oct24 1.5m |  | 1 | 0 | -5.405 |  | -5.417 |  |  | -5.4187 |  |  |  | -4.2461 | Denali | -19.9287
47 | 6 | 3 | C50 28oct24 1.5m |  | 1 | 0 | -5.489 |  | -5.5008 |  |  | -5.5024 |  |  |  | -4.3256 | Alberta | -18.8392
48 | 6 | 4 | C50 28oct24 1.5m |  | 1 | 0 | -5.508 | 0.04507 | -5.5165 | 0.049813 |  | -5.5182 |  |  |  | -4.3406 | MT BHB | -18.431
49 | 7 | 1 | C50 30sep24 1.5m |  | 1 | 0 | -5.063 |  | -5.0257 |  |  | -5.0274 |  |  |  | -3.8746 | Anchorage | -16.4204
50 | 7 | 2 | C50 30sep24 1.5m |  | 1 | 0 | -5.041 |  | -5.0298 |  |  | -5.0315 |  |  |  | -3.8784 | Homer | -14.5399
51 | 7 | 3 | C50 30sep24 1.5m |  | 1 | 0 | -5.109 |  | -5.101 |  |  | -5.1028 |  |  |  | -3.9461 | MT Creek | -13.3562
52 | 7 | 4 | C50 30sep24 1.5m |  | 1 | 0 | -5.027 | 0.03587 | -5.0202 | 0.03807 |  | -5.022 |  |  |  | -3.8694 | Nickel Mines | -7.876
53 | 8 | 1 | CC3 11jul24 1.5m |  | 1 | 0 | -5.52 |  | -5.5613 |  |  | -5.5632 |  |  |  | -4.3833 | Blacksburg | -7.5615
54 | 8 | 2 | CC3 11jul24 1.5m |  | 1 | 0 | -5.562 |  | -5.5749 |  |  | -5.5768 |  |  |  | -4.3962 | Denton's Well | -6.87
55 | 8 | 3 | CC3 11jul24 1.5m |  | 1 | 0 | -5.582 |  | -5.5932 |  |  | -5.5951 |  |  |  | -4.4136 | Hawaii | -3.6108
56 | 8 | 4 | CC3 11jul24 1.5m |  | 1 | 0 | -5.494 | 0.039846 | -5.5006 | 0.040115 |  | -5.5025 |  |  |  | -4.3257 | Myrtle Beach | -2.2908
57 | 9 | 1 | Blacksburg | Drift | 1 | 0 | -8.747 |  | -9.0196 |  | -9.0196 | -9.0216 | -9.0216 |  |  | -7.6675 | Heavy Oxygen | 3.1753
58 | 9 | 2 | Blacksburg | Drift | 1 | 0 | -8.87 |  | -8.9512 |  | -8.9512 | -8.9532 | -8.9532 |  |  | -7.6026 | Heavy Hydrogen | 11.1372
59 | 9 | 3 | Blacksburg | Drift | 1 | 0 | -8.999 |  | -9.0695 |  | -9.0695 | -9.0716 | -9.0716 |  |  | -7.715 | Andres Heavy Water | 17.5664
60 | 9 | 4 | Blacksburg | Drift | 1 | 0 | -8.908 | 0.10445 | -8.9561 | 0.056341 | -8.9561 | -8.9582 | -8.9582 |  |  | -7.6074 | Picarro ZERO | 0.3
61 | 10 | 1 | CS1 17dec24 0.1m |  | 1 | 0 | -7.978 |  | -7.9001 |  |  | -7.9022 |  |  |  | -6.6045 | Picarro MID | -20.6
62 | 10 | 2 | CS1 17dec24 0.1m |  | 1 | 0 | -7.921 |  | -7.8973 |  |  | -7.8994 |  |  |  | -6.6019 | Picarro DEPL | -29.6
63 | 10 | 3 | CS1 17dec24 0.1m |  | 1 | 0 | -7.92 |  | -7.9001 |  |  | -7.9023 |  |  |  | -6.6047
64 | 10 | 4 | CS1 17dec24 0.1m |  | 1 | 0 | -7.878 | 0.041056 | -7.8635 | 0.017888 |  | -7.8657 |  |  |  | -6.5699
65 | 11 | 1 | CC2 17dec24 0.1m |  | 1 | 0 | -7.771 |  | -7.762 |  |  | -7.7643 |  |  |  | -6.4736
66 | 11 | 2 | CC2 17dec24 0.1m |  | 1 | 0 | -7.806 |  | -7.8043 |  |  | -7.8066 |  |  |  | -6.5137
67 | 11 | 3 | CC2 17dec24 0.1m |  | 1 | 0 | -7.712 |  | -7.7087 |  |  | -7.711 |  |  |  | -6.423
68 | 11 | 4 | CC2 17dec24 0.1m |  | 1 | 0 | -7.927 | 0.090712 | -7.9277 | 0.093284 |  | -7.9301 |  |  |  | -6.631
69 | 12 | 1 | C50 17dec24 6m |  | 1 | 0 | -5.665 |  | -5.4755 |  |  | -5.4779 |  |  |  | -4.3023
70 | 12 | 2 | C50 17dec24 6m |  | 1 | 0 | -5.505 |  | -5.4468 |  |  | -5.4492 |  |  |  | -4.2751
71 | 12 | 3 | C50 17dec24 6m |  | 1 | 0 | -5.506 |  | -5.4573 |  |  | -5.4598 |  |  |  | -4.2851
72 | 12 | 4 | C50 17dec24 6m |  | 1 | 0 | -5.619 | 0.081016 | -5.5865 | 0.064412 |  | -5.589 |  |  |  | -4.4078
73 | 13 | 1 | CC4 17dec24 9m |  | 1 | 0 | -5.47 |  | -5.4575 |  |  | -5.4601 |  |  |  | -4.2854
74 | 13 | 2 | CC4 17dec24 9m |  | 1 | 0 | -5.417 |  | -5.4121 |  |  | -5.4147 |  |  |  | -4.2424
75 | 13 | 3 | CC4 17dec24 9m |  | 1 | 0 | -5.525 |  | -5.5231 |  |  | -5.5257 |  |  |  | -4.3478
76 | 13 | 4 | CC4 17dec24 9m |  | 1 | 0 | -5.468 | 0.044113 | -5.4659 | 0.045559 |  | -5.4685 |  |  |  | -4.2935
77 | 14 | 1 | CC4 17dec24 0.1m |  | 1 | 0 | -5.401 |  | -5.3954 |  |  | -5.3981 |  |  |  | -4.2265
78 | 14 | 2 | CC4 17dec24 0.1m |  | 1 | 0 | -5.47 |  | -5.47 |  |  | -5.4728 |  |  |  | -4.2975
79 | 14 | 3 | CC4 17dec24 0.1m |  | 1 | 0 | -5.323 |  | -5.3201 |  |  | -5.3229 |  |  |  | -4.1551
80 | 14 | 4 | CC4 17dec24 0.1m |  | 1 | 0 | -5.42 | 0.061049 | -5.4193 | 0.062403 |  | -5.4221 |  |  |  | -4.2494
81 | 15 | 1 | CC4 30sep24 9m |  | 1 | 0 | -5.989 |  | -6.0367 |  |  | -6.0395 |  |  |  | -4.8357
82 | 15 | 2 | CC4 30sep24 9m |  | 1 | 0 | -6.06 |  | -6.0754 |  |  | -6.0783 |  |  |  | -4.8725
83 | 15 | 3 | CC4 30sep24 9m |  | 1 | 0 | -6.057 |  | -6.0698 |  |  | -6.0727 |  |  |  | -4.8672
84 | 15 | 4 | CC4 30sep24 9m |  | 1 | 0 | -6.078 | 0.039115 | -6.0873 | 0.02167 |  | -6.0902 |  |  |  | -4.8838
85 | 16 | 1 | CC2 11jul24 0.1m |  | 1 | 0 | -5.578 |  | -5.5361 |  |  | -5.5391 |  |  |  | -4.3604
86 | 16 | 2 | CC2 11jul24 0.1m |  | 1 | 0 | -5.595 |  | -5.5834 |  |  | -5.5864 |  |  |  | -4.4054
87 | 16 | 3 | CC2 11jul24 0.1m |  | 1 | 0 | -5.481 |  | -5.469 |  |  | -5.472 |  |  |  | -4.2968
88 | 16 | 4 | CC2 11jul24 0.1m |  | 1 | 0 | -5.564 | 0.050613 | -5.5568 | 0.048877 |  | -5.5598 |  |  |  | -4.3802
89 | 17 | 1 | CP2 28oct24 0.1m |  | 1 | 0 | -5.357 |  | -5.3397 |  |  | -5.3428 |  |  |  | -4.174
90 | 17 | 2 | CP2 28oct24 0.1m |  | 1 | 0 | -5.309 |  | -5.3029 |  |  | -5.306 |  |  |  | -4.1391
91 | 17 | 3 | CP2 28oct24 0.1m |  | 1 | 0 | -5.275 |  | -5.2692 |  |  | -5.2724 |  |  |  | -4.1072
92 | 17 | 4 | CP2 28oct24 0.1m |  | 1 | 0 | -5.297 | 0.034655 | -5.2932 | 0.029269 |  | -5.2965 |  |  |  | -4.13
93 | 18 | 1 | Blacksburg | Drift | 1 | 0 | -8.649 |  | -8.9299 |  | -8.9299 | -8.9331 | -8.9331 |  |  | -7.5835
94 | 18 | 2 | Blacksburg | Drift | 1 | 0 | -8.822 |  | -8.9068 |  | -8.9068 | -8.9101 | -8.9101 |  |  | -7.5616
95 | 18 | 3 | Blacksburg | Drift | 1 | 0 | -8.874 |  | -8.946 |  | -8.946 | -8.9493 | -8.9493 |  |  | -7.5989
96 | 18 | 4 | Blacksburg | Drift | 1 | 0 | -8.856 | 0.10311 | -8.9062 | 0.019318 | -8.9062 | -8.9096 | -8.9096 |  |  | -7.5611
97 | 19 | 1 | CP1 17dec24 0.1m |  | 1 | 0 | -7.892 |  | -7.8112 |  |  | -7.8146 |  |  |  | -6.5214
98 | 19 | 2 | CP1 17dec24 0.1m |  | 1 | 0 | -7.834 |  | -7.8094 |  |  | -7.8129 |  |  |  | -6.5197
99 | 19 | 3 | CP1 17dec24 0.1m |  | 1 | 0 | -7.815 |  | -7.794 |  |  | -7.7975 |  |  |  | -6.5051
100 | 19 | 4 | CP1 17dec24 0.1m |  | 1 | 0 | -7.749 | 0.058938 | -7.7334 | 0.036576 |  | -7.7369 |  |  |  | -6.4476
101 | 20 | 1 | CS2 17dec24 0.1m |  | 1 | 0 | -7.819 |  | -7.8249 |  |  | -7.8284 |  |  |  | -6.5345
102 | 20 | 2 | CS2 17dec24 0.1m |  | 1 | 0 | -7.814 |  | -7.8156 |  |  | -7.8191 |  |  |  | -6.5257
103 | 20 | 3 | CS2 17dec24 0.1m |  | 1 | 0 | -7.818 |  | -7.8194 |  |  | -7.823 |  |  |  | -6.5293
104 | 20 | 4 | CS2 17dec24 0.1m |  | 1 | 0 | -7.872 | 0.027585 | -7.8737 | 0.027167 |  | -7.8774 |  |  |  | -6.581
105 | 21 | 1 | CC3 28oct24 0.1m |  | 1 | 0 | -5.484 |  | -5.2839 |  |  | -5.2876 |  |  |  | -4.1216
106 | 21 | 2 | CC3 28oct24 0.1m |  | 1 | 0 | -5.287 |  | -5.2248 |  |  | -5.2286 |  |  |  | -4.0656
107 | 21 | 3 | CC3 28oct24 0.1m |  | 1 | 0 | -5.408 |  | -5.3584 |  |  | -5.3622 |  |  |  | -4.1924
108 | 21 | 4 | CC3 28oct24 0.1m |  | 1 | 0 | -5.354 | 0.083432 | -5.3185 | 0.056598 |  | -5.3223 |  |  |  | -4.1546
109 | 22 | 1 | C50 17dec24 0.1m |  | 1 | 0 | -5.453 |  | -5.4613 |  |  | -5.4651 |  |  |  | -4.2902
110 | 22 | 2 | C50 17dec24 0.1m |  | 1 | 0 | -5.546 |  | -5.5506 |  |  | -5.5545 |  |  |  | -4.3751
111 | 22 | 3 | C50 17dec24 0.1m |  | 1 | 0 | -5.412 |  | -5.4132 |  |  | -5.4171 |  |  |  | -4.2446
112 | 22 | 4 | C50 17dec24 0.1m |  | 1 | 0 | -5.441 | 0.057948 | -5.4422 | 0.059262 |  | -5.4462 |  |  |  | -4.2722
113 | 23 | 1 | CC4 28oct24 BOT |  | 1 | 0 | -5.778 |  | -5.8062 |  |  | -5.8102 |  |  |  | -4.6179
114 | 23 | 2 | CC4 28oct24 BOT |  | 1 | 0 | -5.705 |  | -5.7113 |  |  | -5.7153 |  |  |  | -4.5278
115 | 23 | 3 | CC4 28oct24 BOT |  | 1 | 0 | -5.729 |  | -5.7348 |  |  | -5.7388 |  |  |  | -4.5501
116 | 23 | 4 | CC4 28oct24 BOT |  | 1 | 0 | -5.655 | 0.051162 | -5.658 | 0.061486 |  | -5.6621 |  |  |  | -4.4773
117 | 24 | 1 | CC3 28oct24 BOT |  | 1 | 0 | -5.438 |  | -5.4198 |  |  | -5.4239 |  |  |  | -4.2511
118 | 24 | 2 | CC3 28oct24 BOT |  | 1 | 0 | -5.318 |  | -5.3099 |  |  | -5.314 |  |  |  | -4.1467
119 | 24 | 3 | CC3 28oct24 BOT |  | 1 | 0 | -5.302 |  | -5.2949 |  |  | -5.2991 |  |  |  | -4.1325
120 | 24 | 4 | CC3 28oct24 BOT |  | 1 | 0 | -5.293 | 0.067628 | -5.2879 | 0.061813 |  | -5.2921 |  |  |  | -4.1259
121 | 25 | 1 | CC2 11jul24 0.1m - DUP |  | 1 | 0 | -5.541 |  | -5.5618 |  |  | -5.566 |  |  |  | -4.386
122 | 25 | 2 | CC2 11jul24 0.1m - DUP |  | 1 | 0 | -5.573 |  | -5.5797 |  |  | -5.584 |  |  |  | -4.4031
123 | 25 | 3 | CC2 11jul24 0.1m - DUP |  | 1 | 0 | -5.634 |  | -5.6409 |  |  | -5.6452 |  |  |  | -4.4612
124 | 25 | 4 | CC2 11jul24 0.1m - DUP |  | 1 | 0 | -5.556 | 0.040817 | -5.5597 | 0.037974 |  | -5.5641 |  |  |  | -4.3842
125 | 26 | 1 | Hawaii |  | 1 | 0 | -4.891 |  | -4.8353 |  |  | -4.8397 |  |  |  | -3.6963
126 | 26 | 2 | Hawaii |  | 1 | 0 | -4.844 |  | -4.8269 |  |  | -4.8313 |  |  |  | -3.6883
127 | 26 | 3 | Hawaii |  | 1 | 0 | -4.778 |  | -4.7623 |  |  | -4.7668 |  |  |  | -3.6271
128 | 26 | 4 | Hawaii |  | 1 | 0 | -4.938 | 0.0683 | -4.9293 | 0.068769 |  | -4.9338 |  |  |  | -3.7856
129 | 27 | 1 | Blacksburg | Drift | 1 | 0 | -8.644 |  | -8.9545 |  | -8.9545 | -8.959 | -8.959 |  |  | -7.6082
130 | 27 | 2 | Blacksburg | Drift | 1 | 0 | -8.896 |  | -8.9912 |  | -8.9912 | -8.9957 | -8.9957 |  |  | -7.643
131 | 27 | 3 | Blacksburg | Drift | 1 | 0 | -8.888 |  | -8.9675 |  | -8.9675 | -8.9721 | -8.9721 |  |  | -7.6205
132 | 27 | 4 | Blacksburg | Drift | 1 | 0 | -8.961 | 0.13939 | -9.0177 | 0.02783 | -9.0177 | -9.0224 | -9.0224 |  |  | -7.6683
Note on G2: Manually change first six injections to ignore "-1".
Note on G3: Manually change first six injections to ignore "-1".
Note on G4: Manually change first six injections to ignore "-1".
Note on G5: Manually change first six injections to ignore "-1".
Note on G6: Manually change first six injections to ignore "-1".
Note on G7: Manually change first six injections to ignore "-1".

## 2H
Line | Sample | Injection | Sample Name | Correction Type | Good | Ignore | dD_H) Picarro Raw | Stdev Raw | d(D_H) Memory Corrected | Stdev Memory Corrected | d(D_H) Memory Corrected Drift | d(D_H) Drift Corrected | Drift Corrected Drift | d(D_H) Drift Corrected High, Low, Drift | d(D_H) Reference Defined | dD Final Corrected Value | Post run corrections:
1 | 1 | 1 | Blacksburg | Drift | 0 | -1 | -45.843 | [9 empty] | (1) Memory correction |  |  | (1) Graphs for memory correction
2 | 1 | 2 | Blacksburg | Drift | 1 | -1 | -44.361 | [9 empty] | Combined Stdev | 0.2702
3 | 1 | 3 | Blacksburg | Drift | 1 | -1 | -44.276 | [9 empty] | Date:
4 | 1 | 4 | Blacksburg | Drift | 1 | -1 | -44.314 | [9 empty] | injection | 18O memory coef
5 | 1 | 5 | Blacksburg | Drift | 1 | -1 | -43.827 | [9 empty] | 1 | 0.85347
6 | 1 | 6 | Blacksburg | Drift | 1 | -1 | -43.933 | [9 empty] | 2 | 0.96225
7 | 1 | 7 | Blacksburg | Drift | 1 | 0 | -43.394 |  | -43.394 |  | -43.394 | -43.4344 | -43.4344 |  |  | -47.6397 | 3 | 0.98311
8 | 1 | 8 | Blacksburg | Drift | 1 | 0 | -43.998 |  | -43.998 |  | -43.998 | -44.0441 | -44.0441 |  |  | -48.2228 | 4 | 0.99117
9 | 1 | 9 | Blacksburg | Drift | 1 | 0 | -43.676 |  | -43.676 |  | -43.676 | -43.7279 | -43.7279 |  |  | -47.9204 | 5 | 0.9942
10 | 1 | 10 | Blacksburg | Drift | 1 | 0 | -44.029 | 0.29956 | -44.029 | 0.29956 | -44.029 | -44.0866 | -44.0866 |  |  | -48.2634 | 6 | 0.99693
11 | 2 | 1 | Myrtle | High Reference | 1 | 0 | -9.111 |  | -3.9944 |  |  | -4.0578 |  | -4.0578 | -10.453 | -9.9846 | 7 | 0.99693
12 | 2 | 2 | Myrtle | High Reference | 1 | 0 | -5.914 |  | -4.4753 |  |  | -4.5445 |  | -4.5445 | -10.453 | -10.4501 | 8 | 0.99785
13 | 2 | 3 | Myrtle | High Reference | 1 | 0 | -4.767 |  | -4.104 |  |  | -4.1789 |  | -4.1789 | -10.453 | -10.1005 | 9 | 0.99976
14 | 2 | 4 | Myrtle | High Reference | 1 | 0 | -4.409 |  | -4.0591 |  |  | -4.1398 |  | -4.1398 | -10.453 | -10.0631 | 10 | 1
15 | 2 | 5 | Myrtle | High Reference | 1 | 0 | -4.163 |  | -3.9316 |  |  | -4.0181 |  | -4.0181 | -10.453 | -9.9467
16 | 2 | 6 | Myrtle | High Reference | 1 | 0 | -4.12 |  | -3.9975 |  |  | -4.0897 |  | -4.0897 | -10.453 | -10.0152 | (2) Drift correction
17 | 2 | 7 | Myrtle | High Reference | 1 | 0 | -4.281 |  | -4.159 |  |  | -4.257 |  | -4.257 | -10.453 | -10.1751 | Slope m: | 0.0057646
18 | 2 | 8 | Myrtle | High Reference | 1 | 0 | -4.096 |  | -4.0099 |  |  | -4.1137 |  | -4.1137 | -10.453 | -10.0381 | Stdev Mem Corr. | 0.40453
19 | 2 | 9 | Myrtle | High Reference | 1 | 0 | -4.093 |  | -4.0833 |  |  | -4.1929 |  | -4.1929 | -10.453 | -10.1138 | Stdev Drift Corr. | 0.30244
20 | 2 | 10 | Myrtle | High Reference | 1 | 0 | -4.029 | 1.585 | -4.029 | 0.15167 |  | -4.1443 |  | -4.1443 | -10.453 | -10.0674
21 | 3 | 1 | Homer | Low Reference | 1 | 0 | -95.259 |  | -108.6271 |  |  | -108.7482 |  | -108.7482 | -110.5169 | -110.098 | (3) Normalization to VSMOW scale
22 | 3 | 2 | Homer | Low Reference | 1 | 0 | -105.092 |  | -108.9067 |  |  | -109.0335 |  | -109.0335 | -110.5169 | -110.3708 | slope m | 0.95628
23 | 3 | 3 | Homer | Low Reference | 1 | 0 | -107.017 |  | -108.7561 |  |  | -108.8887 |  | -108.8887 | -110.5169 | -110.2324 | interception t | -6.1042
24 | 3 | 4 | Homer | Low Reference | 1 | 0 | -107.939 |  | -108.8566 |  |  | -108.995 |  | -108.995 | -110.5169 | -110.334
25 | 3 | 5 | Homer | Low Reference | 1 | 0 | -108.058 |  | -108.6618 |  |  | -108.8059 |  | -108.8059 | -110.5169 | -110.1532 | (4) Quality control |  | offset
26 | 3 | 6 | Homer | Low Reference | 1 | 0 | -108.404 |  | -108.7244 |  |  | -108.8742 |  | -108.8742 | -110.5169 | -110.2185 | QC Check Defined | -13.2294 | accepted ±1.5
27 | 3 | 7 | Homer | Low Reference | 1 | 0 | -108.426 |  | -108.7464 |  |  | -108.9021 |  | -108.9021 | -110.5169 | -110.2452 | QC Check Vial 5 | -13.2623 | -0.032957
28 | 3 | 8 | Homer | Low Reference | 1 | 0 | -108.751 |  | -108.9767 |  |  | -109.1381 |  | -109.1381 | -110.5169 | -110.4709 | QC Check Vial 26 | -13.0259 | 0.20346
29 | 3 | 9 | Homer | Low Reference | 1 | 0 | -108.941 |  | -108.9664 |  |  | -109.1336 |  | -109.1336 | -110.5169 | -110.4665
30 | 3 | 10 | Homer | Low Reference | 1 | 0 | -108.882 | 4.1744 | -108.882 | 0.12411 |  | -109.0549 |  | -109.0549 | -110.5169 | -110.3913 | Drift Defined | -47.446 | accepted ±1.5
31 | 4 | 1 | Blacksburg | Drift | 1 | 0 | -51.762 |  | -43.3921 |  |  | -43.5708 |  | -43.5708 | -47.446 | -47.7701 | Drift Average | -47.9889 | -0.54284
32 | 4 | 2 | Blacksburg | Drift | 1 | 0 | -45.924 |  | -43.5476 |  |  | -43.7321 |  | -43.7321 | -47.446 | -47.9244 | Drift Vial 4 | -48.1766 | -0.73058
33 | 4 | 3 | Blacksburg | Drift | 1 | 0 | -44.616 |  | -43.5308 |  |  | -43.721 |  | -43.721 | -47.446 | -47.9138 | Drift Vial 9 | -47.7114 | -0.26541
34 | 4 | 4 | Blacksburg | Drift | 1 | 0 | -44.29 |  | -43.7196 |  |  | -43.9156 |  | -43.9156 | -47.446 | -48.0999 | Drift Vial 18 | -48.1678 | -0.72174
35 | 4 | 5 | Blacksburg | Drift | 1 | 0 | -43.863 |  | -43.4856 |  |  | -43.6874 |  | -43.6874 | -47.446 | -47.8816 | Drift Vial 27 | -47.8769 | -0.43092
36 | 4 | 6 | Blacksburg | Drift | 1 | 0 | -43.727 |  | -43.527 |  |  | -43.7345 |  | -43.7345 | -47.446 | -47.9267
37 | 4 | 7 | Blacksburg | Drift | 1 | 0 | -43.7 |  | -43.4999 |  |  | -43.7132 |  | -43.7132 | -47.446 | -47.9063 | Reference Defined Values |  | d2H
38 | 4 | 8 | Blacksburg | Drift | 1 | 0 | -44.084 |  | -43.9444 |  |  | -44.1634 |  | -44.1634 | -47.446 | -48.3369 | High Reference | North Myrtle Beach, SC | -10.453
39 | 4 | 9 | Blacksburg | Drift | 1 | 0 | -43.898 |  | -43.8823 |  |  | -44.1071 |  | -44.1071 | -47.446 | -48.283 | Low Reference | Homer, AK | -110.5169
40 | 4 | 10 | Blacksburg | Drift | 1 | 0 | -43.769 | 2.4808 | -43.769 | 0.18601 |  | -43.9996 |  | -43.9996 | -47.446 | -48.1802 | Drift | Blacksburg | -47.446
41 | 5 | 1 | Hawaii | QC Check | 1 | 0 | -11.789 |  | -7.1029 |  |  | -7.3392 |  |  |  | -13.1226 | QC Check | Hawaii | -13.2294
42 | 5 | 2 | Hawaii | QC Check | 1 | 0 | -8.366 |  | -7.0297 |  |  | -7.2718 |  |  |  | -13.0581 | Calibration Date:
43 | 5 | 3 | Hawaii | QC Check | 1 | 0 | -7.992 |  | -7.3879 |  |  | -7.6357 |  |  |  | -13.4061
44 | 5 | 4 | Hawaii | QC Check | 1 | 0 | -7.759 | 1.8916 | -7.441 | 0.20436 |  | -7.6946 |  |  |  | -13.4625
45 | 6 | 1 | C50 28oct24 1.5m |  | 1 | 0 | -19.234 |  | -20.9155 |  |  | -21.1749 |  |  |  | -26.3534
46 | 6 | 2 | C50 28oct24 1.5m |  | 1 | 0 | -20.629 |  | -21.1148 |  |  | -21.38 |  |  |  | -26.5495 | Denali National Park | -153.7172
47 | 6 | 3 | C50 28oct24 1.5m |  | 1 | 0 | -20.759 |  | -20.9785 |  |  | -21.2495 |  |  |  | -26.4247 | Edmonton, Alberta | -146.8245
48 | 6 | 4 | C50 28oct24 1.5m |  | 1 | 0 | -20.341 | 0.69352 | -20.4521 | 0.28769 |  | -20.7288 |  |  |  | -25.9268 | MT BHB | -140.5778
49 | 7 | 1 | C50 30sep24 1.5m |  | 1 | 0 | -19.41 |  | -19.2736 |  |  | -19.556 |  |  |  | -24.8053 | Anchorage, AK | -129.9028
50 | 7 | 2 | C50 30sep24 1.5m |  | 1 | 0 | -19.369 |  | -19.3323 |  |  | -19.6205 |  |  |  | -24.867 | Homer, AK | -110.5169
51 | 7 | 3 | C50 30sep24 1.5m |  | 1 | 0 | -19.418 |  | -19.4024 |  |  | -19.6964 |  |  |  | -24.9395 | MT Creek | -97.7428
52 | 7 | 4 | C50 30sep24 1.5m |  | 1 | 0 | -19.499 | 0.054412 | -19.4916 | 0.093893 |  | -19.7913 |  |  |  | -25.0303 | Nickel Mines, PA | -48.7951
53 | 8 | 1 | CC3 11jul24 1.5m |  | 1 | 0 | -21.289 |  | -21.5513 |  |  | -21.8568 |  |  |  | -27.0055 | Blacksburg | -47.446
54 | 8 | 2 | CC3 11jul24 1.5m |  | 1 | 0 | -21.666 |  | -21.7478 |  |  | -22.0591 |  |  |  | -27.1989 | Denton's Well | -41.771
55 | 8 | 3 | CC3 11jul24 1.5m |  | 1 | 0 | -21.572 |  | -21.607 |  |  | -21.9241 |  |  |  | -27.0698 | Hawaii | -13.2294
56 | 8 | 4 | CC3 11jul24 1.5m |  | 1 | 0 | -21.874 | 0.24286 | -21.895 | 0.15391 |  | -22.2178 |  |  |  | -27.3507 | North Myrtle Beach, SC | -10.453
57 | 9 | 1 | Blacksburg | Drift | 1 | 0 | -40.169 |  | -42.8498 |  | -42.8498 | -43.1784 | -43.1784 |  |  | -47.3949 | WQL Heavy Oxygen | -15.7881
58 | 9 | 2 | Blacksburg | Drift | 1 | 0 | -42.844 |  | -43.6355 |  | -43.6355 | -43.9699 | -43.9699 |  |  | -48.1518 | WQL Heavy Hydrogen | 4.0332
59 | 9 | 3 | Blacksburg | Drift | 1 | 0 | -42.625 |  | -42.9754 |  | -42.9754 | -43.3155 | -43.3155 |  |  | -47.526 | Andres Heavy Water | 20.5382
60 | 9 | 4 | Blacksburg | Drift | 1 | 0 | -43.041 | 1.3446 | -43.2279 | 0.34662 | -43.2279 | -43.5738 | -43.5738 |  |  | -47.773
61 | 10 | 1 | CS1 17dec24 0.1m |  | 1 | 0 | -33.069 |  | -31.6078 |  |  | -31.9594 |  |  |  | -36.6664
62 | 10 | 2 | CS1 17dec24 0.1m |  | 1 | 0 | -32.042 |  | -31.6268 |  |  | -31.9842 |  |  |  | -36.6902
63 | 10 | 3 | CS1 17dec24 0.1m |  | 1 | 0 | -31.588 |  | -31.3946 |  |  | -31.7578 |  |  |  | -36.4736
64 | 10 | 4 | CS1 17dec24 0.1m |  | 1 | 0 | -31.614 | 0.69253 | -31.5131 | 0.10634 |  | -31.882 |  |  |  | -36.5924
65 | 11 | 1 | CC2 17dec24 0.1m |  | 1 | 0 | -31.371 |  | -31.3354 |  |  | -31.7101 |  |  |  | -36.428
66 | 11 | 2 | CC2 17dec24 0.1m |  | 1 | 0 | -31.118 |  | -31.0993 |  |  | -31.4797 |  |  |  | -36.2077
67 | 11 | 3 | CC2 17dec24 0.1m |  | 1 | 0 | -30.863 |  | -30.8503 |  |  | -31.2365 |  |  |  | -35.9751
68 | 11 | 4 | CC2 17dec24 0.1m |  | 1 | 0 | -31.306 | 0.22784 | -31.3033 | 0.22377 |  | -31.6953 |  |  |  | -36.4138
69 | 12 | 1 | C50 17dec24 6m |  | 1 | 0 | -22.684 |  | -21.4206 |  |  | -21.8184 |  |  |  | -26.9687
70 | 12 | 2 | C50 17dec24 6m |  | 1 | 0 | -21.758 |  | -21.3976 |  |  | -21.8011 |  |  |  | -26.9522
71 | 12 | 3 | C50 17dec24 6m |  | 1 | 0 | -21.693 |  | -21.5307 |  |  | -21.94 |  |  |  | -27.085
72 | 12 | 4 | C50 17dec24 6m |  | 1 | 0 | -21.425 | 0.5486 | -21.3377 | 0.080633 |  | -21.7528 |  |  |  | -26.906
73 | 13 | 1 | CC4 17dec24 9m |  | 1 | 0 | -21.198 |  | -21.1647 |  |  | -21.5856 |  |  |  | -26.7461
74 | 13 | 2 | CC4 17dec24 9m |  | 1 | 0 | -21.048 |  | -21.0338 |  |  | -21.4604 |  |  |  | -26.6264
75 | 13 | 3 | CC4 17dec24 9m |  | 1 | 0 | -21.292 |  | -21.2898 |  |  | -21.7221 |  |  |  | -26.8767
76 | 13 | 4 | CC4 17dec24 9m |  | 1 | 0 | -21.209 | 0.10157 | -21.2071 | 0.10684 |  | -21.6452 |  |  |  | -26.8031
77 | 14 | 1 | CC4 17dec24 0.1m |  | 1 | 0 | -21.12 |  | -21.107 |  |  | -21.5508 |  |  |  | -26.7129
78 | 14 | 2 | CC4 17dec24 0.1m |  | 1 | 0 | -21.254 |  | -21.2557 |  |  | -21.7053 |  |  |  | -26.8606
79 | 14 | 3 | CC4 17dec24 0.1m |  | 1 | 0 | -21.368 |  | -21.3707 |  |  | -21.8261 |  |  |  | -26.9761
80 | 14 | 4 | CC4 17dec24 0.1m |  | 1 | 0 | -21.062 | 0.13733 | -21.0607 | 0.14176 |  | -21.5219 |  |  |  | -26.6852
81 | 15 | 1 | CC4 30sep24 9m |  | 1 | 0 | -24.077 |  | -24.5188 |  |  | -24.9857 |  |  |  | -29.9976
82 | 15 | 2 | CC4 30sep24 9m |  | 1 | 0 | -24.496 |  | -24.6256 |  |  | -25.0983 |  |  |  | -30.1053
83 | 15 | 3 | CC4 30sep24 9m |  | 1 | 0 | -24.438 |  | -24.495 |  |  | -24.9735 |  |  |  | -29.9859
84 | 15 | 4 | CC4 30sep24 9m |  | 1 | 0 | -24.12 | 0.21479 | -24.147 | 0.20765 |  | -24.6312 |  |  |  | -29.6586
85 | 16 | 1 | CC2 11jul24 0.1m |  | 1 | 0 | -22.046 |  | -21.7421 |  |  | -22.2321 |  |  |  | -27.3644
86 | 16 | 2 | CC2 11jul24 0.1m |  | 1 | 0 | -21.703 |  | -21.6118 |  |  | -22.1075 |  |  |  | -27.2452
87 | 16 | 3 | CC2 11jul24 0.1m |  | 1 | 0 | -21.67 |  | -21.6286 |  |  | -22.1301 |  |  |  | -27.2669
88 | 16 | 4 | CC2 11jul24 0.1m |  | 1 | 0 | -21.501 | 0.22849 | -21.4779 | 0.10825 |  | -21.9852 |  |  |  | -27.1282
89 | 17 | 1 | CP2 28oct24 0.1m |  | 1 | 0 | -20.724 |  | -20.6101 |  |  | -21.1232 |  |  |  | -26.3039
90 | 17 | 2 | CP2 28oct24 0.1m |  | 1 | 0 | -20.255 |  | -20.208 |  |  | -20.7268 |  |  |  | -25.9249
91 | 17 | 3 | CP2 28oct24 0.1m |  | 1 | 0 | -20.628 |  | -20.6133 |  |  | -21.1378 |  |  |  | -26.3179
92 | 17 | 4 | CP2 28oct24 0.1m |  | 1 | 0 | -20.181 | 0.26902 | -20.1693 | 0.24476 |  | -20.6997 |  |  |  | -25.899
93 | 18 | 1 | Blacksburg | Drift | 1 | 0 | -40.747 |  | -43.7606 |  | -43.7606 | -44.2967 | -44.2967 |  |  | -48.4643
94 | 18 | 2 | Blacksburg | Drift | 1 | 0 | -42.457 |  | -43.2978 |  | -43.2978 | -43.8397 | -43.8397 |  |  | -48.0273
95 | 18 | 3 | Blacksburg | Drift | 1 | 0 | -43.124 |  | -43.5114 |  | -43.5114 | -44.0591 | -44.0591 |  |  | -48.2371
96 | 18 | 4 | Blacksburg | Drift | 1 | 0 | -42.996 | 1.0948 | -43.1975 | 0.2496 | -43.1975 | -43.7509 | -43.7509 |  |  | -47.9424
97 | 19 | 1 | CP1 17dec24 0.1m |  | 1 | 0 | -32.569 |  | -31.0411 |  |  | -31.6003 |  |  |  | -36.323
98 | 19 | 2 | CP1 17dec24 0.1m |  | 1 | 0 | -31.595 |  | -31.1647 |  |  | -31.7296 |  |  |  | -36.4466
99 | 19 | 3 | CP1 17dec24 0.1m |  | 1 | 0 | -31.851 |  | -31.6628 |  |  | -32.2335 |  |  |  | -36.9285
100 | 19 | 4 | CP1 17dec24 0.1m |  | 1 | 0 | -31.223 | 0.5682 | -31.119 | 0.28192 |  | -31.6955 |  |  |  | -36.414
101 | 20 | 1 | CS2 17dec24 0.1m |  | 1 | 0 | -31.413 |  | -31.4408 |  |  | -32.0231 |  |  |  | -36.7273
102 | 20 | 2 | CS2 17dec24 0.1m |  | 1 | 0 | -31.423 |  | -31.4305 |  |  | -32.0185 |  |  |  | -36.7229
103 | 20 | 3 | CS2 17dec24 0.1m |  | 1 | 0 | -31.532 |  | -31.5372 |  |  | -32.131 |  |  |  | -36.8305
104 | 20 | 4 | CS2 17dec24 0.1m |  | 1 | 0 | -31.328 | 0.08367 | -31.3289 | 0.085152 |  | -31.9284 |  |  |  | -36.6368
105 | 21 | 1 | CC3 28oct24 0.1m |  | 1 | 0 | -21.545 |  | -20.1115 |  |  | -20.7168 |  |  |  | -25.9153
106 | 21 | 2 | CC3 28oct24 0.1m |  | 1 | 0 | -20.95 |  | -20.5583 |  |  | -21.1693 |  |  |  | -26.3481
107 | 21 | 3 | CC3 28oct24 0.1m |  | 1 | 0 | -20.663 |  | -20.4829 |  |  | -21.0997 |  |  |  | -26.2815
108 | 21 | 4 | CC3 28oct24 0.1m |  | 1 | 0 | -20.555 | 0.44367 | -20.4599 | 0.19893 |  | -21.0824 |  |  |  | -26.265
109 | 22 | 1 | C50 17dec24 0.1m |  | 1 | 0 | -21.281 |  | -21.3874 |  |  | -22.0157 |  |  |  | -27.1575
110 | 22 | 2 | C50 17dec24 0.1m |  | 1 | 0 | -21.245 |  | -21.271 |  |  | -21.9052 |  |  |  | -27.0517
111 | 22 | 3 | C50 17dec24 0.1m |  | 1 | 0 | -21.155 |  | -21.1651 |  |  | -21.805 |  |  |  | -26.9559
112 | 22 | 4 | C50 17dec24 0.1m |  | 1 | 0 | -21.401 | 0.10187 | -21.4085 | 0.1128 |  | -22.0541 |  |  |  | -27.1942
113 | 23 | 1 | CC4 28oct24 BOT |  | 1 | 0 | -21.928 |  | -22.0052 |  |  | -22.6566 |  |  |  | -27.7703
114 | 23 | 2 | CC4 28oct24 BOT |  | 1 | 0 | -22.179 |  | -22.2084 |  |  | -22.8655 |  |  |  | -27.9701
115 | 23 | 3 | CC4 28oct24 BOT |  | 1 | 0 | -21.948 |  | -21.9572 |  |  | -22.6202 |  |  |  | -27.7355
116 | 23 | 4 | CC4 28oct24 BOT |  | 1 | 0 | -22.396 | 0.22053 | -22.4048 | 0.20518 |  | -23.0735 |  |  |  | -28.169
117 | 24 | 1 | CC3 28oct24 BOT |  | 1 | 0 | -20.438 |  | -20.1511 |  |  | -20.8255 |  |  |  | -26.0193
118 | 24 | 2 | CC3 28oct24 BOT |  | 1 | 0 | -20.291 |  | -20.2115 |  |  | -20.8918 |  |  |  | -26.0826
119 | 24 | 3 | CC3 28oct24 BOT |  | 1 | 0 | -20.468 |  | -20.4354 |  |  | -21.1214 |  |  |  | -26.3023
120 | 24 | 4 | CC3 28oct24 BOT |  | 1 | 0 | -20.459 | 0.082958 | -20.4419 | 0.15064 |  | -21.1336 |  |  |  | -26.3139
121 | 25 | 1 | CC2 11jul24 0.1m - DUP |  | 1 | 0 | -21.671 |  | -21.8486 |  |  | -22.5461 |  |  |  | -27.6647
122 | 25 | 2 | CC2 11jul24 0.1m - DUP |  | 1 | 0 | -21.51 |  | -21.5497 |  |  | -22.253 |  |  |  | -27.3843
123 | 25 | 3 | CC2 11jul24 0.1m - DUP |  | 1 | 0 | -21.338 |  | -21.3528 |  |  | -22.0619 |  |  |  | -27.2016
124 | 25 | 4 | CC2 11jul24 0.1m - DUP |  | 1 | 0 | -21.578 | 0.14061 | -21.5879 | 0.20383 |  | -22.3027 |  |  |  | -27.4319
125 | 26 | 1 | Hawaii |  | 1 | 0 | -8.458 |  | -6.5355 |  |  | -7.2561 |  |  |  | -13.0431
126 | 26 | 2 | Hawaii |  | 1 | 0 | -6.978 |  | -6.4269 |  |  | -7.1533 |  |  |  | -12.9448
127 | 26 | 3 | Hawaii |  | 1 | 0 | -6.872 |  | -6.6237 |  |  | -7.3558 |  |  |  | -13.1384
128 | 26 | 4 | Hawaii |  | 1 | 0 | -6.582 | 0.8405 | -6.4496 | 0.089666 |  | -7.1874 |  |  |  | -12.9775
129 | 27 | 1 | Blacksburg | Drift | 1 | 0 | -38.379 |  | -43.0383 |  | -43.0383 | -43.7819 | -43.7819 |  |  | -47.972
130 | 27 | 2 | Blacksburg | Drift | 1 | 0 | -41.522 |  | -42.8408 |  | -42.8408 | -43.5902 | -43.5902 |  |  | -47.7887
131 | 27 | 3 | Blacksburg | Drift | 1 | 0 | -42.456 |  | -43.0618 |  | -43.0618 | -43.817 | -43.817 |  |  | -48.0055
132 | 27 | 4 | Blacksburg | Drift | 1 | 0 | -42.463 | 1.9351 | -42.7799 | 0.14093 | -42.7799 | -43.5408 | -43.5408 |  |  | -47.7414
Note on G2: Manually change first six injections to ignore "-1".
Note on G3: Manually change first six injections to ignore "-1".
Note on G4: Manually change first six injections to ignore "-1".
Note on G5: Manually change first six injections to ignore "-1".
Note on G6: Manually change first six injections to ignore "-1".
Note on G7: Manually change first six injections to ignore "-1".

## RESULTS
| Analyst: | DWH |  |  |  | Complete Sequence |  |  | Run Summary | 18O | 18O | 2H | 2H
| Analysis Date: | 2025-06-05 |  |  |  | Line | Inj Number | Identifier | Ignore | Ignore | d18O Final Value | d2H Final Value | Ignore
| Project: | CCR |  |  |  | 1 | 1 | Blacksburg | -1 | -1 |  |  | -1
[6 empty] | 2 | 2 | Blacksburg | 0 | -1 |  |  | -1
[6 empty] | 3 | 3 | Blacksburg | 0 | -1 |  |  | -1
Vial | Sample Name | d18O | 1 s.d. | d2H | 1 s.d. | 4 | 4 | Blacksburg | 0 | -1 |  |  | -1
|  | ‰ vs. VSMOW |  | ‰ vs. VSMOW |  | 5 | 5 | Blacksburg | 0 | -1 |  |  | -1
1 | Blacksburg | -7.6124 | 0.049562 | -48.0116 | 0.29136 | 6 | 6 | Blacksburg | 0 | -1 |  |  | -1
2 | Myrtle | -2.2821 | 0.052396 | -10.0955 | 0.14081 | 7 | 7 | Blacksburg | 0 | 0 | -7.5903 | -47.6397 | 0
3 | Homer | -14.5327 | 0.024533 | -110.2981 | 0.12823 | 8 | 8 | Blacksburg | 0 | 0 | -7.6407 | -48.2228 | 0
4 | Blacksburg | -7.5775 | 0.02526 | -48.0223 | 0.19008 | 9 | 9 | Blacksburg | 0 | 0 | -7.6644 | -47.9204 | 0
5 | Hawaii | -3.6944 | 0.029006 | -13.2623 | 0.20162 | 10 | 10 | Blacksburg | 0 | 0 | -7.5543 | -48.2634 | 0
6 | C50 28oct24 1.5m | -4.3155 | 0.047311 | -26.3136 | 0.27029 | 11 | 1 | Myrtle | 0 | 0 | -2.2632 | -9.9846 | 0
7 | C50 30sep24 1.5m | -3.8921 | 0.03616 | -24.9105 | 0.096875 | 12 | 2 | Myrtle | 0 | 0 | -2.2535 | -10.4501 | 0
8 | CC3 11jul24 1.5m | -4.3797 | 0.038072 | -27.1562 | 0.15256 | 13 | 3 | Myrtle | 0 | 0 | -2.3648 | -10.1005 | 0
9 | Blacksburg | -7.6481 | 0.053496 | -47.7114 | 0.3328 | 14 | 4 | Myrtle | 0 | 0 | -2.3406 | -10.0631 | 0
10 | CS1 17dec24 0.1m | -6.5953 | 0.016954 | -36.6056 | 0.097384 | 15 | 5 | Myrtle | 0 | 0 | -2.2213 | -9.9467 | 0
11 | CC2 17dec24 0.1m | -6.5103 | 0.08861 | -36.2562 | 0.21268 | 16 | 6 | Myrtle | 0 | 0 | -2.212 | -10.0152 | 0
12 | C50 17dec24 6m | -4.3176 | 0.061198 | -26.978 | 0.076114 | 17 | 7 | Myrtle | 0 | 0 | -2.277 | -10.1751 | 0
13 | CC4 17dec24 9m | -4.2922 | 0.043281 | -26.7631 | 0.10565 | 18 | 8 | Myrtle | 0 | 0 | -2.2723 | -10.0381 | 0
14 | CC4 17dec24 0.1m | -4.2321 | 0.059253 | -26.8087 | 0.13559 | 19 | 9 | Myrtle | 0 | 0 | -2.3478 | -10.1138 | 0
15 | CC4 30sep24 9m | -4.8648 | 0.020616 | -29.9368 | 0.19311 | 20 | 10 | Myrtle | 0 | 0 | -2.268 | -10.0674 | 0
16 | CC2 11jul24 0.1m | -4.3607 | 0.046409 | -27.2512 | 0.096968 | 21 | 1 | Homer | 0 | 0 | -14.5228 | -110.098 | 0
17 | CP2 28oct24 0.1m | -4.1376 | 0.027762 | -26.1114 | 0.2307 | 22 | 2 | Homer | 0 | 0 | -14.5115 | -110.3708 | 0
18 | Blacksburg | -7.5763 | 0.018336 | -48.1678 | 0.2333 | 23 | 3 | Homer | 0 | 0 | -14.5751 | -110.2324 | 0
19 | CP1 17dec24 0.1m | -6.4984 | 0.034696 | -36.528 | 0.27206 | 24 | 4 | Homer | 0 | 0 | -14.5274 | -110.334 | 0
20 | CS2 17dec24 0.1m | -6.5426 | 0.02583 | -36.7294 | 0.079239 | 25 | 5 | Homer | 0 | 0 | -14.513 | -110.1532 | 0
21 | CC3 28oct24 0.1m | -4.1335 | 0.053771 | -26.2025 | 0.19479 | 26 | 6 | Homer | 0 | 0 | -14.4914 | -110.2185 | 0
22 | C50 17dec24 0.1m | -4.2955 | 0.056259 | -27.0898 | 0.10776 | 27 | 7 | Homer | 0 | 0 | -14.5487 | -110.2452 | 0
23 | CC4 28oct24 BOT | -4.5433 | 0.058351 | -27.9112 | 0.20053 | 28 | 8 | Homer | 0 | 0 | -14.5545 | -110.4709 | 0
24 | CC3 28oct24 BOT | -4.1641 | 0.058662 | -26.1795 | 0.15076 | 29 | 9 | Homer | 0 | 0 | -14.5354 | -110.4665 | 0
25 | CC2 11jul24 0.1m - DUP | -4.4086 | 0.03607 | -27.4206 | 0.19058 | 30 | 10 | Homer | 0 | 0 | -14.5471 | -110.3913 | 0
26 | Hawaii | -3.6993 | 0.065323 | -13.0259 | 0.085415 | 31 | 1 | Blacksburg | 0 | 0 | -7.5743 | -47.7701 | 0
27 | Blacksburg | -7.635 | 0.026462 | -47.8769 | 0.1313 | 32 | 2 | Blacksburg | 0 | 0 | -7.562 | -47.9244 | 0
[6 empty] | 33 | 3 | Blacksburg | 0 | 0 | -7.5907 | -47.9138 | 0
[6 empty] | 34 | 4 | Blacksburg | 0 | 0 | -7.5201 | -48.0999 | 0
[6 empty] | 35 | 5 | Blacksburg | 0 | 0 | -7.5943 | -47.8816 | 0
| Quality Assurance Verification |  |  |  |  | 36 | 6 | Blacksburg | 0 | 0 | -7.5587 | -47.9267 | 0
| Metric |  | Reviewer 1 | Reviewer 2 |  | 37 | 7 | Blacksburg | 0 | 0 | -7.5941 | -47.9063 | 0
| d18O Cal Curve R^2 | 0.999+ | pass |  |  | 38 | 8 | Blacksburg | 0 | 0 | -7.5875 | -48.3369 | 0
| d18O QC checks | < 0.15 | pass |  |  | 39 | 9 | Blacksburg | 0 | 0 | -7.5846 | -48.283 | 0
| dD Cal Curve R^2 | 0.999+ | pass |  |  | 40 | 10 | Blacksburg | 0 | 0 | -7.6085 | -48.1802 | 0
| dD QC checks | < 1.5 | pass |  |  | 41 | 1 | Hawaii | 0 | 0 | -3.7305 | -13.1226 | 0
| Initials |  | DWH |  |  | 42 | 2 | Hawaii | 0 | 0 | -3.6662 | -13.0581 | 0
| Date |  | 2025-06-10 |  |  | 43 | 3 | Hawaii | 0 | 0 | -3.6765 | -13.4061 | 0
| Comments: |  |  |  |  | 44 | 4 | Hawaii | 0 | 0 | -3.7045 | -13.4625 | 0
[6 empty] | 45 | 1 | C50 28oct24 1.5m | 0 | 0 | -4.3496 | -26.3534 | 0
[6 empty] | 46 | 2 | C50 28oct24 1.5m | 0 | 0 | -4.2461 | -26.5495 | 0
[6 empty] | 47 | 3 | C50 28oct24 1.5m | 0 | 0 | -4.3256 | -26.4247 | 0
[6 empty] | 48 | 4 | C50 28oct24 1.5m | 0 | 0 | -4.3406 | -25.9268 | 0
[6 empty] | 49 | 1 | C50 30sep24 1.5m | 0 | 0 | -3.8746 | -24.8053 | 0
[6 empty] | 50 | 2 | C50 30sep24 1.5m | 0 | 0 | -3.8784 | -24.867 | 0
[6 empty] | 51 | 3 | C50 30sep24 1.5m | 0 | 0 | -3.9461 | -24.9395 | 0
[6 empty] | 52 | 4 | C50 30sep24 1.5m | 0 | 0 | -3.8694 | -25.0303 | 0
[6 empty] | 53 | 1 | CC3 11jul24 1.5m | 0 | 0 | -4.3833 | -27.0055 | 0
[6 empty] | 54 | 2 | CC3 11jul24 1.5m | 0 | 0 | -4.3962 | -27.1989 | 0
[6 empty] | 55 | 3 | CC3 11jul24 1.5m | 0 | 0 | -4.4136 | -27.0698 | 0
[6 empty] | 56 | 4 | CC3 11jul24 1.5m | 0 | 0 | -4.3257 | -27.3507 | 0
[6 empty] | 57 | 1 | Blacksburg | 0 | 0 | -7.6675 | -47.3949 | 0
[6 empty] | 58 | 2 | Blacksburg | 0 | 0 | -7.6026 | -48.1518 | 0
[6 empty] | 59 | 3 | Blacksburg | 0 | 0 | -7.715 | -47.526 | 0
[6 empty] | 60 | 4 | Blacksburg | 0 | 0 | -7.6074 | -47.773 | 0
[6 empty] | 61 | 1 | CS1 17dec24 0.1m | 0 | 0 | -6.6045 | -36.6664 | 0
[6 empty] | 62 | 2 | CS1 17dec24 0.1m | 0 | 0 | -6.6019 | -36.6902 | 0
[6 empty] | 63 | 3 | CS1 17dec24 0.1m | 0 | 0 | -6.6047 | -36.4736 | 0
[6 empty] | 64 | 4 | CS1 17dec24 0.1m | 0 | 0 | -6.5699 | -36.5924 | 0
[6 empty] | 65 | 1 | CC2 17dec24 0.1m | 0 | 0 | -6.4736 | -36.428 | 0
[6 empty] | 66 | 2 | CC2 17dec24 0.1m | 0 | 0 | -6.5137 | -36.2077 | 0
[6 empty] | 67 | 3 | CC2 17dec24 0.1m | 0 | 0 | -6.423 | -35.9751 | 0
[6 empty] | 68 | 4 | CC2 17dec24 0.1m | 0 | 0 | -6.631 | -36.4138 | 0
[6 empty] | 69 | 1 | C50 17dec24 6m | 0 | 0 | -4.3023 | -26.9687 | 0
[6 empty] | 70 | 2 | C50 17dec24 6m | 0 | 0 | -4.2751 | -26.9522 | 0
[6 empty] | 71 | 3 | C50 17dec24 6m | 0 | 0 | -4.2851 | -27.085 | 0
[6 empty] | 72 | 4 | C50 17dec24 6m | 0 | 0 | -4.4078 | -26.906 | 0
[6 empty] | 73 | 1 | CC4 17dec24 9m | 0 | 0 | -4.2854 | -26.7461 | 0
[6 empty] | 74 | 2 | CC4 17dec24 9m | 0 | 0 | -4.2424 | -26.6264 | 0
[6 empty] | 75 | 3 | CC4 17dec24 9m | 0 | 0 | -4.3478 | -26.8767 | 0
[6 empty] | 76 | 4 | CC4 17dec24 9m | 0 | 0 | -4.2935 | -26.8031 | 0
[6 empty] | 77 | 1 | CC4 17dec24 0.1m | 0 | 0 | -4.2265 | -26.7129 | 0
[6 empty] | 78 | 2 | CC4 17dec24 0.1m | 0 | 0 | -4.2975 | -26.8606 | 0
[6 empty] | 79 | 3 | CC4 17dec24 0.1m | 0 | 0 | -4.1551 | -26.9761 | 0
[6 empty] | 80 | 4 | CC4 17dec24 0.1m | 0 | 0 | -4.2494 | -26.6852 | 0
[6 empty] | 81 | 1 | CC4 30sep24 9m | 0 | 0 | -4.8357 | -29.9976 | 0
[6 empty] | 82 | 2 | CC4 30sep24 9m | 0 | 0 | -4.8725 | -30.1053 | 0
[6 empty] | 83 | 3 | CC4 30sep24 9m | 0 | 0 | -4.8672 | -29.9859 | 0
[6 empty] | 84 | 4 | CC4 30sep24 9m | 0 | 0 | -4.8838 | -29.6586 | 0
[6 empty] | 85 | 1 | CC2 11jul24 0.1m | 0 | 0 | -4.3604 | -27.3644 | 0
[6 empty] | 86 | 2 | CC2 11jul24 0.1m | 0 | 0 | -4.4054 | -27.2452 | 0
[6 empty] | 87 | 3 | CC2 11jul24 0.1m | 0 | 0 | -4.2968 | -27.2669 | 0
[6 empty] | 88 | 4 | CC2 11jul24 0.1m | 0 | 0 | -4.3802 | -27.1282 | 0
[6 empty] | 89 | 1 | CP2 28oct24 0.1m | 0 | 0 | -4.174 | -26.3039 | 0
[6 empty] | 90 | 2 | CP2 28oct24 0.1m | 0 | 0 | -4.1391 | -25.9249 | 0
[6 empty] | 91 | 3 | CP2 28oct24 0.1m | 0 | 0 | -4.1072 | -26.3179 | 0
[6 empty] | 92 | 4 | CP2 28oct24 0.1m | 0 | 0 | -4.13 | -25.899 | 0
[6 empty] | 93 | 1 | Blacksburg | 0 | 0 | -7.5835 | -48.4643 | 0
[6 empty] | 94 | 2 | Blacksburg | 0 | 0 | -7.5616 | -48.0273 | 0
[6 empty] | 95 | 3 | Blacksburg | 0 | 0 | -7.5989 | -48.2371 | 0
[6 empty] | 96 | 4 | Blacksburg | 0 | 0 | -7.5611 | -47.9424 | 0
[6 empty] | 97 | 1 | CP1 17dec24 0.1m | 0 | 0 | -6.5214 | -36.323 | 0
[6 empty] | 98 | 2 | CP1 17dec24 0.1m | 0 | 0 | -6.5197 | -36.4466 | 0
[6 empty] | 99 | 3 | CP1 17dec24 0.1m | 0 | 0 | -6.5051 | -36.9285 | 0
[6 empty] | 100 | 4 | CP1 17dec24 0.1m | 0 | 0 | -6.4476 | -36.414 | 0
[6 empty] | 101 | 1 | CS2 17dec24 0.1m | 0 | 0 | -6.5345 | -36.7273 | 0
[6 empty] | 102 | 2 | CS2 17dec24 0.1m | 0 | 0 | -6.5257 | -36.7229 | 0
[6 empty] | 103 | 3 | CS2 17dec24 0.1m | 0 | 0 | -6.5293 | -36.8305 | 0
[6 empty] | 104 | 4 | CS2 17dec24 0.1m | 0 | 0 | -6.581 | -36.6368 | 0
[6 empty] | 105 | 1 | CC3 28oct24 0.1m | 0 | 0 | -4.1216 | -25.9153 | 0
[6 empty] | 106 | 2 | CC3 28oct24 0.1m | 0 | 0 | -4.0656 | -26.3481 | 0
[6 empty] | 107 | 3 | CC3 28oct24 0.1m | 0 | 0 | -4.1924 | -26.2815 | 0
[6 empty] | 108 | 4 | CC3 28oct24 0.1m | 0 | 0 | -4.1546 | -26.265 | 0
[6 empty] | 109 | 1 | C50 17dec24 0.1m | 0 | 0 | -4.2902 | -27.1575 | 0
[6 empty] | 110 | 2 | C50 17dec24 0.1m | 0 | 0 | -4.3751 | -27.0517 | 0
[6 empty] | 111 | 3 | C50 17dec24 0.1m | 0 | 0 | -4.2446 | -26.9559 | 0
[6 empty] | 112 | 4 | C50 17dec24 0.1m | 0 | 0 | -4.2722 | -27.1942 | 0
[6 empty] | 113 | 1 | CC4 28oct24 BOT | 0 | 0 | -4.6179 | -27.7703 | 0
[6 empty] | 114 | 2 | CC4 28oct24 BOT | 0 | 0 | -4.5278 | -27.9701 | 0
[6 empty] | 115 | 3 | CC4 28oct24 BOT | 0 | 0 | -4.5501 | -27.7355 | 0
[6 empty] | 116 | 4 | CC4 28oct24 BOT | 0 | 0 | -4.4773 | -28.169 | 0
[6 empty] | 117 | 1 | CC3 28oct24 BOT | 0 | 0 | -4.2511 | -26.0193 | 0
[6 empty] | 118 | 2 | CC3 28oct24 BOT | 0 | 0 | -4.1467 | -26.0826 | 0
[6 empty] | 119 | 3 | CC3 28oct24 BOT | 0 | 0 | -4.1325 | -26.3023 | 0
[6 empty] | 120 | 4 | CC3 28oct24 BOT | 0 | 0 | -4.1259 | -26.3139 | 0
[6 empty] | 121 | 1 | CC2 11jul24 0.1m - DUP | 0 | 0 | -4.386 | -27.6647 | 0
[6 empty] | 122 | 2 | CC2 11jul24 0.1m - DUP | 0 | 0 | -4.4031 | -27.3843 | 0
[6 empty] | 123 | 3 | CC2 11jul24 0.1m - DUP | 0 | 0 | -4.4612 | -27.2016 | 0
[6 empty] | 124 | 4 | CC2 11jul24 0.1m - DUP | 0 | 0 | -4.3842 | -27.4319 | 0
[6 empty] | 125 | 1 | Hawaii | 0 | 0 | -3.6963 | -13.0431 | 0
[6 empty] | 126 | 2 | Hawaii | 0 | 0 | -3.6883 | -12.9448 | 0
[6 empty] | 127 | 3 | Hawaii | 0 | 0 | -3.6271 | -13.1384 | 0
[6 empty] | 128 | 4 | Hawaii | 0 | 0 | -3.7856 | -12.9775 | 0
[6 empty] | 129 | 1 | Blacksburg | 0 | 0 | -7.6082 | -47.972 | 0
[6 empty] | 130 | 2 | Blacksburg | 0 | 0 | -7.643 | -47.7887 | 0
[6 empty] | 131 | 3 | Blacksburg | 0 | 0 | -7.6205 | -48.0055 | 0
[6 empty] | 132 | 4 | Blacksburg | 0 | 0 | -7.6683 | -47.7414 | 0